## Sheet1
id | name | actions | Index
1362 | Fit Front Panel Logo | Clean Front Panel | 1
1362 | Fit Front Panel Logo | Prepare Logo | 2
1362 | Fit Front Panel Logo | Fitting Logo | 3
11246 | Production and Supplier Paperwork | All Paperwork Sorted and Stapled Together | 4
1328 | SCS Check | Check SCS Folder/Sales Order | 5
1328 | SCS Check | SCS Number (If applicable) | 6
10545 | SmartPI | Correct PC - 349535-9001-020 Installed | 7
10545 | SmartPI | Print Out SmartPI Check List (Link in Documents) | 8
10545 | SmartPI | Install SmartSEM Licences | 9
10545 | SmartPI | Install MIL Redistribution | 10
10545 | SmartPI | MIL Dongle Serial Number | 11
10545 | SmartPI | Fit MIL Dongle | 12
10545 | SmartPI | SmartPI Software Version | 13
10545 | SmartPI | Install SmartPI Software | 14
10545 | SmartPI | I/O-SCAN Card Fitted (Bruker Only) | 15
10545 | SmartPI | I/O-SCAN Card Calibrated | 16
10545 | SmartPI | Setup & Run Application | 17
10545 | SmartPI | Check Results & Image Montage | 18
10545 | SmartPI | Store Compressed Database,Montage Image, Log files etc | 19
3316 | SmartPI | Correct PC - 347823-9024-660 Installed | 20
3316 | SmartPI | Print Out SmartPI Check List (Link in Documents) | 21
3316 | SmartPI | Install SmartSEM Licences | 22
3316 | SmartPI | Install MIL Redistribution | 23
3316 | SmartPI | MIL Dongle Serial Number | 24
3316 | SmartPI | Fit MIL Dongle | 25
3316 | SmartPI | SmartPI Software Version | 26
3316 | SmartPI | Install SmartPI Software | 27
3316 | SmartPI | Calibrate I/O-SCAN card | 28
3316 | SmartPI | Setup & Run Application | 29
3316 | SmartPI | Check Results & Image Montage | 30
950 | Insitu stage | Fit the Insitu substage | 31
950 | Insitu stage | Test the Insitu substage | 32
950 | Insitu stage | System vac <3.0e-06mbar | 33
950 | Insitu stage | X-ray test | 34
950 | Insitu stage | Repackage all hardware excluding the chamber feed through in Deben packaging | 35
10733 | Bay | Clean and Tidy Assembly Bay | 36
766 | Material Availability | Are there any shortages? | 37
10734 | Fit Stage Assembly | Select Stage Type | 38
10734 | Fit Stage Assembly | Stage Serial Number | 39
10734 | Fit Stage Assembly | Shim Stage Door | 40
10734 | Fit Stage Assembly | Check Stage Pumps with No Hissing | 41
10734 | Fit Stage Assembly | Set Stage Clamp Height | 42
11414 | Print System Label | Confirm label has CSA mark and SN and PN are correct. | 43
10767 | Power Up | Rotary Pump Pumping | 44
10767 | Power Up | Penning Gauge LED is On | 45
10771 | Power Supply Set Up (ON LOAD) | Vac 5.15V (P1) | 46
10771 | Power Supply Set Up (ON LOAD) | Vac 5.00V (P10) | 47
10771 | Power Supply Set Up (ON LOAD) | h-ASAP +- 12.5V (Set on PSU) | 48
10771 | Power Supply Set Up (ON LOAD) | Stage 5.15V | 49
10771 | Power Supply Set Up (ON LOAD) | Stage 21.5V | 50
10771 | Power Supply Set Up (ON LOAD) | Stage -21.5V | 51
10771 | Power Supply Set Up (ON LOAD) | EO 5.15V | 52
10774 | Pressure Test | No Leaks Detected | 53
986 | Confirm 4-eye check by additional operator | Checked Carried Our By: | 54
10957 | Column Pumping Jig | Clean and Prepare the Jig and Fit the O ring | 55
10735 | Fit Stage Cabling | Fit Stage Cable | 56
10735 | Fit Stage Cabling | Fit Specimen Touch Cable | 57
10968 | C3 Outer Cone | Fit C3 Cone into Jig (Vac seal facing you) | 58
752 | Material Availability | Are there any shortages? | 59
10969 | C3 Inner Cone | Fit the Sealing Plate onto the Inner C3 Cone | 60
10969 | C3 Inner Cone | Fit Handles and Lower into the Outer Cone | 61
10969 | C3 Inner Cone | Secure the Outer C3 Cone to the Jig | 62
10969 | C3 Inner Cone | Secure the Inner C£ Cone to the Outer Cone | 63
10969 | C3 Inner Cone | Fit Scan Coils | 64
10969 | C3 Inner Cone | Fit FoV Coils | 65
10969 | C3 Inner Cone | Fit O Ring to Top of Scan Coil | 66
10969 | C3 Inner Cone | Perform Vacuum Leak Check | 67
10970 | Column Pumping Body | Fit Column Pumping Body onto the C3 Assembly | 68
10970 | Column Pumping Body | Fit Test Square Blanking Plate | 69
10970 | Column Pumping Body | Fit Mu-metal Shield | 70
10970 | Column Pumping Body | Fit Aperture Changer | 71
10970 | Column Pumping Body | Fit Circular Blank (PCM) | 72
10970 | Column Pumping Body | Fit Rectangular Blank (CIV) | 73
10974 | C1/C2 Body | Fit C1/C2 Body | 74
10974 | C1/C2 Body | Perform Vacuum Leak Check | 75
10975 | C3 Continuity Test Box | Scan Coil - LED is On | 76
10975 | C3 Continuity Test Box | Press Button and LED Goes Off | 77
10976 | Gun Head | Fit Firing Unit into Gun Head | 78
10976 | Gun Head | Fit O Ring into Top of C!/C2 Body | 79
10976 | Gun Head | Fit Gun Head to C1/C2 Body | 80
10977 | C1/C2 Body External Parts | Fit Gun Safety Switch | 81
10977 | C1/C2 Body External Parts | Fit Gun Align Cable To Bracket | 82
10977 | C1/C2 Body External Parts | Fit Assembled Bracket To C1/C2 Body | 83
10978 | C1/C2 Body Internal Parts | Fit Pole Pieces | 84
10978 | C1/C2 Body Internal Parts | Fit Bottom Plate | 85
10978 | C1/C2 Body Internal Parts | Fit Gun Align Coils | 86
10978 | C1/C2 Body Internal Parts | Fit Anode | 87
10978 | C1/C2 Body Internal Parts | Check for Cleanliness Using UV Torch | 88
10980 | Final Vacuum Check | Complete Column Vauum Checked | 89
10980 | Final Vacuum Check | Vacuum Isolation Valve Closed | 90
11117 | SCS Check | Check SCS Folder/Sales Order | 91
11117 | SCS Check | SCS Number (If applicable) | 92
961 | Declaration of conformity (EVO 15) | Please print 'Declaration of Conformity' from URL link | 93
962 | Declaration of conformity (EVO 25) | Please print 'Declaration of Conformity' from URL link | 94
10979 | System Type | Select System Type | 95
10979 | System Type | Column Type | 96
10942 | Circular Blank | Clean and Inspect O Ring | 97
10942 | Circular Blank | Fit O Ring into Blank | 98
10943 | Rectangular Blank | Clean and Inspect O Ring | 99
10943 | Rectangular Blank | Fit O Ring into Blank | 100
10798 | Clean Chamber | Clean internal surfaces with Carbaflow | 101
10798 | Clean Chamber | Clean internal surfaces with IPA | 102
10944 | Sealing Plate | Clean and Inspect O Rings | 103
10944 | Sealing Plate | Fit O Rings into Plate | 104
10945 | C3 Cone Vac Seal | Clean and Inspect O Ring | 105
10945 | C3 Cone Vac Seal | Fit O Ring into Vac Seal | 106
10945 | C3 Cone Vac Seal | Insert Vac Seal into C3 Cone | 107
10946 | Lens End Cap Assembly | Clean and Inspect O Ring | 108
10946 | Lens End Cap Assembly | Fit O Ring it to the Lens Cap | 109
10946 | Lens End Cap Assembly | Mount the End Cap Onto the C3 Cone | 110
10946 | Lens End Cap Assembly | Tighten the Screws | 111
10946 | Lens End Cap Assembly | End Cap and C3 Cone Gap Checked | 112
10947 | Gun Align Coil | Liner Tube Inspected | 113
10947 | Gun Align Coil | Inspect O Ring Sealing Surfaces | 114
10947 | Gun Align Coil | Liner Tube Cleaned | 115
10947 | Gun Align Coil | Clean and Inspect O Rings | 116
10947 | Gun Align Coil | Fit O Rings onto Coils | 117
479 | Scan Coil Auto Test | Auto Inductance Test Passed | 118
479 | Scan Coil Auto Test | Auto Flux Test Passed | 119
10948 | Scan Coils | Liner Tube Inspected | 120
10948 | Scan Coils | Inspect O Ring Sealing Surfaces | 121
10948 | Scan Coils | Liner Tube Cleaned | 122
10948 | Scan Coils | Clean and Inspect O Rings | 123
10948 | Scan Coils | Fit O Rings onto Coils | 124
10949 | Pumping Body | Clean and Inspect O Ring | 125
10949 | Pumping Body | Fit O Ring to the Underside of the Body | 126
10949 | Pumping Body | Fit Pumping Guide | 127
10949 | Pumping Body | Check Pumping Guide Alignment | 128
10949 | Pumping Body | Fit M3x10 Countersunk Screws | 129
10949 | Pumping Body | Fit Chamber Pumping Shield | 130
10949 | Pumping Body | Clean and Inspect O Ring | 131
10949 | Pumping Body | Fit O Ring to Upperside of the Body | 132
11198 | Fit Ski Jump and CIV | Fit Ski Jump Onto Plate | 133
11198 | Fit Ski Jump and CIV | Using Jig Align Ski Slope Bore With Plate Bore. | 134
11198 | Fit Ski Jump and CIV | Fit Bottom Plate Using 4x M4X8 CSK Screws | 135
11198 | Fit Ski Jump and CIV | Fit Metal Shield | 136
11198 | Fit Ski Jump and CIV | Insert Ski Jump Plate and Secure | 137
11198 | Fit Ski Jump and CIV | Fit O Ring | 138
11198 | Fit Ski Jump and CIV | Clean and Check CIV Sealing O Ring | 139
11198 | Fit Ski Jump and CIV | Clean, Check and Fit CIV Body O Ring | 140
11198 | Fit Ski Jump and CIV | Check CIV Height Using Letterbox Jig | 141
11198 | Fit Ski Jump and CIV | Fit CIV Onto Pumping Body | 142
11198 | Fit Ski Jump and CIV | Check CIV Can Open and Close Without Touching Insert Plate | 143
10798 | Clean Chamber | Clean internal surfaces with Carbaflow | 144
10798 | Clean Chamber | Clean internal surfaces with IPA | 145
11199 | Ion Pump | Fit Heater and Cover Plate | 146
11199 | Ion Pump | Fit the Large Cover Plate and Temperature Label | 147
11199 | Ion Pump | Remove Ion Pump Transit Blank | 148
11199 | Ion Pump | Fit Copper Seal | 149
11199 | Ion Pump | Fit and Secure Baffle | 150
11199 | Ion Pump | Align and Secure T Pipe | 151
11199 | Ion Pump | Fit Ion Pump Assy. Onto Column | 152
10950 | Aperture Changer | Inspect Aperture Blade | 153
10950 | Aperture Changer | Install the Apertures | 154
10951 | Firing Unit | Disassemble the Firing Unit | 155
10951 | Firing Unit | Clean Listed Parts in Ultrasonic Tank | 156
10951 | Firing Unit | Reassemble Firing Unit and Fit Filament | 157
10951 | Firing Unit | Check Filament Height and Centre | 158
10952 | Bottom Plate and Aperture Clamp (W) | Inspect Bores of Plate and Clamp | 159
10952 | Bottom Plate and Aperture Clamp (W) | Fit 600um Aperture | 160
10952 | Bottom Plate and Aperture Clamp (W) | Clean and Inspect O Ring and Fit to the Underside of Plate | 161
10954 | Bottom Plate and Aperture Clamp (LaB6) | Inspect Bores of Plate and Clamp | 162
10954 | Bottom Plate and Aperture Clamp (LaB6) | Fit 300um Aperture | 163
10954 | Bottom Plate and Aperture Clamp (LaB6) | Clean and Inspect O Ring and Fit to the Underside of Plate | 164
10953 | C1/C2 Pole Pieces | Fit an Aperture Insert Into Each Pole Piece | 165
10953 | C1/C2 Pole Pieces | Check Aperture Inserts are Securely Fitted | 166
416 | Vacuum Leak Test | No Leaks Detected (With Turbo running) when in high sensitivity | 167
416 | Vacuum Leak Test | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng | 168
10816 | Fit Column | Correct Chamber Bung Fitted | 169
416 | Vacuum Leak Test | No Leaks Detected (With Turbo running) when in high sensitivity | 170
416 | Vacuum Leak Test | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng | 171
1134 | Print Shipment, Packing Labels and Identify System Material | Print Shipment Labels | 172
1134 | Print Shipment, Packing Labels and Identify System Material | Label Material Boxes | 173
1134 | Print Shipment, Packing Labels and Identify System Material | Place left over stickes in Document Holder | 174
960 | Declaration of conformity (EVO 10) | Please print 'Declaration of Conformity' from URL link | 175
11096 | SCS Check | Check SCS Folder/Sales Order | 176
11096 | SCS Check | SCS Number (If applicable) | 177
402 | Label serial port cable com ports | Fit labels to PC serial cable com port 4 to 7 | 178
6935 | PC MAC Address | Record MAC address (Internet not MIC-NET) | 179
890 | Change USB Card | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details) | 180
10463 | Record PC Details | Select PC | 181
10463 | Record PC Details | PC Serial No | 182
10467 | Install SmartSEM + ASAP Engine | Install SmartSEM + ASAP Engine | 183
10467 | Install SmartSEM + ASAP Engine | Copy the SmartSEM sotware into the SmartSEM Software folder (C:Support/Data/SmartSEM Software) | 184
10467 | Install SmartSEM + ASAP Engine | SmartSEM Version | 185
10467 | Install SmartSEM + ASAP Engine | SmartSEM Service Pack | 186
10467 | Install SmartSEM + ASAP Engine | SmartSEM Hotfix | 187
10467 | Install SmartSEM + ASAP Engine | ASAP Driver Scan Control Service Version | 188
10467 | Install SmartSEM + ASAP Engine | Installation of Zen Core EM if Hasap is fitted | 189
1167 | Install Production Test Licence | Check that Time & Date on PC are set to current UK | 190
1167 | Install Production Test Licence | Download Licence from SharePoint | 191
1167 | Install Production Test Licence | Delete the Licence in SharePoint | 192
1167 | Install Production Test Licence | Install the licence using the Zeiss License Management Tool. | 193
1167 | Install Production Test Licence | hASAP Only - Installation of ZEN Core EM licence using the Zeiss activation tool | 194
993 | Wiring, Micro switches and Cable form Fitment. (Set Micro-switch Positions) | Z Microswitch Position Set | 195
993 | Wiring, Micro switches and Cable form Fitment. (Set Micro-switch Positions) | T Microswitch Position Set | 196
993 | Wiring, Micro switches and Cable form Fitment. (Set Micro-switch Positions) | R Microswitch Position Set | 197
994 | Fit ZTR Motors and place the appropriate sticker in the box provided be-low: | Z Motor SeriaL Number: | 198
994 | Fit ZTR Motors and place the appropriate sticker in the box provided be-low: | T Motor Serial Number: | 199
994 | Fit ZTR Motors and place the appropriate sticker in the box provided be-low: | R Motor Serial Number: | 200
10465 | Install and SetUp Windows 10 | Windows 10 Version | 201
10465 | Install and SetUp Windows 10 | Disable Drivers and Windows Updates | 202
10464 | Power Up and Enter Bios | Check That AC Loss Policy Depend On, is Set To Power On | 203
10464 | Power Up and Enter Bios | Disable Windows Fast Start UP | 204
10464 | Power Up and Enter Bios | Set the Time and Date on the PC to current UK | 205
1276 | Setup the Support folder | Select the Support folder in the root dir (c:) | 206
1276 | Setup the Support folder | Create a Data folder within the Support folder | 207
1276 | Setup the Support folder | In the Data folder create the following new sub-folders- Images, Licences, Log files, SmartSEM Software. | 208
11272 | Install Predictive Service Software | Predictive Service S/W installed on Customer PC. | 209
11272 | Install Predictive Service Software | Do Not Activate | 210
11272 | Install Predictive Service Software | Copy the Predictive Service.json file into C:Support/Data/Licence folder. | 211
8699 | Cable Up | System Mains Voltage | 212
8699 | Cable Up | IF LaB6 and 120V - Transformer Installed | 213
8699 | Cable Up | Cable Up | 214
8699 | Cable Up | Label Cables | 215
8699 | Cable Up | Tape PC Configuration Image to PC | 216
10470 | Administrator Factory Setup-Serial Number | Enter System Serial number into Administrator. | 217
10470 | Administrator Factory Setup-Serial Number | Check the Correct Serial number is on the Rear Silver S/N Label | 218
759 | Run EO Test From Service Panel | Run EO Test | 219
10476 | Check Firmware/Hardware Status of Hardware | Check ASAP connection | 220
10476 | Check Firmware/Hardware Status of Hardware | Check Stage Connection | 221
10476 | Check Firmware/Hardware Status of Hardware | Check EHT Set connection | 222
10476 | Check Firmware/Hardware Status of Hardware | Check Vacuum PCB connection | 223
10476 | Check Firmware/Hardware Status of Hardware | Check EO PCB connection | 224
1285 | Set EO Temp Limit to 40 - Hittech Stage Only | Set EO Temp Limit to 40 deg C | 225
10481 | Ion Pump System Selected | Fit Link on Vacuum PCB | 226
10481 | Ion Pump System Selected | Gun Chamber Valve Motorised=Yes | 227
10481 | Ion Pump System Selected | Check that when Ion Pump is Pumping, the Vacuum is Improving. | 228
10481 | Ion Pump System Selected | Gun/ Chamber Valve Closes ≈ 2.0e-06 | 229
10482 | Zero MKS Baratron Gauge | Baratron Gauge Set To Zero | 230
10482 | Zero MKS Baratron Gauge | 2nd Rotary Pump Is NOT Running | 231
10483 | Set Up MKS Driver Unit | UnPlug PL4 on Vac PCB | 232
10483 | Set Up MKS Driver Unit | Check That 2nd Rotary Pump is OFF in HV | 233
10483 | Set Up MKS Driver Unit | Go To VP, 2nd Rotary Pump is ON | 234
10483 | Set Up MKS Driver Unit | Set Up Driver Box | 235
10483 | Set Up MKS Driver Unit | Go TO HV, 2nd Rotary Pump Switched OFF | 236
10483 | Set Up MKS Driver Unit | ReFit PL4 on Vac PCB | 237
10485 | Overnight Stage Macro ( DO Not Do On Hittech Stages) | Minimum of 8 Hours (Cartesian Stage) | 238
10487 | Condition Gun - Tungsten | Connect the Hitek Test Tool EHT Set | 239
10487 | Condition Gun - Tungsten | Gun Monitor Setup | 240
10487 | Condition Gun - Tungsten | Test No. 96 Selected | 241
10487 | Condition Gun - Tungsten | EHT Stable at 30KV - ≥15 minutes | 242
10487 | Condition Gun - Tungsten | 30KV Gun Monitor Traced Saved | 243
10488 | Obtain Image | Ensure Gun Monitor is Running and Test 96 Selected. | 244
10488 | Obtain Image | Carry Out Initial Gun Run Up | 245
10488 | Obtain Image | Connect the Customers Spellman EHT Set | 246
10488 | Obtain Image | Carry Out Gun Run Up | 247
10488 | Obtain Image | X Gun Tilt ≤50% | 248
10488 | Obtain Image | Y Gun Tilt ≤50% | 249
10488 | Obtain Image | X Gun Shift ≤50% | 250
10488 | Obtain Image | Y Gun Shift ≤50% | 251
1273 | Mains Sync- Set to OFF | Mains Sync set to OFF | 252
10486 | Set Up Anti-Vibration Air Feet and Jack Feet | Jacking Feet Set | 253
10486 | Set Up Anti-Vibration Air Feet and Jack Feet | Air Feet Inflated | 254
10486 | Set Up Anti-Vibration Air Feet and Jack Feet | Rear Plinth Fan is Running in Operate and Standby | 255
10489 | Check Function of SE Detector | Scintillator | 256
10489 | Check Function of SE Detector | Collector Volts | 257
10489 | Check Function of SE Detector | Check that the Signal Level in Line Mode is at Black Level with Brightness=50% and Contrast =0% | 258
10489 | Check Function of SE Detector | Check that the Signal Level in Line Mode is at White Level with Brigfhtness =62% and Contrast =0% | 259
10492 | Column Alignment | C3 Reversal-≤20µm | 260
10492 | Column Alignment | 30KV-5KV-≤20µm | 261
10492 | Column Alignment | Image Shift-≤3000nm | 262
10494 | Column Calibration C1/C2 | C1 Cal | 263
10494 | Column Calibration C1/C2 | C1 Rem | 264
10494 | Column Calibration C1/C2 | C1 Xtalk | 265
10494 | Column Calibration C1/C2 | C2 Cal | 266
10494 | Column Calibration C1/C2 | C2 Rem | 267
10494 | Column Calibration C1/C2 | C2 Xtalk | 268
10495 | Calibrate C3 | C3 Calibration | 269
10495 | Calibrate C3 | 30KV to 2KV Focus Check- ≤0.2mm change in WD. | 270
10496 | Initial/Final X-Ray Check | Accelerating Voltage=30KV | 271
10496 | Initial/Final X-Ray Check | Specimen Current | 272
10496 | Initial/Final X-Ray Check | Monitor Model | 273
10496 | Initial/Final X-Ray Check | Serial Number | 274
10496 | Initial/Final X-Ray Check | Cal. Expiry Date | 275
10496 | Initial/Final X-Ray Check | No X-ray Leaks Detected | 276
10496 | Initial/Final X-Ray Check | Complete & Sign X-Ray Certificate | 277
10760 | X-Ray Test | Serial No. (or N/A) | 278
10760 | X-Ray Test | No X-Ray Leakage Detected | 279
10497 | Fit Lab6 Firing Unit/Bakeout | System Vacuum & Gun Vacuum is within Spec | 280
10497 | Fit Lab6 Firing Unit/Bakeout | Gun Chamber has been cleaned and inspected using UV light | 281
10497 | Fit Lab6 Firing Unit/Bakeout | EHT set socket well has been cleaned and inspected using UV light | 282
10497 | Fit Lab6 Firing Unit/Bakeout | Clean Anode in Plasma Oven for 20 Minutes | 283
10497 | Fit Lab6 Firing Unit/Bakeout | Check Cathode is aligned before fitting | 284
10497 | Fit Lab6 Firing Unit/Bakeout | Lab6 Firing Unit Fitted | 285
10497 | Fit Lab6 Firing Unit/Bakeout | Overnight Bakeout | 286
10498 | Condition Gun - LAB6 | Conect the Hitek EHT Set | 287
10498 | Condition Gun - LAB6 | Gun Monitor Setup | 288
10498 | Condition Gun - LAB6 | Test No. 96 Selected | 289
10498 | Condition Gun - LAB6 | Run up to 30KV and leave for 15 minute | 290
10498 | Condition Gun - LAB6 | Shutdown and run up to 30KV, Repeat 5 times | 291
10498 | Condition Gun - LAB6 | Run up to 30KV and leave for 30 minutes (There must be no spikes) | 292
10498 | Condition Gun - LAB6 | 30KV Gun Monitor Traced Saved | 293
10500 | Lab6 Firing Unit Run Up | Gun Monitor Setup | 294
10500 | Lab6 Firing Unit Run Up | Test No. 96 Selected | 295
10500 | Lab6 Firing Unit Run Up | New Filament Selected | 296
10500 | Lab6 Firing Unit Run Up | Long Fil. Life Selected | 297
10500 | Lab6 Firing Unit Run Up | Carry Out Initial LaB6 Gun Run Up to 30KV | 298
10500 | Lab6 Firing Unit Run Up | Connect the Customers Spellman EHT Set | 299
10500 | Lab6 Firing Unit Run Up | Run the LaB6 to 30KV | 300
10500 | Lab6 Firing Unit Run Up | X Gun Tilt ≤50% | 301
10500 | Lab6 Firing Unit Run Up | Y Gun Tilt ≤50% | 302
10500 | Lab6 Firing Unit Run Up | X Gun Shift ≤50% | 303
10500 | Lab6 Firing Unit Run Up | Y Gun Shift ≤50% | 304
11195 | LaB6 Beam Current | Check That 20uA (Spec: ±3) is seen in the Gun Monitor (Beam I Monitor) | 305
10502 | Stigmator Check - Standard | 30KV @ 20mm X- ≤35% | 306
10502 | Stigmator Check - Standard | 30KV @ 20mm Y- ≤35% | 307
10502 | Stigmator Check - Standard | 2KV @ 20mm X- ≤35% | 308
10502 | Stigmator Check - Standard | 2KV @ 20mm Y- ≤35% | 309
10502 | Stigmator Check - Standard | 30KV @ 5mm X- ≤35% | 310
10502 | Stigmator Check - Standard | 30KV @ 5mm Y- ≤35% | 311
10502 | Stigmator Check - Standard | 2KV @ 5mm X- ≤35% | 312
10502 | Stigmator Check - Standard | 2KV @ 5mm Y- ≤35% | 313
10504 | Stigmator Check - VP Aperture | VP 30KV @ 5mm (100um) X- ≤35% | 314
10504 | Stigmator Check - VP Aperture | VP 30KV @ 5mm (100um) Y- ≤35% | 315
10505 | Stimator Check- VP Aperture (140um) | VP 30KV @ 5mm (140um) X- ≤35% | 316
10505 | Stimator Check- VP Aperture (140um) | VP 30KV @ 5mm (140um) Y- ≤35% | 317
10506 | VP Functional Checks | Vp Functional Check | 318
10506 | VP Functional Checks | Vac Status = Ready and EHT Vac Ready = Yes, are always present at all VP values. (i.e. The beam remains on/can be switched on). | 319
10506 | VP Functional Checks | Fit Label to the Black Plastic Aperture Box | 320
10507 | Stigmator & Alignment Check - EP Apertures | EP 500um 30KV @ 5mm X- ≤35% | 321
10507 | Stigmator & Alignment Check - EP Apertures | EP 500um 30KV @ 5mm Y- ≤35% | 322
10507 | Stigmator & Alignment Check - EP Apertures | EP 500um Alignment- ≤55% | 323
10507 | Stigmator & Alignment Check - EP Apertures | EP 1000um 30KV @ 5mm X- ≤35% | 324
10507 | Stigmator & Alignment Check - EP Apertures | EP 1000um 30KV @ 5mm Y- ≤35% | 325
10507 | Stigmator & Alignment Check - EP Apertures | EP 1000um Alignment- ≤55% | 326
10656 | TTL Functional Check | Vp Functional Check | 327
10656 | TTL Functional Check | Vac Status = Ready and EHT Vac Ready = Yes, are always present at all VP values. (i.e. The beam remains on/can be switched on). | 328
10508 | Stigmator & Alignment Check -BeamSleeve (500um) | BeamSleeve 500um 30KV @ 8mm X- ≤35% | 329
10508 | Stigmator & Alignment Check -BeamSleeve (500um) | BeamSleeve 500um 30KV @ 8mm Y- ≤35% | 330
10508 | Stigmator & Alignment Check -BeamSleeve (500um) | BeamSleeve 500um Alignment- ≤70% | 331
10509 | Stigmator & Alignment Check - BeamSleeve (1000um) | BeamSleeve 1000um 30KV @ 8mm X- ≤35% | 332
10509 | Stigmator & Alignment Check - BeamSleeve (1000um) | BeamSleeve 1000um 30KV @ 8mm Y- ≤35% | 333
10509 | Stigmator & Alignment Check - BeamSleeve (1000um) | BeamSleeve 1000um Alignment- ≤70% | 334
10510 | EP & Water Function Check | 500um Lower Aperture Fitted | 335
10510 | EP & Water Function Check | 2nd Rotary Pump is ON in EP Mode | 336
10510 | EP & Water Function Check | 10Pa Chamber Vacuum Achieved | 337
10510 | EP & Water Function Check | 3000Pa Vacuum Achieved (Air). | 338
10510 | EP & Water Function Check | Vac Status = Ready and EHT Vac Ready = Yes, at 3000Pa (i.e. The beam remains on/can be switched on) | 339
10510 | EP & Water Function Check | Water Function Checked | 340
10510 | EP & Water Function Check | Approx. 2000Pa Chamber Vacuum Achieved with Water | 341
10510 | EP & Water Function Check | Vac Status = Ready and EHT Vac Ready = Yes, at 2000Pa (i.e. The beam remains on/can be switched on) | 342
10510 | EP & Water Function Check | Go To HV, 2nd Rotary Pump is OFF | 343
1226 | Hasap scanning calibration using Zen | carry out scanning calibration (Hasap only) | 344
10512 | Stigmator Check-400um EasyVP Aperture | 30KV @ 5mm X- ≤35% | 345
10512 | Stigmator Check-400um EasyVP Aperture | 30KV @ 5mm Y- ≤35% | 346
10513 | Stigmator Check-750um EasyVP Aperture | 30KV @ 5mm X- ≤35% | 347
10513 | Stigmator Check-750um EasyVP Aperture | 30KV @ 5mm Y- ≤35% | 348
10880 | Stigmator Check-750um Easy Align Aperture | 30KV @ 5mm X- ≤35% | 349
10880 | Stigmator Check-750um Easy Align Aperture | 30KV @ 5mm Y- ≤35% | 350
10761 | EasyVP/Easy Align Functional Check | Functional Check | 351
10762 | Easy VP Magnification Calibration | Easy VP Mag. Cal Complete | 352
1283 | Check All Stage Axis For Travel | X Axis | 353
1283 | Check All Stage Axis For Travel | Y Axis | 354
1283 | Check All Stage Axis For Travel | Z Axis | 355
1283 | Check All Stage Axis For Travel | T Axis | 356
1284 | Stage X and Y Macro - Hittech Stage Only | X/Y Macro Run For 30 Minutes | 357
10929 | Calibrate Compucentric Stage | Calibrate Stage Centre (Right Click on Stgae tab) | 358
10929 | Calibrate Compucentric Stage | Copy the Centre Pos X and Y values into the Adminitrator, Stage Post Init. Position. | 359
10929 | Calibrate Compucentric Stage | Tilt Centre Calibration (Located in Docking panel) | 360
1250 | Stage Repeatability Macro - Hittech Stage Only | Download Macro folder | 361
1250 | Stage Repeatability Macro - Hittech Stage Only | Run Kompass 3.08.exe | 362
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 1 ($[StageRepeatability]) | 363
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 2 ($[StageRepeatability]) | 364
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 3 ($[StageRepeatability]) | 365
1250 | Stage Repeatability Macro - Hittech Stage Only | Image of Macro Results | 366
10516 | Stage Acceptance Check X and Y Axis - Cartesian Stages Only | X Axis | 367
10516 | Stage Acceptance Check X and Y Axis - Cartesian Stages Only | Y Axis | 368
10516 | Stage Acceptance Check X and Y Axis - Cartesian Stages Only | X+Y Axis | 369
10516 | Stage Acceptance Check X and Y Axis - Cartesian Stages Only | Y+X Axis | 370
10516 | Stage Acceptance Check X and Y Axis - Cartesian Stages Only | 10mm Square | 371
10517 | Stage Acceptance Check Z Axis - Cartesian Stages Only | Z Axis | 372
10518 | Stage Backlash Measurement X & Y - Cartesian Stages Only | X Axis | 373
10518 | Stage Backlash Measurement X & Y - Cartesian Stages Only | Y Axis | 374
10519 | Compucentric Stage Test Rotation | Rotate=20 deg | 375
10519 | Compucentric Stage Test Rotation | Rotate=45 deg | 376
10519 | Compucentric Stage Test Rotation | Rotate=90 deg | 377
10519 | Compucentric Stage Test Rotation | Rotate=180 deg | 378
10520 | Compucentric Stage Tests - Tilt | Tilt=5 deg | 379
10520 | Compucentric Stage Tests - Tilt | Tilt=10 deg | 380
10520 | Compucentric Stage Tests - Tilt | Tilt=20 deg | 381
10520 | Compucentric Stage Tests - Tilt | Tilt=45 deg | 382
631 | VP/EP Function | VP/EP Function Checked | 383
631 | VP/EP Function | Check Beam Is On at Max VP/EP Pressure | 384
10524 | W System Pumpdown Time | Pumpdown Time | 385
10525 | Lab6 System Pumpdown Time | Pumpdown Time | 386
1278 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 387
1278 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 388
1278 | ZEN Core EM | Install Customer's ZEN Licence | 389
1278 | ZEN Core EM | Setup SmartSEM Toggle Button | 390
1278 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 391
1278 | ZEN Core EM | Mains Sync Set To OFF | 392
1278 | ZEN Core EM | Select Your ZEN core EM Runtime Profile | 393
1278 | ZEN Core EM | Create User Supervisor EM | 394
1278 | ZEN Core EM | Create User Supervisor LM | 395
1278 | ZEN Core EM | Place both users as “Member of” Supervisor group | 396
1278 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 397
1278 | ZEN Core EM | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems. | 398
1278 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 399
1278 | ZEN Core EM | Configue Airlock Cam and Setup Exchange Position $Exchange (use Upper case E in ZEN) | 400
1278 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 401
1278 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 402
1278 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 403
1278 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 404
1278 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 405
1278 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 406
84 | Xplore 30EDS | Language Set to English (United States | 407
84 | Xplore 30EDS | Ethernet Connection Setup | 408
84 | Xplore 30EDS | Install Aztec Software | 409
84 | Xplore 30EDS | AZtec Software Version | 410
84 | Xplore 30EDS | Install Licences | 411
84 | Xplore 30EDS | Run AZtec Assist+ | 412
84 | Xplore 30EDS | Connected to Test Tool Module | 413
10526 | Chamberscope | Number of Chamberscope's on Order | 414
10526 | Chamberscope | Chamberscope 1 - Select Port | 415
10526 | Chamberscope | Check Camera for focus | 416
10526 | Chamberscope | Functional Check-Black/White + Colour Mode | 417
10526 | Chamberscope | Chamberscope 2 - Select Port or Not Fitted | 418
10526 | Chamberscope | Chamberscope 2 - Check Camera 2 for Focus | 419
10526 | Chamberscope | Chamberscope 2 - Functional Check-Black/White + Colour Mode | 420
10526 | Chamberscope | Chamberscope 2 - Run wxPropview to Remove Defective Pixels | 421
10526 | Chamberscope | Chamberscope 2 - Add Icon to Toolbar | 422
11100 | Chamberscope 2 | Select Port | 423
11100 | Chamberscope 2 | Check Camera for focus | 424
11100 | Chamberscope 2 | Functional Check-Black/White + Colour Mode | 425
11100 | Chamberscope 2 | Add Icon to Toolbar | 426
10869 | Camera Stand | Camera Model | 427
10869 | Camera Stand | Install Camera Software | 428
632 | 2nd Rotary Pump Check | 2nd Rotary Pump is OFF in HV Mode | 429
632 | 2nd Rotary Pump Check | 2nd Rotary Pump is ON in VP/EP Mode | 430
10866 | Additional Hard Disk & Cloning Software | Install Hard Disk | 431
10866 | Additional Hard Disk & Cloning Software | Format Hard Disk (NTFS) | 432
10866 | Additional Hard Disk & Cloning Software | Hard Disk Is Seen In File Explorer | 433
10528 | Control Panel | Installation | 434
10528 | Control Panel | select Language | 435
10528 | Control Panel | Function of all knobs and buttons | 436
10528 | Control Panel | Function of all keyboard keys | 437
10530 | VPSE G4 | Calibration of Bias | 438
10530 | VPSE G4 | Functional Check | 439
10530 | VPSE G4 | Function Image | 440
10530 | VPSE G4 | Check Beam Is On at Max VP Pressure | 441
10530 | VPSE G4 | Store VPSE Functional Image | 442
10529 | BSD/BSD1 | BSD Diode Pre-test | 443
10529 | BSD/BSD1 | Select Type | 444
10529 | BSD/BSD1 | Select Amplifier Type | 445
10529 | BSD/BSD1 | ES-BSD Only - Using ASAP Test Panel Update the ID to 336 | 446
10529 | BSD/BSD1 | Black Level Calibration | 447
10529 | BSD/BSD1 | Check Diode Segment Function | 448
10529 | BSD/BSD1 | Check All Gain ranges | 449
10529 | BSD/BSD1 | Check Auto range | 450
10529 | BSD/BSD1 | Check Auto Level (Does Not Function With BSD1) | 451
10529 | BSD/BSD1 | Setup and Check TOPO | 452
10529 | BSD/BSD1 | Take 3 images - Atomic, 5pA, 2KV | 453
10529 | BSD/BSD1 | Store BSD Functional Images-Atomic, 5pA, 2KV | 454
10531 | C2D/C2D F2 | Select Type | 455
10531 | C2D/C2D F2 | Black Level Cal | 456
10531 | C2D/C2D F2 | C2D Bias Target Check | 457
10531 | C2D/C2D F2 | Gain Range Check | 458
10531 | C2D/C2D F2 | Auto Level Check | 459
10531 | C2D/C2D F2 | Auto Range Check | 460
10531 | C2D/C2D F2 | Function Image | 461
10531 | C2D/C2D F2 | Check Beam Is On at Max VP Pressure | 462
10531 | C2D/C2D F2 | Store C2D Functional Image | 463
10532 | Fisheye | Set-Up (Mag.Cal. Wizard | 464
10532 | Fisheye | Functional Check | 465
10533 | Specimen Current Meter | Functional Check | 466
10534 | C2DX | Alignment of Ring (Using Jig) | 467
10534 | C2DX | Black Level Calibration | 468
10534 | C2DX | Functional Check | 469
10534 | C2DX | Store C2DX Functional Images - Air, water | 470
10536 | Beam Deceleration | Check 5KV is achievable using the Bias Slider | 471
10536 | Beam Deceleration | safety Check - Remove PL1 (Vac Ready from vac switch | 472
10536 | Beam Deceleration | Safety Check - Remove PL2 (Turbo Speed signal) | 473
10536 | Beam Deceleration | Safety Check - Remove PL4 (CAN Bus signal) | 474
10536 | Beam Deceleration | Functional Check and Images (Using the BSD Detector) | 475
10536 | Beam Deceleration | Store Beam De-Cel Functional Images - OV and 5KV | 476
10645 | Indigo Filter | Filter Has Been Fitted | 477
10535 | Specimen Current Detector | Black Level Calibration | 478
10535 | Specimen Current Detector | Functional Check | 479
10535 | Specimen Current Detector | Function Image | 480
10535 | Specimen Current Detector | Store SCD Functional Image | 481
10537 | CL Detector | Functional Check | 482
10537 | CL Detector | Function Image | 483
10537 | CL Detector | Store CL Functional Image | 484
10539 | Probe Current Measure | Probe Current Measurment Select in Administrator | 485
10539 | Probe Current Measure | Functional Check | 486
10539 | Probe Current Measure | Probe Current Monitor and Specimen Current Meter Read the same | 487
10539 | Probe Current Measure | Checked at 500pA. 1nA, 10nA | 488
10539 | Probe Current Measure | Picoammeter and Probe Current Monitor read the same | 489
10540 | Beam Blanking | Fit 300um Aperture - Required for Tunsten Systems only | 490
10540 | Beam Blanking | Installation | 491
10540 | Beam Blanking | Cable Open Circuit to Ground | 492
10540 | Beam Blanking | Blanking Tests Using 20um Aperture | 493
10540 | Beam Blanking | Blanking Tests Using 30um Aperture | 494
10540 | Beam Blanking | Blank/UnBlank Ratio is better than 0.1% for all settings | 495
10540 | Beam Blanking | Low Mag Blanking Check | 496
10541 | Quiet Mode | Installation of Quiet Mode Vessel | 497
10541 | Quiet Mode | Quiet Mode Fitted Selected in SEM Status | 498
10541 | Quiet Mode | Compare On/Off Times With Graph | 499
10541 | Quiet Mode | Pump Starts When Quiet Mode is Deselected | 500
10541 | Quiet Mode | Run Gun Monitor-Turbo Power and System Vacuum | 501
10542 | Video Input Kit | Tested Using Video Mix | 502
10543 | Nav Cam | Ferrite Fitted To Cables - PC End | 503
10543 | Nav Cam | $NavCam Postion Setup | 504
10543 | Nav Cam | Image Navigation Setup | 505
10543 | Nav Cam | Functional Check | 506
10527 | Dual Joystick | Basic Movements-T/Z Joystick | 507
10527 | Dual Joystick | Basic Movements-X/Y/R Joystick | 508
10527 | Dual Joystick | Initialise Stage | 509
10650 | Styled Joystick | Basic Movements-T/Z Joystick | 510
10650 | Styled Joystick | Basic Movements-X/Y/Z Joystick | 511
10650 | Styled Joystick | Initialise Stage | 512
10650 | Styled Joystick | Check Break Button | 513
10544 | SmartStitch | Installation | 514
10544 | SmartStitch | Functional Test | 515
10544 | SmartStitch | Store SmartStitch Functional Image | 516
10544 | SmartStitch | SmartStitch Folder is present on Desktop | 517
10651 | Smart Browse | Installation | 518
10651 | Smart Browse | Functional Test | 519
10652 | SmartEDX | Type of Detector | 520
10652 | SmartEDX | Detector Serial number | 521
10652 | SmartEDX | DDP Module Serial Number | 522
10652 | SmartEDX | Fit Flange & Mount Detector | 523
10652 | SmartEDX | Install EDS PCB's Into PC | 524
10652 | SmartEDX | Installation of DPP Module & Cables | 525
10652 | SmartEDX | Configure PC Network & Test | 526
10652 | SmartEDX | Disable Windows10 Fast Startup | 527
10652 | SmartEDX | Install ApexKey.edx Licences | 528
10652 | SmartEDX | Install Genesis.edx Licences (if Applicable) | 529
10652 | SmartEDX | Fit Dongle | 530
10652 | SmartEDX | APEX Software Version | 531
10652 | SmartEDX | Install APEX Software | 532
10652 | SmartEDX | Install Remoting | 533
10652 | SmartEDX | Setup RConfigure | 534
10652 | SmartEDX | Run API Test | 535
10652 | SmartEDX | Configure Imaging-Run SGShell | 536
10652 | SmartEDX | Calibrate Detector-Run eDPP Shell | 537
10652 | SmartEDX | Save Calibration Report | 538
10652 | SmartEDX | Detector Resolution (measured using Cu) ≈ 129 eV | 539
10652 | SmartEDX | Store Calibration Report | 540
10652 | SmartEDX | Save Calibration on Exit | 541
10652 | SmartEDX | Run APEX | 542
10652 | SmartEDX | Check Spectrum | 543
10652 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected | 544
10652 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected | 545
10652 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected | 546
10652 | SmartEDX | Check Mapping | 547
10652 | SmartEDX | Check Line Scanning | 548
10652 | SmartEDX | Save Aquired Spectrum, Mappimg & Line Scanning to the Image Folder | 549
10652 | SmartEDX | Store Aquired Spectrum (STD, Quick, HD), Mappimg & Line Scanning Images | 550
10652 | SmartEDX | Save Calibration Data to the folder C: Support/Data/ Log Files | 551
10652 | SmartEDX | Store Calibration Data File | 552
10652 | SmartEDX | Save Calibration Data to the SmartSEM Image Folder | 553
10653 | Bruker EDS Detector Install | Number of Detectors | 554
10653 | Bruker EDS Detector Install | Detector Serial Number | 555
10653 | Bruker EDS Detector Install | Detector Resolution | 556
10653 | Bruker EDS Detector Install | Detector Area | 557
10653 | Bruker EDS Detector Install | Detector Installed | 558
10653 | Bruker EDS Detector Install | In/Out Stops Set | 559
10653 | Bruker EDS Detector Install | Isolation Check of Detector & Column >2MΩ | 560
10653 | Bruker EDS Detector Install | Install Remoting & Setup Network Connection | 561
10653 | Bruker EDS Detector Install | Bruker Esprit Software Version | 562
10653 | Bruker EDS Detector Install | Install Bruker Esprit Software | 563
10653 | Bruker EDS Detector Install | Bruker EDS Manager Version | 564
10653 | Bruker EDS Detector Install | Install EDS Manager | 565
10653 | Bruker EDS Detector Install | Install Bruker Licences | 566
10653 | Bruker EDS Detector Install | Setup Pulse Processor | 567
10653 | Bruker EDS Detector Install | Check Standard Timing Cards Fitted | 568
10653 | Bruker EDS Detector Install | List Optional Timing Cards Fitted | 569
10653 | Bruker EDS Detector Install | Install Detector Data | 570
10653 | Bruker EDS Detector Install | Check Microscope driver | 571
10653 | Bruker EDS Detector Install | Setup SEM Conditions | 572
10653 | Bruker EDS Detector Install | Determine Analytical Working Distance | 573
10653 | Bruker EDS Detector Install | Calibrate Detector | 574
10653 | Bruker EDS Detector Install | Calibrate Image in Espirit | 575
11346 | Oxford EDS Detector Install | Setup Ethernet Connection | 576
11346 | Oxford EDS Detector Install | Install Aztec Software | 577
11346 | Oxford EDS Detector Install | Aztec Software Version | 578
11346 | Oxford EDS Detector Install | Licence Installation | 579
11346 | Oxford EDS Detector Install | Run Aztec Assist+ | 580
11346 | Oxford EDS Detector Install | Install Detector | 581
11346 | Oxford EDS Detector Install | Imaging Setup | 582
11346 | Oxford EDS Detector Install | Stage Calibration Setup | 583
11346 | Oxford EDS Detector Install | Calibrate Detector | 584
11346 | Oxford EDS Detector Install | Save Aztec Sign Off PDF and Calibration Folder to the Product Folder (O: drive) | 585
11346 | Oxford EDS Detector Install | Store Aztec Sign Off PDF and Calibration Folder | 586
10545 | SmartPI | Correct PC - 349535-9001-020 Installed | 587
10545 | SmartPI | Print Out SmartPI Check List (Link in Documents) | 588
10545 | SmartPI | Install SmartSEM Licences | 589
10545 | SmartPI | Install MIL Redistribution | 590
10545 | SmartPI | Fit MIL Dongle | 591
10545 | SmartPI | SmartPI Software Version | 592
10545 | SmartPI | Install SmartPI Software | 593
10545 | SmartPI | I/O-SCAN Card Fitted (Bruker Only) | 594
10545 | SmartPI | I/O-SCAN Card Calibrated | 595
10545 | SmartPI | Setup & Run Application | 596
10545 | SmartPI | Check Results & Image Montage | 597
10545 | SmartPI | Store Compressed Database,Montage Image, Log files etc | 598
10654 | Mineralogic | Print Out Mineralogic Check List (Link in Documents) | 599
10654 | Mineralogic | Correct PC - Installed | 600
10654 | Mineralogic | Install SmartSEM Licences | 601
10654 | Mineralogic | Instal MIL Redistribution | 602
10654 | Mineralogic | Fit MIL Dongle | 603
10654 | Mineralogic | Install Mineralogic Software | 604
10654 | Mineralogic | Install MAC Standard Recipe Files | 605
10654 | Mineralogic | Calibrate I/O-SCAN card | 606
10654 | Mineralogic | MAC Standard & Stage Postion Set | 607
10654 | Mineralogic | Adjust Recipe for the System | 608
10654 | Mineralogic | Run 20 min Rapid MAC STD Measure | 609
10654 | Mineralogic | Check Results & Image Montage - 20 min Rapid MAC STD Measure | 610
10654 | Mineralogic | Run 90 min High Mag Rapid MAC STD Measure | 611
10654 | Mineralogic | Check Results and Compare with 20 min Run | 612
10654 | Mineralogic | 10 Hour Overnight Run | 613
10654 | Mineralogic | Check Results and Compare with 20 and 90 min Run | 614
10654 | Mineralogic | Store Compressed Database,Montage Image, Log files etc | 615
11200 | Large ZTR Module | Select EVO25-L-ZTR. in Administaror as Stage Type | 616
11200 | Large ZTR Module | Remove Standard ZTR | 617
11200 | Large ZTR Module | Fit The Two Inserts to the XY Platform | 618
11200 | Large ZTR Module | Fit The Large Z Module | 619
11200 | Large ZTR Module | Remove the Transit Clamp and Fit the Linear Bearing Blanks | 620
11200 | Large ZTR Module | In Stage Nav Settings, Set Spacer Offset and Stage Thickness to 0 | 621
11200 | Large ZTR Module | Check Function of Touch Alarm | 622
11200 | Large ZTR Module | Initialise Stage | 623
11200 | Large ZTR Module | Check Full Movement of Each Axis on ZTR | 624
10546 | SmartTouch | Chamberscope Fitted to Stage Door | 625
10546 | SmartTouch | SmartTouch Software Version | 626
10546 | SmartTouch | Install SmartTouch Software | 627
10546 | SmartTouch | Install Easy VP Aperture | 628
10546 | SmartTouch | Capture Sample Image Using the Nav Cam | 629
10546 | SmartTouch | Select Sample Holder and Sample Type | 630
10546 | SmartTouch | Set Yellow Line To 8.5mm WD | 631
10546 | SmartTouch | Sample Setup Complete | 632
10546 | SmartTouch | In the Image Tab, Check Sliders Function | 633
10546 | SmartTouch | Save Image - Photomode 1, Scan speed 8 | 634
10546 | SmartTouch | Save Image - Photomode 2, Single Image | 635
10546 | SmartTouch | Save Image - Photomode 3, Scan Speed 2, Multiple Frames | 636
10546 | SmartTouch | Check Auto Saturation (W Filament Only) | 637
10546 | SmartTouch | Check Auto Stigmation | 638
10546 | SmartTouch | Check Manual Stigmation | 639
10546 | SmartTouch | Check Annotate | 640
10546 | SmartTouch | Check All Detectors Are Presaent | 641
10546 | SmartTouch | Select the Automate tab & Create Folder as Factory | 642
10546 | SmartTouch | Setup 3 Point Calibration (Nav Cam ) | 643
10546 | SmartTouch | Automation Run Completed | 644
10546 | SmartTouch | Select Browse Tab and Check Functions | 645
10960 | 3DSM | Install 3DSM Software and Record Software Version | 646
10960 | 3DSM | Set Up Image Using One segment of the BSD | 647
10960 | 3DSM | Run 3DSM and Setup | 648
10960 | 3DSM | Save 3DSM Project | 649
10960 | 3DSM | Store 3DSM Project | 650
10547 | Zen | ZEN Version | 651
10547 | Zen | Correct Version of Windows 10 Installed | 652
10547 | Zen | Install ZEN | 653
10547 | Zen | Install Licence | 654
10547 | Zen | Install Remote API Client | 655
10547 | Zen | Functional Test | 656
10547 | Zen | Save Snapped Image | 657
10547 | Zen | Store Snapped Image | 658
10547 | Zen | Store Entitlement Certificate | 659
10548 | Atlas 5 | Connect Silver Box and Cables | 660
10548 | Atlas 5 | Download latest Version Of Software (see link in docs) | 661
10548 | Atlas 5 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info | 662
10548 | Atlas 5 | Request Licence File from Fibics | 663
10548 | Atlas 5 | Install Atlas Software | 664
10548 | Atlas 5 | Install Licence File | 665
10548 | Atlas 5 | Calibrate SEM | 666
10548 | Atlas 5 | Obtain Mosaic | 667
10548 | Atlas 5 | Set ATLAS Rotation Offset | 668
10548 | Atlas 5 | Save ATLAS Licence Document | 669
10549 | YAG BSD | Install YAG BSD Lightpipe | 670
10549 | YAG BSD | Alignment | 671
10549 | YAG BSD | Select YAG BSD in the Administrator | 672
10549 | YAG BSD | 30KV Function Image | 673
10549 | YAG BSD | Store YAG BSD Functional Image | 674
10549 | YAG BSD | Obtain a Image at 3KV | 675
11201 | Stage Acceptance Check X and Y Axis - large ZTR Module | X Axis | 676
11201 | Stage Acceptance Check X and Y Axis - large ZTR Module | Y Axis | 677
11201 | Stage Acceptance Check X and Y Axis - large ZTR Module | X+Y Axis | 678
11201 | Stage Acceptance Check X and Y Axis - large ZTR Module | Y+X Axis | 679
11201 | Stage Acceptance Check X and Y Axis - large ZTR Module | 10mm Square | 680
11202 | Stage Acceptance Check Z Axis - Large ZTR Module | Z Axis | 681
11203 | Calibrate Compucentric Stage - Large ZTR Module | Calibrate Stage Centre (Right Click on Stgae tab) | 682
11203 | Calibrate Compucentric Stage - Large ZTR Module | Tilt Centre Calibration (Located in Docking panel) | 683
11204 | Compucentric Stage Test Rotation - Large ZTR Module | Rotate=20 deg | 684
11204 | Compucentric Stage Test Rotation - Large ZTR Module | Rotate=45 deg | 685
11204 | Compucentric Stage Test Rotation - Large ZTR Module | Rotate=90 deg | 686
11204 | Compucentric Stage Test Rotation - Large ZTR Module | Rotate=180 deg | 687
11205 | Compucentric Stage Tests - Tilt - Large ZTR Module | Tilt=5 deg | 688
11205 | Compucentric Stage Tests - Tilt - Large ZTR Module | Tilt=10 deg | 689
11205 | Compucentric Stage Tests - Tilt - Large ZTR Module | Tilt=20 deg | 690
11205 | Compucentric Stage Tests - Tilt - Large ZTR Module | Tilt=45 deg | 691
10550 | Panel Up System | Fit All Panels | 692
10550 | Panel Up System | Fit All Earth Leads | 693
10550 | Panel Up System | Front Column Assy. Fits Correctly | 694
10551 | CIV Macro | Macro Passed | 695
10551 | CIV Macro | Gun Monitor Saved to Product Folder | 696
10551 | CIV Macro | Store CIV Gun Monitor Files (Parram. Mon file & Mon Export File) | 697
207 | Fit Test Tool SCM | Test Tool SCM Fitted | 698
10588 | SE Detector | Collector Volts | 699
10588 | SE Detector | Scintillator | 700
10588 | SE Detector | With the beam blanked, select SE Detector, set contrast and brightness to 50%, scan speed 3 and ensure that there are no white streaks across the screen | 701
10552 | LAB6 Stability Run | Conditions set - 20KV, 1nA, Long Filament Life | 702
10552 | LAB6 Stability Run | Gun Monitor Setup and Running | 703
10552 | LAB6 Stability Run | Stability Run Passed | 704
10552 | LAB6 Stability Run | Gun Monitor Saved to Product Folder | 705
10552 | LAB6 Stability Run | Store STAB Gun Monitor Files (Parram. Mon file & Mon Export File) | 706
208 | Remove Test Tool SCM | Remove the Test Tool SCM | 707
208 | Remove Test Tool SCM | Re-Connect Specimen Touch Cable to X13 on the Vac PCB | 708
206 | Re-fit Dummy SCM | Dummy SCM Fitted and Aligned Correctly | 709
206 | Re-fit Dummy SCM | Specimen Touch Cable Connected to Dummy SCM | 710
10553 | Standard W Resoltion Checks | 30KV Res- ≤2.2nm | 711
10553 | Standard W Resoltion Checks | 3KV Res- ≤8nm | 712
10553 | Standard W Resoltion Checks | 3KV Res- ≤7.5nm - If Unable to Achieve Contact Man Eng. | 713
10553 | Standard W Resoltion Checks | 1KV Res- ≤15nm | 714
10553 | Standard W Resoltion Checks | 1KV Res- ≤14nm - If Unable To Achieve Contact Man Eng. | 715
10553 | Standard W Resoltion Checks | 200V Res- Not measured | 716
10553 | Standard W Resoltion Checks | Store Standard Resolution Images | 717
11097 | SCS Completed | Check SCS Folder/Sales Order | 718
10554 | 2KV W - VP & EasyVP Aperture Resolution Check | 2KV Res - 100um VP Aperture- ≤16nm | 719
10554 | 2KV W - VP & EasyVP Aperture Resolution Check | 2KV Res - 400um Easy VP Aperture- ≤16nm | 720
10554 | 2KV W - VP & EasyVP Aperture Resolution Check | Store VP & Easy VP Aperture Resolution Images | 721
10611 | Standard LaB6 Resolution Checks | 30KV Res- ≤2nm | 722
10611 | Standard LaB6 Resolution Checks | 3KV Res- ≤6nm | 723
10611 | Standard LaB6 Resolution Checks | 1KV Res- ≤9nm | 724
10611 | Standard LaB6 Resolution Checks | 200V Res- Not Measured | 725
10611 | Standard LaB6 Resolution Checks | Store Standard LaB6 Resolution | 726
10612 | 2KV LaB6 - VP Aperture (100um) Resolution Check | 2KV Res - 100um VP Apertutre- ≤12nm | 727
10612 | 2KV LaB6 - VP Aperture (100um) Resolution Check | Store VP Aperture (100u) Resolution Images | 728
10881 | 2KV LAB6 EasyVP Aperture (750um) Resolution Check | 2KV Res - 750um Easy VP Aperture- ≤12nm | 729
10881 | 2KV LAB6 EasyVP Aperture (750um) Resolution Check | Store Easy VP Aperture (750u) Resolution Images | 730
10614 | 2KV Lab6 - Easy Align Aperture (750um) Resolution Check | 2KV Res - 750um Easy Align Aperture- ≤12nm | 731
10614 | 2KV Lab6 - Easy Align Aperture (750um) Resolution Check | Store Easy Align Aperture (750u) Resolution Images | 732
11206 | 30KV Resolution - Large ZTR Module | 30KV Res- ≤2nm | 733
11206 | 30KV Resolution - Large ZTR Module | Store 30KV ( Large ZTR)Resolution Image | 734
10875 | 2KV W Easy Align Aperture (750um) Resolution Check | 2KV Res - 750um Easy Align Aperture- ≤16nm | 735
10875 | 2KV W Easy Align Aperture (750um) Resolution Check | Store Easy Align Aperture (750u) Resolution Images | 736
10658 | 2KV LaB6 - VP Aperture Resolution Checks | 2KV - 100um VP Aperture- ≤12nm | 737
10658 | 2KV LaB6 - VP Aperture Resolution Checks | 2KV - 140um VP Aperture- ≤12nm | 738
10658 | 2KV LaB6 - VP Aperture Resolution Checks | Store 2KV Aperture Resolution Images | 739
10555 | BSD Resolution Check (W) | BSD Resolution ≤4nm. | 740
10555 | BSD Resolution Check (W) | Store BSD Resolution Image | 741
230 | BSD Resolution Check (EVO18) | BSD Resolution ≤4nm. | 742
230 | BSD Resolution Check (EVO18) | Store BSD Resolution Image | 743
203 | BSD Resolution Check (LaB6) | BSD Resolution 3.8nm | 744
203 | BSD Resolution Check (LaB6) | Store BSD Resolution Image | 745
10556 | VPSE Resolution (W) Check - Easy VP Aperture | 30KV Resolution - Easy VP- ≤3.4nm | 746
10556 | VPSE Resolution (W) Check - Easy VP Aperture | Store VPSE Resolution Image | 747
204 | VPSE Resolution (LaB6) Check - Easy VP Aperture | VPSE Resolution 3.0nm | 748
204 | VPSE Resolution (LaB6) Check - Easy VP Aperture | Store VPSE Resolution Image | 749
10659 | VPSE Resolution Check (EVO18) - VP Aperture | 30KV Resolution- ≤5.5nm | 750
10659 | VPSE Resolution Check (EVO18) - VP Aperture | Store VPSE Resolution Image | 751
10557 | Beam Deceleration Resolution | Resolution @ 7KV- ≤9nm | 752
10557 | Beam Deceleration Resolution | Store Beam Deceleration Resolution Image | 753
10769 | Run Up SmartSEM | ASAP FW/HW Present | 754
10769 | Run Up SmartSEM | Stage FW/HW Present | 755
10769 | Run Up SmartSEM | Vacuum FW/HW Present | 756
10769 | Run Up SmartSEM | EO PCB FW/HW Present | 757
10769 | Run Up SmartSEM | Vacuum Reading OK | 758
10769 | Run Up SmartSEM | Initialise Stage | 759
10770 | Run SmartSEM Service Centre | PSU Tests Passed | 760
10770 | Run SmartSEM Service Centre | EO Tests Passed | 761
10776 | System Cabling | All Cables Are Neatly Routed | 762
10776 | System Cabling | All Plugs Are Secured Correctly | 763
10776 | System Cabling | All Earth Cables Are Tightly Secured | 764
10776 | System Cabling | Check Rear Panel Cables Are Correctly Located | 765
10776 | System Cabling | Check Opto Devices Are Not Damaged On All PCBs & Mains Box | 766
10561 | Final X-Ray Check | Complete & Sign Test Certificate | 767
10565 | Copy Standard Rersolution Images | Copy the standard resolution images into the C:Support/Data/Images folder. | 768
10566 | Attach Stickers | Safety Labels | 769
10566 | Attach Stickers | Language | 770
10566 | Attach Stickers | China RoHS Sticker | 771
971 | Fit Front Panel Logo | Clean Front Panel | 772
971 | Fit Front Panel Logo | Prepare Logo | 773
971 | Fit Front Panel Logo | Fitting Logo | 774
10567 | Install Customer Licences | Delete the Production Test Licences | 775
10567 | Install Customer Licences | Obtain the customer Entitlement Certificate | 776
10567 | Install Customer Licences | Download into the C:Support/Data/Licence folder | 777
10567 | Install Customer Licences | Install the Customer Licences | 778
10567 | Install Customer Licences | Copy Contents of Image Folder to the Product Folder | 779
10567 | Install Customer Licences | Copy EMServer.db to the Product Folder | 780
10567 | Install Customer Licences | Copy EMServer.db to the C: Support/Data/Log Files folder | 781
10567 | Install Customer Licences | Print out the Customer Entitlement Certicate - To be sent to the customer | 782
10567 | Install Customer Licences | Store EM Server.db | 783
10567 | Install Customer Licences | Copy EMServer.log to the Product Folder | 784
10567 | Install Customer Licences | Store EM Server Log | 785
10567 | Install Customer Licences | Copy EMServer Screening.txt to the Product Folder | 786
10567 | Install Customer Licences | Store EM Server Screening | 787
10568 | Add SmartEDX Calibration File | Calibration File Added to C: Support/Data/Log Files | 788
10568 | Add SmartEDX Calibration File | Copy Calibration File to the Product Folder | 789
10568 | Add SmartEDX Calibration File | Store SmartEDX Calibration File | 790
10569 | LAB6 Settings | Check thet the following coditions are set= I Beam=20uA, Fil I=1.60A, Long Fil. Life selected | 791
10569 | LAB6 Settings | Secure Gun Head with 3 M6x18(000000-0004-529) screws | 792
10570 | Run SmartSEM as Guest User | Set Language in User Preferences | 793
10571 | Run SmartSEM as System User | Set Language in User Preferences | 794
10572 | Select Keyboard & System Destination | Select Keyboard Language | 795
10572 | Select Keyboard & System Destination | Set Keyboard Language in Windows Control Panel | 796
10572 | Select Keyboard & System Destination | Destination of System | 797
10574 | Ensure Easy VP/ Easy Align Aperture is Fitted | Easy VP/Easy Align Aperture Fitted | 798
1124 | Sign Quality Certificate | Quality Certificate Signed By: | 799
10576 | Final Quality Checks | QA Engineer | 800
10577 | Check System Configuration | Configuration Matches the Sales Order | 801
10577 | Check System Configuration | Sales Order is the Latest Version | 802
1289 | Check System Serial Number Label | Correct Serial Number label fitted and correct. | 803
10579 | Standby to Operate | Rear Plinth Fan Running in Standby (yellow button) | 804
10579 | Standby to Operate | Rear Plinth Fan Running in Operate (green button) | 805
10579 | Standby to Operate | EO Fans Running in Operate (green button) | 806
10579 | Standby to Operate | PC Runs Up | 807
10580 | Check SmartSEM Version | SmartSEM Version | 808
10580 | Check SmartSEM Version | Check SmartSEM is in the folder C:Support/Data/SmartSEM Software | 809
10587 | Stage Functions | Initialise Stage | 810
10587 | Stage Functions | Stage Movements operate Correctly | 811
10589 | Aperture Changer | Apertures in the Correct Position | 812
10589 | Aperture Changer | Both apertures can be zero wobbled | 813
10590 | VPSE Detector | VPSE Is Selected When Going Into VP Mode | 814
10590 | VPSE Detector | Image Brightness Changes When Changing Bias | 815
10590 | VPSE Detector | Breakdown Of Bias Occurs at 100% | 816
10585 | C2D Detector | C2D Is Selected When Going Into VP Mode | 817
10585 | C2D Detector | Image Brightness Changes When Changing Bias | 818
10585 | C2D Detector | Breakdown Of Bias Occurs at 100% | 819
996 | Confirm 4-eye check by additional operator | Checked Carried Our By: | 820
1008 | Confirm 4-eye check by additional operator | Checked Carried Our By: | 821
10777 | Detectors & Accessories | Fitted To Correct ports | 822
10777 | Detectors & Accessories | Connected to Correct ASAP Ports & Labeled | 823
10591 | C2DX Detector | C2DX Is Selected When Going Into EP Mode | 824
10591 | C2DX Detector | Image Brightness Changes When Changing Bias | 825
10591 | C2DX Detector | Breakdown Of Bias Occurs at 100% | 826
10592 | BSD/BSD1 Detector | Check Diode Segments | 827
10592 | BSD/BSD1 Detector | Check Gain Ranges | 828
10592 | BSD/BSD1 Detector | Check Auto Range | 829
10592 | BSD/BSD1 Detector | Check TOPO | 830
10741 | Extended Pressure (EP) | EP/Water Kit Fitted | 831
10741 | Extended Pressure (EP) | MKS Driver Box Link Set to 0-10V | 832
10741 | Extended Pressure (EP) | TTL Kit Fitted | 833
10737 | Fit Lens Mounted BSD Diode Holder | Diode Holder Assembly aligned using jig. | 834
10737 | Fit Lens Mounted BSD Diode Holder | All screw fully tightened | 835
10737 | Fit Lens Mounted BSD Diode Holder | Check alignment is Ok again using jig | 836
10742 | Label Up Valve Cables | Check all valve cables have been labelled | 837
10742 | Label Up Valve Cables | All cable earth screens are in electrical contact with the grounding bracket | 838
11193 | Fit SE Detector | Serial Number (or NA) | 839
11193 | Fit SE Detector | Check that the Assembly is the Correct Part Number | 840
11193 | Fit SE Detector | Correct Collector Beak Fitted | 841
11193 | Fit SE Detector | Fit Detector Assembly | 842
11193 | Fit SE Detector | Apply Emersion Oil to PM Tube Face | 843
11193 | Fit SE Detector | Fit PM Head Amp Assembly | 844
10778 | Air System | Air Pipes Are Correctly Routed and Kink Free | 845
10599 | Control Panel | Check Keyboard Language | 846
10599 | Control Panel | Check the operation of the control panel and the Keyboard | 847
10884 | SmartEDX | Check Aquire Spectrum | 848
10884 | SmartEDX | Check Line Scanning | 849
10884 | SmartEDX | Check Mapping | 850
10884 | SmartEDX | Check EDX Icon Is Present In All User Toolbars | 851
10884 | SmartEDX | EDAX Calibration File Saved in the folder C:Support/Data/Log Files | 852
10887 | EMO Box | Power Off Button Tested | 853
10888 | Camera Stand | Software Loaded | 854
10889 | Additional Hard Disk & Cloning Software | Hard Disk Is Seen In File Manager | 855
10889 | Additional Hard Disk & Cloning Software | Cloning Software Installed | 856
1050 | Micro Switch Checks | Check Set Point X Axis Micro Switch | 857
1050 | Micro Switch Checks | Check Set Point Y Axis Micro Switch | 858
1050 | Micro Switch Checks | Check Set Point Z Axis Micro Switch | 859
1050 | Micro Switch Checks | Check Set point T Axis (-10.8º) | 860
1051 | Stage Initialisation | Check Stage Initialises Correctly | 861
1052 | EVO 25 Travel | Confirm X Travel (0 - 130mm) | 862
1052 | EVO 25 Travel | Confirm Y Travel (0 - 130mm) | 863
1052 | EVO 25 Travel | Confirm Z Travel (0 - 50mm) | 864
1052 | EVO 25 Travel | Check Tilt Travel -10° to +90° | 865
1054 | EVO 10 Travel | Confirm X Travel (0 - 80mm) | 866
1054 | EVO 10 Travel | Confirm Y Travel (0 - 80mm) | 867
1054 | EVO 10 Travel | Confirm Z Travel (0 - 35mm) | 868
1054 | EVO 10 Travel | Check Tilt Travel -10° to +90° | 869
1056 | Stage Axis Cycle Tests | Run Axis Cycle Test (min - 1Hr) | 870
1056 | Stage Axis Cycle Tests | Stage Axis Cycle Test with 0.5Kg Weight for 10 Mins | 871
1058 | EV025 Stage Kit | Stage Kit | 872
1058 | EV025 Stage Kit | 27mm Spacer Block | 873
1058 | EV025 Stage Kit | 54mm Spacer Block | 874
1058 | EV025 Stage Kit | 81mm Spacer Block - Fitted | 875
1059 | EVO15/ Sigma Stage Kit | Stage Kit | 876
1059 | EVO15/ Sigma Stage Kit | 27mm Spacer Block Fitted | 877
11115 | SCS Check | Check SCS Folder/Sales Order | 878
11115 | SCS Check | SCS Number (If applicable) | 879
760 | Material Availability | Are there any shortages? | 880
10780 | Interlocks | Gun Switch Cable is Pluuged into VICE PCB | 881
10897 | Probe Current Measurement | PCM and SCM read the same @ 500pA | 882
1053 | EVO 15 Travel | Confirm X Travel (0 - 125mm) | 883
1053 | EVO 15 Travel | Confirm Y Travel (0 - 125mm) | 884
1053 | EVO 15 Travel | Confirm Z Travel (0 - 50mm) | 885
1053 | EVO 15 Travel | Check Tilt Travel -10° to +90° | 886
1060 | EVO10 Stage Kit | Stage Kit | 887
1062 | Stage Type & Serial Number | Select Stage Type | 888
1062 | Stage Type & Serial Number | Stage Serial Number: | 889
1063 | Quality sign off. | Quality Sign Off Carried Out By: | 890
11106 | Record Plinth S/No | Serial Number | 891
10683 | Inspect Damper Plate | Damper Plate is Correct Version | 892
10683 | Inspect Damper Plate | Damper Plate is Located Correctly | 893
10684 | Inspect Table Top | Table Top is Correct Version | 894
10684 | Inspect Table Top | Table Top is Free from Damage | 895
10684 | Inspect Table Top | Remove Punchout (Dependant on system type) | 896
10685 | Fit Damper Plate Metalwork | Earth Plate | 897
10685 | Fit Damper Plate Metalwork | Cable Plates | 898
10685 | Fit Damper Plate Metalwork | Bellows | 899
10685 | Fit Damper Plate Metalwork | Bellows is concentric to hole in the desktop | 900
10686 | Fit Turbo Controller to Mains Box | Turbo Controller Serial Number | 901
11399 | Installation of Quiet Mode PCB | Fit Quiet Mode PCB | 902
11399 | Installation of Quiet Mode PCB | Fit Cable X45 | 903
11399 | Installation of Quiet Mode PCB | Fit Cable X11 | 904
10898 | Beam Blanking | Blanking Unit Installed | 905
10898 | Beam Blanking | Blanking Cable Connected To Column Connection | 906
10899 | Video Input Kit | Video Input Kit Installed | 907
10900 | Nav Cam | Stage Moves To Correct Position When Arm Is Opened | 908
10900 | Nav Cam | Image Is Captured | 909
10900 | Nav Cam | Functional Check | 910
10904 | Mineralogic | Mineralogic Software Installed and Tested | 911
10903 | SmartPI | SmartPI Software Installed and Tested | 912
10905 | SmartTouch | Run Up & Obtain An Image | 913
1181 | SmartStitch | Check SmartStitch can be run | 914
1181 | SmartStitch | SmartStitch folder is present on the Desktop | 915
10906 | ZEN | ZEN Software Installed | 916
10907 | ATLAS 5 | Atlas Hardware Installed | 917
10907 | ATLAS 5 | Atlas Software Installed | 918
10907 | ATLAS 5 | Atlas Licence Saved to the C:Support/Data/Log Files | 919
10602 | Magnification Calibration | Check Calibration in each mag range | 920
10602 | Magnification Calibration | Check Stage Angle has been setup correctly | 921
10603 | 30KV Image Check | X Stigmator | 922
10603 | 30KV Image Check | Y Stigmator | 923
10603 | 30KV Image Check | X Gun Tilt ≤50% | 924
10603 | 30KV Image Check | Y Gun Tilt ≤50% | 925
10603 | 30KV Image Check | X Gun Shift ≤50% | 926
10603 | 30KV Image Check | Y Gun Shift ≤50% | 927
11269 | Check Column Align | Check 30KV to 5KV image Shift is ≤20µm | 928
11397 | C3 Cal Check | Change in Working Distance ≤0.2mm | 929
11394 | 1KV Image Check | X Stigmator | 930
11394 | 1KV Image Check | Y Stigmator | 931
11394 | 1KV Image Check | X Gun Tilt ≤50% | 932
11394 | 1KV Image Check | Y Gun Tilt ≤50% | 933
11394 | 1KV Image Check | X Gun Shift ≤50% | 934
11394 | 1KV Image Check | Y Gun Shift ≤50% | 935
1171 | Beamsleeve | Nano VP funtions | 936
1171 | Beamsleeve | Collector Bias functions | 937
669 | Support Folder - Check Customer Entitlement Certificate and EM Server.db | The Customer Entitlement Certificate is present in C:Support/Data/Licence folder | 938
669 | Support Folder - Check Customer Entitlement Certificate and EM Server.db | The EM Server.db is present in C:Support/Data/Log Files folder | 939
1179 | Run Calibration Extraction Untility | Download the CEU folder and copy to the Desktop. | 940
1179 | Run Calibration Extraction Untility | Install RClientinstaller.exe (if not already installed) | 941
1179 | Run Calibration Extraction Untility | Run SEM_Calibration Extraction Tool. | 942
1179 | Run Calibration Extraction Untility | Store the SEM_CEU.txt and EVO**-**-**.csv files | 943
1179 | Run Calibration Extraction Untility | Copy the SEM_CEU.txt and EVO**-**-**.csv files to the C:Support/Data/Log Files folder | 944
1179 | Run Calibration Extraction Untility | If the RClientinstaller has been installed for the the CEU only, then uninstall | 945
1277 | Copy the Support folder to the (D:) Drive | Copy the Support folder and its sub folders to the (D:) drive on the PC | 946
1277 | Copy the Support folder to the (D:) Drive | Verify that all the folders and their contents have been copied | 947
1290 | Paint work and Cladding | Ensure paint finish matches and is free from chips and scratches | 948
1290 | Paint work and Cladding | Cladding is secured correctly and not able to move. | 949
1290 | Paint work and Cladding | Cladding Front Hood Aasembly Fits Correctly. | 950
1290 | Paint work and Cladding | Safety stickers are applied correctly and are the correct language. | 951
10607 | Internal Plinth Check | Check Cables | 952
10607 | Internal Plinth Check | Check Electrical Connectors | 953
10607 | Internal Plinth Check | Check Vacuum Connections | 954
10608 | Final System Vacuum Reading | System Vacuum Reading (≤3.00e-06 mbar) | 955
10759 | Final Gun Vacuum Reading | Gun Vacuum Reading (≤5.00e-07 mbar) | 956
11197 | Remove Items From Quick Access | Remove Files | 957
11197 | Remove Items From Quick Access | Remove Folders | 958
10638 | SCS 4 Eye Check | Special Fitted | 959
10638 | SCS 4 Eye Check | If Fitted, Checked by SCS Team Member | 960
11063 | Clear Download Folder & Internet History | Download Folder Images Deleted | 961
11063 | Clear Download Folder & Internet History | Clear Internet History | 962
893 | Remove MAC Address | Remove MAC address (Internet) | 963
893 | Remove MAC Address | Remove MAC address (MIC-NET) | 964
11420 | Earth Continuity Test | PAT Tester Serial Number | 965
11420 | Earth Continuity Test | PAT Tester Cal Expiry Date | 966
11420 | Earth Continuity Test | Earth resistance – measured value must be < 0.1 Ω | 967
11420 | Earth Continuity Test | Insulation resistance – measured value must be > 1.0M Ω | 968
11420 | Earth Continuity Test | Test Certificate Completed & Signed | 969
213 | Photograph Oxford Sales Number Labels | Take Photo of Label | 970
213 | Photograph Oxford Sales Number Labels | Oxford Sales Order | 971
213 | Photograph Oxford Sales Number Labels | Store Oxford Sales Order Label Images | 972
10618 | Copy Radiation Test Certificate, Customer Entitlement Certificate and Packing Lists | Radiation Cert | 973
10618 | Copy Radiation Test Certificate, Customer Entitlement Certificate and Packing Lists | Customer Entitlement Certificate | 974
10618 | Copy Radiation Test Certificate, Customer Entitlement Certificate and Packing Lists | Itemes removed | 975
10618 | Copy Radiation Test Certificate, Customer Entitlement Certificate and Packing Lists | Items retained | 976
10618 | Copy Radiation Test Certificate, Customer Entitlement Certificate and Packing Lists | System and accessories packing list | 977
10619 | Language of EVO User Manual | User manual language | 978
10620 | Fit Damper Plate Transit Brackets | Fit LH Bracket | 979
10620 | Fit Damper Plate Transit Brackets | Fit RH Bracket | 980
10620 | Fit Damper Plate Transit Brackets | Fit Rear Bracket | 981
10620 | Fit Damper Plate Transit Brackets | Fit Front Bracket | 982
10621 | Check Damper Plate Transit Brackets | Check LH Bracket | 983
10621 | Check Damper Plate Transit Brackets | Check RH Bracket | 984
10621 | Check Damper Plate Transit Brackets | Check Rear Bracket | 985
10621 | Check Damper Plate Transit Brackets | Check Front Bracket | 986
10621 | Check Damper Plate Transit Brackets | Checked By (Engineer's Name) | 987
1288 | Fit Upper Cladidng Foam Packing Pieces | Fit Right Hand Foam | 988
1288 | Fit Upper Cladidng Foam Packing Pieces | Fit Left Hand Foam | 989
11421 | Fit PSU Transit Spacers | Fit EO/Stage PSU Spacers | 990
11421 | Fit PSU Transit Spacers | Fit Vac/ASAP PSU Spacers | 991
193 | Check 2 Off Lindy Mini DP to DP cables Supplied - | Lindy Cables Supplied | 992
11322 | Main System Crate | Weight | 993
11322 | Main System Crate | Dimensions Length | 994
11322 | Main System Crate | Dimensions Width | 995
11322 | Main System Crate | Dimensions Height | 996
11322 | Main System Crate | Security Seals Serial No | 997
11322 | Main System Crate | Shock Watch Serial Number | 998
11322 | Main System Crate | Tilt Watch Serial No. | 999
11322 | Main System Crate | Packed By | 1000
1133 | Print Shipment Labels and Identify System Material | Print Shipment Labels | 1001
1133 | Print Shipment Labels and Identify System Material | Label Material Boxes | 1002
1133 | Print Shipment Labels and Identify System Material | Place left over stickes in Document Holder | 1003
952 | Declaration of conformity (Gem 360) | Please print 'Declaration of Conformity' from URL link | 1004
953 | Declaration of conformity (Gem 560) | Please print 'Declaration of Conformity' from URL link | 1005
954 | Declaration of conformity (Gem 460) | Please print 'Declaration of Conformity' from URL link | 1006
11255 | SCS Check | Check SCS Folder/Sales Order | 1007
11255 | SCS Check | SCS Number (If applicable) | 1008
425 | PC MAC Address | Record MAC address (Internet not MIC-NET) | 1009
10930 | FIB set up and Disabled | Run the ‘Disable FIB–all GIS.reg’ file | 1010
11021 | Configure OEM PC | PC Serial Number: | 1011
11021 | Configure OEM PC | Connect Up OEM Test Tool PC | 1012
11021 | Configure OEM PC | Unistall SmartSEM | 1013
11021 | Configure OEM PC | Install SmartSEM Version | 1014
11021 | Configure OEM PC | Install Latest Service Pack | 1015
11021 | Configure OEM PC | Install Gemini Support Tools-Version | 1016
11021 | Configure OEM PC | Run ASAP Screening Program | 1017
11021 | Configure OEM PC | System Specific Datazone Loaded | 1018
252 | Configure PC ASAP | PC S/N | 1019
252 | Configure PC ASAP | PC/Model (SAP No.) | 1020
252 | Configure PC ASAP | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448 | 1021
252 | Configure PC ASAP | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448 , then the PC needs reimaging – follow the attached WI. The correct image version is V2. | 1022
252 | Configure PC ASAP | Hard disk Image (Version) | 1023
252 | Configure PC ASAP | Record Operating System | 1024
252 | Configure PC ASAP | Check the Entitlement certificate to determine which version of Smartsem to install | 1025
252 | Configure PC ASAP | Install Smart SEM (See WI 9693) Record Version | 1026
252 | Configure PC ASAP | Check for available SmartSEMService Packs and install them. (N/A if none) | 1027
252 | Configure PC ASAP | Check for available SmartSEM hotfixes and install them. (N/A if none) | 1028
252 | Configure PC ASAP | Install CZ Support Tools. Ver | 1029
252 | Configure PC ASAP | Disable Automatic Updates | 1030
252 | Configure PC ASAP | Run ASAP Screening Program | 1031
252 | Configure PC ASAP | System Specific Datazone Loaded | 1032
252 | Configure PC ASAP | In SEM status set 'Mains Sync' to off | 1033
11038 | Configure PC SVEM | PC S/N | 1034
11038 | Configure PC SVEM | If SEM Enhanced PC(349535-9001-020) has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2. | 1035
11038 | Configure PC SVEM | PC/Model (SAP No.) | 1036
11038 | Configure PC SVEM | Hard disk Image (Version) | 1037
11038 | Configure PC SVEM | Record Operating System | 1038
11038 | Configure PC SVEM | Install Firewire Driver | 1039
11038 | Configure PC SVEM | Check the Entitlement certificate to determine which version of Smartsem to install | 1040
11038 | Configure PC SVEM | Install Smart SEM. Record Version | 1041
11038 | Configure PC SVEM | Check for available SmartSEMService Packs and install them. (N/A if none) | 1042
11038 | Configure PC SVEM | Check for available SmartSEM hotfixes and install them. (N/A if none) | 1043
11038 | Configure PC SVEM | Install CZ Support Tools. Ver | 1044
11038 | Configure PC SVEM | Disable Automatic Updates | 1045
11038 | Configure PC SVEM | Approved Datazone Loaded | 1046
254 | Install Customer Language Option | Language | 1047
972 | Install Service centre Software | Service centre S/W installed on Customer PC. | 1048
972 | Install Service centre Software | Confirm Data Logging is running and External Logging is enabled. | 1049
257 | Start SmartSEM Administrator (Login as Factory) | In Factory Tab: Omit Dongle Check | 1050
257 | Start SmartSEM Administrator (Login as Factory) | In Factory Tab: Enter S/N | 1051
257 | Start SmartSEM Administrator (Login as Factory) | In License Tab: Set All Licenses | 1052
11270 | Install Predictive Service Software | Predictive Service S/W installed on Customer PC. | 1053
11270 | Install Predictive Service Software | Predictive Service has been activated | 1054
255 | Connect Plinth Services | Air (Pressure = 6 Bar) | 1055
255 | Connect Plinth Services | Water | 1056
255 | Connect Plinth Services | Confirm that the Water Flow Switch is set to 0.9 LPM | 1057
255 | Connect Plinth Services | Mains | 1058
255 | Connect Plinth Services | Safety Earth | 1059
255 | Connect Plinth Services | N2 | 1060
255 | Connect Plinth Services | Ensure Test tool Quiet Mode is set to 'ON' in service cupboard (Unless customer Quiet mode fitted) | 1061
11427 | Connect Plinth Services (Raith) | Air (Pressure = 6 Bar) | 1062
11427 | Connect Plinth Services (Raith) | Water | 1063
11427 | Connect Plinth Services (Raith) | Check Water Flow switch set to 0.9 LPM | 1064
11427 | Connect Plinth Services (Raith) | Mains | 1065
11427 | Connect Plinth Services (Raith) | Safety Earth | 1066
11427 | Connect Plinth Services (Raith) | N2 | 1067
11427 | Connect Plinth Services (Raith) | Remove Blank and fit Test Tool Stage | 1068
11427 | Connect Plinth Services (Raith) | Fit Test Tool Stage Controller and cable up | 1069
11427 | Connect Plinth Services (Raith) | Mount Test Tool SE2 Detector onto chamber | 1070
11427 | Connect Plinth Services (Raith) | Mount PM Tube (with Immersion oil) onto SE2 Detector | 1071
11427 | Connect Plinth Services (Raith) | Mount SCM to port on Stage Door | 1072
424 | Install Quiet Mode | Install customer quiet mode | 1073
424 | Install Quiet Mode | Set Test tool Quiet Mode to 'OFF' (In service cupboard) | 1074
11025 | Fit Ferrites | Fit 2 Off Ferrites to Liner Cable | 1075
11321 | Raith Firmware Check | Correct Firmware installed | 1076
256 | Run System checker and upgrade board FW to latest versions. | Ensure the Stage Controller version is compatible with the Stage Axes version (See compatible verions in drop down list) | 1077
256 | Run System checker and upgrade board FW to latest versions. | Store System checker full report | 1078
1268 | Vacuum System Configuration Check | Vacuum System is Configured correctly | 1079
389 | SmartSEM Administrator (Column / Stage Tab) | Tick option 'Low KV Objective Cap' | 1080
258 | Start SmartSEM User Interface | Confirm there are no Errors in EMserver (Apart from Dongle check and Stage Init) | 1081
187 | Enable Z Move on vent | Z Move on Vent Enabled | 1082
385 | SmartSEM Adminstrator (Factory Tab) | Select correct System type (HV / VP) | 1083
385 | SmartSEM Adminstrator (Factory Tab) | If Pressure Stage is fitted Tick Box In Administrator | 1084
259 | Set up (Standard rear) Camera | Check camera Focus and Rotation | 1085
259 | Set up (Standard rear) Camera | Eliminate camera image tilt | 1086
259 | Set up (Standard rear) Camera | Ty=wrap and Label Cables | 1087
387 | Check for sales order revisions | Order revision in use or 'n/a' | 1088
11026 | Plasma Clean Chamber | Carry out a 3hrs Plasma clean (Remove all samples from chamber!) | 1089
11028 | On/Off Switch Panel | Curved Switch Panel Fitted | 1090
895 | Water Flow checks | Set Water Flow Meter to 0.9 lpm | 1091
260 | BUILD / CONFIG SIGN OFF | System Configuration correct? | 1092
260 | BUILD / CONFIG SIGN OFF | Build quality satisfactory? | 1093
11029 | Prepare Open X Configuration | Clear EPROM on VAC Extension | 1094
11030 | Configure Open X Interfaces X21 & X121 | Configue Raith Open X Interface | 1095
11031 | Set Joystick Angle | Set joystick angle to 180 degrees. | 1096
262 | Bakeout | Opening of column manual valve/s required to achieve gun vacuum | 1097
262 | Bakeout | Carry out overnight ‘Full service Bakeout’ and note actual duration. | 1098
262 | Bakeout | Confirm that all Thermocouples are reading. | 1099
262 | Bakeout | Confirm that all Ion pump heater elements and flange heaters are getting hot. | 1100
262 | Bakeout | Gun Vacuum post Bakeout. Ensure column has cooled prior to taking reading. | 1101
262 | Bakeout | After completion of Bakeout Torque the Gun head bolts to 25NM. | 1102
264 | Input Standard Tip Data | Tip No. | 1103
264 | Input Standard Tip Data | Confirm that the Tip S/N on the Denka datasheet is the same as the Tip S/N on the column sticker. | 1104
264 | Input Standard Tip Data | Filled out Normal Nominal Values in Gun Parameters. | 1105
421 | Set Expert Gun Mode | Select "Tools and User Preferences": Set "Expert Gun Mode = Yes" | 1106
963 | Column Mechanical Alignment Check (DC) | Click through Gemini / Crossbeam alignment wizard and confirm mechanical alignment is good. | 1107
963 | Column Mechanical Alignment Check (DC) | Confirm Min- Max current deviation = <10um | 1108
266 | Carry Out High Voltage Gun Conditioning | Test Power Supply Serial Number | 1109
266 | Carry Out High Voltage Gun Conditioning | Use the 35kV External supply for 30mins. | 1110
266 | Carry Out High Voltage Gun Conditioning | Save Gun Moitor Log as GunConditioning_TIP SN | 1111
266 | Carry Out High Voltage Gun Conditioning | Tighten gun cable screws and confirm they are a minimum torque of 1Nm with torque driver. | 1112
10689 | Fit Vacuum PCB | Vacuum PCB Serial Number | 1113
11260 | Install C2D Detector | Detector Switch Box SN | 1114
11260 | Install C2D Detector | Detector Switch Box Qualification Marks Checked | 1115
10743 | Install Beam Decleration | Fit Stage Elecrtical Feedthrough | 1116
10743 | Install Beam Decleration | PCB Fitted | 1117
10743 | Install Beam Decleration | PSU Qualification Marks Checked | 1118
10743 | Install Beam Decleration | PSU Fitted | 1119
10743 | Install Beam Decleration | Cables Fitted | 1120
11267 | Install SmartEDX Detector | Detector Type | 1121
11267 | Install SmartEDX Detector | Detector Serial Number | 1122
267 | Carry out Liner Conditioning | Use Autokathode tool. | 1123
268 | Carry out Gun run up procedure | Carry Out Gun Run UP | 1124
268 | Carry out Gun run up procedure | Save Gun Monitor Log as Gun Run Up_TIP SN | 1125
79 | Carry out Gun run up procedure (Raith) | Confirm that you are aware that the Filament I value used on Raith systems is 20mA less than the Datasheet value. | 1126
79 | Carry out Gun run up procedure (Raith) | Carry Out Gun Run UP | 1127
79 | Carry out Gun run up procedure (Raith) | Save Gun Monitor Log as Gun Run Up_TIP SN | 1128
1138 | Configure 2nd Inlens detector (DJEL) | The 2nd Inlens detector has been configured as an ESB | 1129
272 | Lower chamber onto foam blocks | Chamber lowered onto foam blocks (N/A to XB 550L) | 1130
273 | Set up Pendulum system | Level Frame | 1131
273 | Set up Pendulum system | Check pendulum height (with spacer) | 1132
273 | Set up Pendulum system | Set up ‘float’ height | 1133
273 | Set up Pendulum system | Check inlet / outlet valves | 1134
273 | Set up Pendulum system | Set up docking pins | 1135
269 | ESB / InLens-Duo Conditioning Macro | Run ESB conditioning MACRO prior to switching on EHT. | 1136
11035 | Raith On Axis Alignment Procedure | With Cond. Off Max Aper. X & Y < 30% | 1137
11035 | Raith On Axis Alignment Procedure | Gun Head Clamping Screws Secure | 1138
11035 | Raith On Axis Alignment Procedure | Tighten Gun Head clamping screws and confirm the 3 out of 4 accessible are a minimum torque of 1Nm with torque driver. | 1139
11035 | Raith On Axis Alignment Procedure | Condenser Clamping Screws Secure | 1140
11035 | Raith On Axis Alignment Procedure | Tighten condensor clamping screws and confirm they are a minimum torque of 1Nm with torque driver. | 1141
11381 | Photograph Macro | Run overnight ESB Run Up & Photograph Macro | 1142
11381 | Photograph Macro | Analyse images for drift, breakdown, ESB issues | 1143
426 | Column Mu Metal Shield | Water pipes / connectors clear. | 1144
11410 | Set default Aligning System Offset Angle | Align offset angle set to -110 degrees. | 1145
11401 | Enter Aperture Sizes | 30um aperture entered on central aperture | 1146
216 | Enter Aperture Sizes 'Low kV Lens' | 20um aperture entered on central aperture | 1147
427 | Column Alignment | Confirm that the contents of Tablefiller are deleted before starting the procedure. Stored values can affect the alignment routine. | 1148
427 | Column Alignment | Mechanically align the column using the Merlin Alignment Wizard | 1149
427 | Column Alignment | Tighten Gun Head clamping screws and confirm the 3 out of 4 accessible are a minimum torque of 1Nm with torque driver. | 1150
427 | Column Alignment | Once adjusters are backed off, tighten C1 + C2 clamping screws and confirm they are a minimum torque of 1Nm with torque driver. | 1151
275 | Align Centre Aperture - 20nA system | Are Gun Head clamping screws secure? | 1152
275 | Align Centre Aperture - 20nA system | Tighten Gun Head clamping screws and confirm the 3 out of 4 accessible are a minimum torque of 1Nm with torque driver. | 1153
275 | Align Centre Aperture - 20nA system | Are Condensor clamping screws secure? | 1154
275 | Align Centre Aperture - 20nA system | Tighten condensor clamping screws and confirm they are a minimum torque of 1Nm with torque driver. | 1155
275 | Align Centre Aperture - 20nA system | Store 1kV emission image (Condensor off) | 1156
275 | Align Centre Aperture - 20nA system | Check emission image against 'emissions distortion criteria' | 1157
275 | Align Centre Aperture - 20nA system | Store 10kV emission image (Condensor off) | 1158
11313 | Align Centre Aperture - 100nA system | Are Gun Head clamping screws secure? | 1159
11313 | Align Centre Aperture - 100nA system | Tighten Gun Head clamping screws and confirm the 3 out of 4 accessible are a minimum torque of 1Nm with torque driver. | 1160
11313 | Align Centre Aperture - 100nA system | Are Condensor clamping screws secure? | 1161
11313 | Align Centre Aperture - 100nA system | Tighten condensor clamping screws and confirm they are a minimum torque of 1Nm with torque driver. | 1162
11313 | Align Centre Aperture - 100nA system | Store 1kV emission image (Condensor off) | 1163
11313 | Align Centre Aperture - 100nA system | Store 10kV emission image (Condensor off) | 1164
292 | Insight Column | Ensure no charging particles outside of specification | 1165
704 | Raith 500V Column Alignment Check | Gun X | 1166
704 | Raith 500V Column Alignment Check | Gun Y | 1167
704 | Raith 500V Column Alignment Check | Aper X | 1168
704 | Raith 500V Column Alignment Check | Aper Y | 1169
219 | Emission Cut Off Check | Normally no cut off is visible @ 10kV, SS9 on the central aperture of the emission image (Condenser off and alignment corrections zeroed). However, some cut off is allowable if certain criteria are met. Please refer to WI for details. | 1170
198 | 10kv Column Alignment Check | Gun X | 1171
198 | 10kv Column Alignment Check | Gun Y | 1172
198 | 10kv Column Alignment Check | Aper X | 1173
198 | 10kv Column Alignment Check | Aper Y | 1174
400 | 10kV On Axis Stig Checks | Stig X | 1175
400 | 10kV On Axis Stig Checks | Stig Y | 1176
197 | 500v Column Alignment Check | Gun X | 1177
197 | 500v Column Alignment Check | Gun Y | 1178
197 | 500v Column Alignment Check | Aper X | 1179
197 | 500v Column Alignment Check | Aper Y | 1180
199 | 99v Column Alignment Check | No cut off on image at minimum Mag. | 1181
199 | 99v Column Alignment Check | Wobble removed | 1182
200 | 99v Alignment check | Were you able to Align column at 99v without column intervention | 1183
11036 | Check EBBU Blanking Position. | Large Aperture in Top Left Quadrant | 1184
11036 | Check EBBU Blanking Position. | Save Emission Image with Condenser On | 1185
11036 | Check EBBU Blanking Position. | Save Emission Image with Condenser Off | 1186
276 | Check Specimen Current @ 10kV WD 5mm (standard objective) | Ext V = | 1187
276 | Check Specimen Current @ 10kV WD 5mm (standard objective) | Measured Speciment Current = | 1188
80 | Check Accuracy of Specimen Current Meter (Raith) | Specimen I measured with SCM | 1189
80 | Check Accuracy of Specimen Current Meter (Raith) | Specimen I measured with Pico-ammeter | 1190
80 | Check Accuracy of Specimen Current Meter (Raith) | Is the SCM measurement within +/-5pa of the Pico-ammeter measurement? | 1191
293 | ESB Concentricity | ESB concentricity in spec | 1192
293 | ESB Concentricity | ESB-Inlens concentricity in spec | 1193
277 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | Ext V = | 1194
277 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | Measured Speciment Current = | 1195
278 | Check Specimen Current @ 10KV 5mmWD for Crossbeam 350 | Ext V = | 1196
278 | Check Specimen Current @ 10KV 5mmWD for Crossbeam 350 | Measured Speciment Current = | 1197
279 | Input High Res Tip Data into Gun Parameters | Fil I | 1198
279 | Input High Res Tip Data into Gun Parameters | Ext V | 1199
428 | C2 Polarity Check | Polarity is correct | 1200
284 | C2 Lens Polarity Check (20nA configuration) | Confirm the 60um aperture emission image still visible in High current mode | 1201
404 | C2 Lens Polarity Check (100nA configuration) | 100nA-Emmision rotates ACW with I increase | 1202
286 | Calibrate Stage Angle | Stage Angle = | 1203
11316 | C3 Lens Polarity Check | No Image rotation with WD change | 1204
11072 | Calibrate Stage Angle | Low mag mode off Angle = | 1205
11072 | Calibrate Stage Angle | Low mag mode on Angle = | 1206
11037 | Calibrate Stage Angle | Stage Angle is close to 133Degs | 1207
429 | Calibrate System Offset Angle (Alignment Wizard) | Angle = | 1208
294 | Image Jumps | Review image jumps run and confirm acceptable | 1209
1618 | BSD 4 Detector | Pneumatics functional | 1210
1618 | BSD 4 Detector | SmartSEM Control tested | 1211
1618 | BSD 4 Detector | Images saved | 1212
10941 | Manually Set Up Column Conditions | 10KV Optimum Mode 0% DOF | 1213
10941 | Manually Set Up Column Conditions | 10KV Analytic Mode 100% DOF 100pA | 1214
10941 | Manually Set Up Column Conditions | 10KV Analytic Mode 100% DOF 5nA | 1215
10941 | Manually Set Up Column Conditions | 10KV Analytic Mode 0% DOF 100pA | 1216
10941 | Manually Set Up Column Conditions | 2KV optimum Mode 0% DOF | 1217
430 | C1/C2/C3 Cal (Alignment Wizard) | C1 Cal = | 1218
430 | C1/C2/C3 Cal (Alignment Wizard) | C2 Cal = | 1219
430 | C1/C2/C3 Cal (Alignment Wizard) | C3 Cal = | 1220
430 | C1/C2/C3 Cal (Alignment Wizard) | Are Calibration factors typical? (see table of typical values in WI) | 1221
406 | Align 60um Aperture in normal and high current mode | Alignment done and confirmed. | 1222
289 | C2/C3 Calibration (Alignment Wizard) | C2 Calibration Value | 1223
289 | C2/C3 Calibration (Alignment Wizard) | C3 Calibration Value | 1224
289 | C2/C3 Calibration (Alignment Wizard) | Are Calibration factors typical? (see table of typical values in WI) | 1225
431 | IProbe and Anode aperture calibration (Alignment Wizard) | Nominal Extractor Voltage = | 1226
431 | IProbe and Anode aperture calibration (Alignment Wizard) | Anode Diameter = | 1227
434 | Check Column Mechanical Alignment | Image Shift (um) = | 1228
435 | C2 Calibration Check | Focus Deviation (um) = | 1229
436 | C3 Calibration Check | Focus Deviation (um) = | 1230
484 | Aperture Finder (X-Beam 350) | Aperture Finder run succesfully | 1231
485 | Stig Checks XB350 | Stig: 30μm, 99V, 10kV, 20kV is ≤ 20% | 1232
485 | Stig Checks XB350 | Stig: Off Axis Aper, 10kV is ≤ 30% | 1233
485 | Stig Checks XB350 | Aper/Gun 30μm, 500V, 10kV, 20kV is ≤ 60% | 1234
306 | Gun value | Actual Gun Axis X = | 1235
306 | Gun value | Actual Gun Axis Y = | 1236
307 | Align1 value | Actual Align 1 X = | 1237
307 | Align1 value | Actual Align 1 Y = | 1238
308 | Stig Value | Actual 1kV X Axis = | 1239
308 | Stig Value | Actual 1kV Y Axis = | 1240
308 | Stig Value | Actual 20kV X Axis = | 1241
308 | Stig Value | Actual 20kV Y Axis = | 1242
482 | Run Tablefiller | DC Tablefiller successfully completed . | 1243
1230 | HP Stage Mapping | Stage mapping completed successfuly | 1244
1224 | Hasap scanning calibration using Zen | carry out scanning calibration (Hasap only) | 1245
296 | Calibration Wizard | Calibration wizard completed. | 1246
297 | Mag Cal Accuracy Check | Check Mag cal errors are <1% in X & Y axes in all 3 Mag Ranges. Orthoganality ± 0.1deg. | 1247
297 | Mag Cal Accuracy Check | Store accuracy images (Mag 1, Mag 2, Mag 3 and Orthog) | 1248
11213 | Widefield Mode Calibration | Calibration wizard completed. | 1249
11071 | Mag Cal Accuracy check (G500/560) | Check Mag cal errors are <2% in X & Y axes in all 3 Mag Ranges. Orthog. ± 0.1deg | 1250
11071 | Mag Cal Accuracy check (G500/560) | Store accuracy images (Mag 1, Mag 2, Mag 3 and Orthog) | 1251
298 | Stage Axis cycle test | Macro (runs 30 mins.) completed. | 1252
300 | Stage Centre Wizard | X = | 1253
300 | Stage Centre Wizard | Y = | 1254
299 | Stage Repeatability Macro | Macro Result 1 | 1255
299 | Stage Repeatability Macro | Macro Result 2 | 1256
299 | Stage Repeatability Macro | Macro Result 3 | 1257
299 | Stage Repeatability Macro | Image of Macro Results | 1258
456 | Eucentric Stage Set up | Eucentricity set up | 1259
456 | Eucentric Stage Set up | Focus Deviation (should be less than 20um) | 1260
302 | Compucentric Height Wizard | Wizard successfully completed. | 1261
303 | Stage Drift Program | Actual (Spec: 6.7nm / min) | 1262
303 | Stage Drift Program | Store stage drift macro results image | 1263
11069 | Dynamic Focus Test | Dynamic focus function check - Whole image in focus | 1264
11069 | Dynamic Focus Test | Y - Axis measurement (400μm +/- 2%) | 1265
11069 | Dynamic Focus Test | Orthogonality | 1266
305 | Sample Set up Holder Height | Holder Height | 1267
305 | Sample Set up Holder Height | Holder Rotation | 1268
305 | Sample Set up Holder Height | Select Customer Sample Holders | 1269
305 | Sample Set up Holder Height | Confirm Holder touch alarm works (Alarm tone and error message) | 1270
906 | R axis Test | Store Images in Image gallery Stage Folder | 1271
906 | R axis Test | Compare photographs to confirm that position of specimen holder has not changed | 1272
907 | XB Tilt Test | XB Tilt Test carried out | 1273
1141 | Concentricity with 2 InLens Detectors | InLens Detector Concentricity is acceptable | 1274
281 | ESB concentricity & hole diameter check (standard objective) | Check concentricity of ESB wrt scanning system. | 1275
281 | ESB concentricity & hole diameter check (standard objective) | Record the ESB serial number - This is the S/N that is shown on the ESB label (1026-XXX or 1125-XXX). | 1276
281 | ESB concentricity & hole diameter check (standard objective) | Measure the ESB hole diameter. | 1277
281 | ESB concentricity & hole diameter check (standard objective) | Store ESB Concentricity image. Add ESB serial number as a text annotation in SmartSEM. (filename must be exactly esbdiameter.tif) | 1278
282 | InLens mechanically aligned to ESB. | Type in measured value (limit: <= 15um) | 1279
282 | InLens mechanically aligned to ESB. | Store Inlens to ESB image | 1280
10647 | Check Calibration Figures in SEM Status are acceptable | Confirm SEM Status that (C1 Double only), C2 & C3 Cal are reasonable | 1281
10647 | Check Calibration Figures in SEM Status are acceptable | System Align Offset Angle Approx: (-21 for X-Beam) (-51 for Sphinx) (-110 for Raith) (N/A for D/C) | 1282
396 | ESB concentricity & hole diameter check (low kV objective) | Check concentricity of ESB wrt scanning system. (Limit: <= 50um) | 1283
396 | ESB concentricity & hole diameter check (low kV objective) | Measure the ESB hole diameter. | 1284
396 | ESB concentricity & hole diameter check (low kV objective) | Record the ESB serial number - This is the S/N that is shown on the ESB label (1026-XXX or 1125-XXX). | 1285
396 | ESB concentricity & hole diameter check (low kV objective) | Store ESB Concentricity image. Add ESB serial number as a text annotation in SmartSEM. (filename must be exactly esbdiameter.tif) | 1286
1549 | Insight | Objective Cap - No particles visible | 1287
1549 | Insight | Lower Aperture - No large particles visible (See WI for details) | 1288
1549 | Insight | Upper Aperture - No cut-off on image and no particles visible. | 1289
283 | Check for ESB charging and Flashing. | Check for ESB charging and Flashing. | 1290
339 | Run 30kV Cycle Conditioning Macro | After conditioning confirm there are no EMServer error messages | 1291
339 | Run 30kV Cycle Conditioning Macro | Store Gun monitor trace as 30kV Cycle Conditioning.mn2 | 1292
311 | Objective Drift test (Macro) | Limit is 600nm .......................Actual = | 1293
311 | Objective Drift test (Macro) | Store objective drift results | 1294
444 | Objective Shift Limit | Actual @ 1kV WD5 = | 1295
444 | Objective Shift Limit | Actual @ 20kV WD2 = | 1296
622 | Objective Shift Limit (ASB) | Actual @ 1kV WD5 = | 1297
622 | Objective Shift Limit (ASB) | Actual @ 20kV WD2 = | 1298
314 | WD Test (1kV – 20kV) | Normal mode (5mm WD) | 1299
314 | WD Test (1kV – 20kV) | High Current mode (5mm WD) | 1300
1269 | Interim QA Checklist | System is ready for QA Engineer to perform Interim QA | 1301
408 | Run Probe I Stability Macro | Actual = | 1302
408 | Run Probe I Stability Macro | After conditioning confirm there are no EMServer error messages | 1303
408 | Run Probe I Stability Macro | Store Gun monitor trace as Probestability.mn2 | 1304
700 | EM Server Screening | Confirm EM Server screeening programme is installed and running | 1305
700 | EM Server Screening | Confirm no repeated EM Server screening errors have been flagged | 1306
290 | Gun Conditions | Gun running to Datasheet | 1307
290 | Gun Conditions | Gun Vacuum better than 5e-9mbar (EHT on/off) | 1308
290 | Gun Conditions | Nominal Ext V acceptable | 1309
291 | Column Alignment | Mechanical Alignment ok | 1310
291 | Column Alignment | Column clamping secure | 1311
409 | Column Mechanical Alignment Check | Gun X | 1312
409 | Column Mechanical Alignment Check | Gun Y | 1313
409 | Column Mechanical Alignment Check | Aper X | 1314
409 | Column Mechanical Alignment Check | Aper Y | 1315
771 | Column Mechanical Alignment Check (Raith) | Gun X | 1316
771 | Column Mechanical Alignment Check (Raith) | Gun Y | 1317
771 | Column Mechanical Alignment Check (Raith) | Aper X | 1318
771 | Column Mechanical Alignment Check (Raith) | Aper Y | 1319
894 | Pendulum Setup | Check pins are centered | 1320
894 | Pendulum Setup | Dampers are concentric | 1321
894 | Pendulum Setup | Rise and sink times are correct | 1322
316 | Initial Resolution Images | Actual for 1kV = | 1323
316 | Initial Resolution Images | Store 1KV Resolution Images and Pate image | 1324
316 | Initial Resolution Images | Actual for 15kV = | 1325
316 | Initial Resolution Images | Store 15KV Resolution Images and Pate image | 1326
11062 | Resolution (Raith) | 1kV = | 1327
11062 | Resolution (Raith) | Store 1KV Resolution Images and Pate image | 1328
11062 | Resolution (Raith) | 15kV = | 1329
11062 | Resolution (Raith) | Store 15KV Resolution Images and Pate image | 1330
317 | Take 200V Image | Image quality good and Stigs in spec. | 1331
318 | Find Image Jumps (Macro) | Jumps - Limit is <= 1 Event / 2 Hours............................................... Actual = | 1332
318 | Find Image Jumps (Macro) | Stigmation change - Limit <= 1% change / 12 Hours....................... Actual = | 1333
319 | Align Apertures (20nA column configuration) | Aperture Finder successfully run | 1334
321 | Probe I minimum check – Analytic mode | Max dev = Measured value +3pa of slider | 1335
322 | Probe I maximum check – Analytic mode | Max dev = Measured value +/-15% of slider (Use Gun Aligns to maximise Probe I) | 1336
323 | Probe I check – 1nA Analytic mode | Max dev= Measured value +/-15% of slider. | 1337
324 | Probe I maximum check – Optimum mode | Max Dev= Measured value +/-15% of demanded current | 1338
325 | Check the Maximum Probe I. | Configuration dependent: | 1339
325 | Check the Maximum Probe I. | Actual Value | 1340
326 | Check the Minimum Probe I | Actual < 7pA +/- 2pA | 1341
328 | Run SC Tablefiller | Tablefiller successfully completed. | 1342
1170 | Run SC Tablefiller (Raith) | Tablefiller successfully completed. | 1343
11110 | SC Manual Tablefiller Entries | 99v | 1344
11110 | SC Manual Tablefiller Entries | 22kv | 1345
11110 | SC Manual Tablefiller Entries | 30kv | 1346
461 | SC Manual Tablefiller Entries (Raith) | 21k - All Apertures at Normal and High current WD5 | 1347
461 | SC Manual Tablefiller Entries (Raith) | 30kv - Aperture 1 and 7 at Normal and High current at WD5 | 1348
11233 | SC Manual Tablefiller VP Mode | 10kv VP | 1349
11234 | SC Manual Tablefiller Widefield Mode | 10kv Widefield Table entries complete for all Apertures? | 1350
11234 | SC Manual Tablefiller Widefield Mode | Beamshift calibration complete for all Apertures.? | 1351
11314 | 100nA Tablefiller | Tablefiller successfully completed. | 1352
329 | Stigmation / Aperture Checks | Stig: 30μm, All EHT’s is ≤ 20% | 1353
329 | Stigmation / Aperture Checks | Stig: Off Axis Aper, All EHT’s is ≤ 30% | 1354
329 | Stigmation / Aperture Checks | Aper/Gun: 30μm, All EHT’s is ≤ 60% | 1355
330 | Stigmation / Aperture Checks | Stig: 20μm Aper, 10kV is ≤ 35% | 1356
330 | Stigmation / Aperture Checks | Stig: Off Axis Aper, 10kV is ≤ 45% | 1357
330 | Stigmation / Aperture Checks | Stig: 20μm Aper, 1kV is ≤ 20% | 1358
330 | Stigmation / Aperture Checks | Stig: Off Axis Aper, 1kV is ≤ 30% | 1359
330 | Stigmation / Aperture Checks | Aper/Gun: 20μm, All EHT’s is ≤ 60% | 1360
705 | Raith Stigmation /Aperture Checks | Stigs: 30um Aperture, All EHT ≤20% | 1361
705 | Raith Stigmation /Aperture Checks | Stigs: Off Axis Apertures, All EHT ≤30% | 1362
705 | Raith Stigmation /Aperture Checks | Aper & Gun: 30um Aperture, All EHT ≤50% | 1363
902 | Fit dowel pins to EBBU mounting plate | Dowel pins have been mounted correctly to the mounting plate | 1364
11040 | Beamshift Calibration | Beamshift Cal. Done On All Entries In Table | 1365
11041 | Blanking Positions | Save An Entry For Each In Table Editor (Four entries) | 1366
11043 | Record Specimen Current | 10KV HC - Typical=434pA | 1367
11043 | Record Specimen Current | 10KV 7um HC - Typical=26pA | 1368
11043 | Record Specimen Current | 10KV 10um HC - Typical=55pA | 1369
11043 | Record Specimen Current | 10KV 15um HC - Typical=118pA | 1370
11043 | Record Specimen Current | 10KV 20um HC - Typical=200pA | 1371
11043 | Record Specimen Current | 10KV 60um HC - Typical=1840pA | 1372
11043 | Record Specimen Current | 10KV 1200um HC - Typical=7195pA | 1373
11043 | Record Specimen Current | 1KV - Typical=222pA | 1374
11043 | Record Specimen Current | 12KV - Typical=300pA | 1375
11043 | Record Specimen Current | 20KV - Typical=418pA | 1376
11043 | Record Specimen Current | 20KV HC - Typical=680pA | 1377
11043 | Record Specimen Current | 20.01KV - Typical=300pA | 1378
11043 | Record Specimen Current | 20.01KV HC - Typical=483pA | 1379
11043 | Record Specimen Current | 30KV - Typical=447pA | 1380
11045 | Touch Alarm | Touch Alarm Function Ok | 1381
1187 | Check the Maximum Probe I. | Configuration dependent: | 1382
1187 | Check the Maximum Probe I. | Actual Value | 1383
333 | Calibrate Max. Extractor Voltage. | 30kV Max Ext Voltage = | 1384
334 | Input Max. Extractor value into Gun Parameters | Fil I = | 1385
334 | Input Max. Extractor value into Gun Parameters | Ext V = | 1386
335 | Run OPTI Probe | Measured Specimen Current @ 20kV = | 1387
335 | Run OPTI Probe | Store Optiprobe log | 1388
10648 | Set EHT | Set EHT to 30KV | 1389
752 | Material Availability | Are there any shortages? | 1390
337 | Automatic Beam Shift and Blanking Test | Blanking Drift Distance Median (Limit <= 300nm) = | 1391
337 | Automatic Beam Shift and Blanking Test | Blanking Drift Time Median (Limit <= 3.5 Sec) = | 1392
337 | Automatic Beam Shift and Blanking Test | Corner Test Max Drift (Limit <= 300nm) = | 1393
337 | Automatic Beam Shift and Blanking Test | Store Autobeamshift results PDF | 1394
11315 | Blanked Probe Current Check | Measure I Probe with beam blanked - Spec <=2pA | 1395
11315 | Blanked Probe Current Check | Difference between EHT on and off < 1pA. | 1396
1657 | Set Probe Current to $[MaxCurrent]nA. | Measured probe current | 1397
1659 | Check Column, Chamber, and Plinth for X-ray Leakage. | Accelerating Voltage=30KV | 1398
1659 | Check Column, Chamber, and Plinth for X-ray Leakage. | Specimen Current | 1399
1659 | Check Column, Chamber, and Plinth for X-ray Leakage. | Monitor Model | 1400
1659 | Check Column, Chamber, and Plinth for X-ray Leakage. | Serial Number | 1401
1659 | Check Column, Chamber, and Plinth for X-ray Leakage. | Cal. Expiry Date | 1402
1659 | Check Column, Chamber, and Plinth for X-ray Leakage. | No X-Ray Leaks Detected | 1403
11120 | Check Column, Chamber, and Plinth for X-ray Leakage. | Accelerating Voltage=30KV | 1404
11120 | Check Column, Chamber, and Plinth for X-ray Leakage. | Monitor Model | 1405
11120 | Check Column, Chamber, and Plinth for X-ray Leakage. | Serial Number | 1406
11120 | Check Column, Chamber, and Plinth for X-ray Leakage. | Cal. Expiry Date | 1407
11120 | Check Column, Chamber, and Plinth for X-ray Leakage. | UHV Chamber No X-Ray Leaks Detected | 1408
11120 | Check Column, Chamber, and Plinth for X-ray Leakage. | Specimen Chamber No X-Ray Leaks Detected | 1409
11120 | Check Column, Chamber, and Plinth for X-ray Leakage. | Condenser Housing No X-Ray Leaks Detected | 1410
11120 | Check Column, Chamber, and Plinth for X-ray Leakage. | VP section No X-Ray Leaks Detected | 1411
10649 | Complete X-Ray Certificate | Complete X-Ray certificate | 1412
1317 | Fit SCM | Fit SCM to Vac PCB | 1413
11244 | X-Ray Test | Serial No. (or N/A) | 1414
11244 | X-Ray Test | No X-Ray Leakage detected. | 1415
1280 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 1416
1280 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 1417
1280 | ZEN Core EM | Install Customer's ZEN Licence | 1418
1280 | ZEN Core EM | Setup SmartSEM Toggle Button | 1419
1280 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 1420
1280 | ZEN Core EM | Mains Sync Set To OFF | 1421
1280 | ZEN Core EM | Choose your ZEN core EM runtime profile | 1422
1280 | ZEN Core EM | Create User Supervisor EM | 1423
1280 | ZEN Core EM | Create User Supervisor LM | 1424
1280 | ZEN Core EM | Create User Supervisor Volutome (only applicable for Volutome systems) | 1425
1280 | ZEN Core EM | Place all users as “Member of” Supervisor group | 1426
1280 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 1427
1280 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 1428
1280 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 1429
1280 | ZEN Core EM | Configue $NavCam (Capital N and C) | 1430
1280 | ZEN Core EM | Configue $Exchange (capital E) | 1431
1280 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 1432
1280 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 1433
1280 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 1434
1280 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 1435
1280 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 1436
1280 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 1437
1251 | aSTEM Resolution | Resolution Result | 1438
756 | Material Availability | Are there any shortages? | 1439
1241 | Rapid BSD | Fit the Rapid BSD | 1440
1241 | Rapid BSD | Software setup of the Rapid BSD | 1441
1241 | Rapid BSD | Test the Rapid BSD | 1442
1241 | Rapid BSD | X-ray check Rapid BSD | 1443
1241 | Rapid BSD | System vac in spec (<2.0e-06mbar) | 1444
1241 | Rapid BSD | Remove detector before shipping | 1445
948 | Insitu stage | Fit the Insitu substage | 1446
948 | Insitu stage | Test the Insitu substage | 1447
948 | Insitu stage | System vac <2.0e-06mbar | 1448
948 | Insitu stage | X-ray test | 1449
948 | Insitu stage | Repackage all hardware excluding the chamber feed through in Deben packaging | 1450
945 | HTBSD | Fit the HT-BSD | 1451
945 | HTBSD | Software setup of the HTBSD | 1452
945 | HTBSD | Test the HTBSD | 1453
945 | HTBSD | X-ray check HT-BSD | 1454
945 | HTBSD | System vac in spec (<2.0e-06mbar) | 1455
945 | HTBSD | Remove detector before shipping | 1456
916 | Volutome | 1. Unpack Volutome and check all parts on the packing list are present (packing list link can be found in attachments) | 1457
916 | Volutome | 1. Upload data from USB stick to the O drive product folder | 1458
916 | Volutome | 1. Remove transit clamps from Volutome stage | 1459
916 | Volutome | 1. Ensure both safety switches are fitted to the Volutome door | 1460
916 | Volutome | 2. Ensure hardware has been configured according to work instructions | 1461
916 | Volutome | 2. Fit Volutome stage | 1462
916 | Volutome | 2. Check protection circuit for Piezo is fitted (349567-9013-000) | 1463
916 | Volutome | 2. Connect Volutome stage and controller to the SEM | 1464
916 | Volutome | 2. Fit Stemi holder bracket to Volutome stage door (STEMI adjustment bracket needs to fit / disconnect without force from door) | 1465
916 | Volutome | 3. Ethernet Connection Setup | 1466
916 | Volutome | 4. Install the latest version of Zen - Ensure to configure it for Volutome (Refer to Zen installation WI) | 1467
916 | Volutome | 4. Ensure Athena inspector manager has all green lights, Physik microtome and Physik stage are present - save a screenshot in the Volutome folder | 1468
916 | Volutome | 4. Install PI Micromove | 1469
916 | Volutome | 4. Run getversions.exe and save text file in the Volutome folder | 1470
916 | Volutome | 4. Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator | 1471
916 | Volutome | 4. Set up the Volutome workbench in Zen | 1472
916 | Volutome | 5. PID parameters configured | 1473
916 | Volutome | 6. Record VolumeBSD serial number | 1474
916 | Volutome | 6. Record Volutome stage serial number | 1475
916 | Volutome | 6. Record VolumeBSD amplifier serial number | 1476
916 | Volutome | 6. Record VolumeBSD mechanics serial number | 1477
916 | Volutome | 6. Record C-885 PI controller serial number | 1478
916 | Volutome | 6. Record E-727 PI controller serial number | 1479
916 | Volutome | 7. Volutome 'internal LED light' function check | 1480
916 | Volutome | 8. Volutome 'External cold light' function check | 1481
916 | Volutome | 9. Adjust Volutome stage centre (Crosshairs should be within 4 central squares) | 1482
916 | Volutome | 10. Install Nav cam in-line connections (If applicable) | 1483
916 | Volutome | 11. Perform the microtome and stage functionality checks | 1484
916 | Volutome | 12. Fit and align the VolumeBSD diode | 1485
916 | Volutome | 12. Check VolumeBSD alignment repeatability | 1486
916 | Volutome | 12. Check all gain ranges of the VolumeBSD (normal/inverted) | 1487
916 | Volutome | 12. Perform Signal/noise ratio check | 1488
916 | Volutome | 12. Ensure BSD diode protector is fitted and adjusted in chamber | 1489
916 | Volutome | 12. Install cable tie to loop BSD signal cable | 1490
916 | Volutome | 12. Test BSD functionality in Zen - Insert, retract and gain ranges all functional | 1491
916 | Volutome | 13. Check SCM with Volutome stage | 1492
916 | Volutome | 14. Oscillation check - Check there is no visible oscillation when imaging with a pixel size of 5nm | 1493
916 | Volutome | 15. Perform X-ray check | 1494
916 | Volutome | 16. Pump Down Time | 1495
916 | Volutome | 16. Ultimate Vacuum | 1496
916 | Volutome | 17. Focal CC fitted to Volutome stage | 1497
916 | Volutome | 17. Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar | 1498
916 | Volutome | 17. Remove carbon needle from focal CC after testing and place in the engineers toolbox container | 1499
916 | Volutome | 18. Generate live image from Axiocam | 1500
916 | Volutome | 19. Dummy experiment completed with no erros reported in Zen - Results saved in Volutome folder | 1501
916 | Volutome | 20. Re-package Volutome and check it's completeness | 1502
916 | Volutome | 21. Check functionality of all additional software installed e.g Atlas | 1503
916 | Volutome | 22. Ensure Arivis Software Licence is available (1 year or permanent) | 1504
916 | Volutome | 23. ManEng/ Senior Eng/ Production Manger Sign off | 1505
11273 | VP Functional Checks | Correct Penning Guage Fitted | 1506
11273 | VP Functional Checks | Chamber Pressure Regulation good | 1507
186 | 3DSM | Install 3DSM Software and Record Software Version | 1508
186 | 3DSM | Set Up Image Using One segment of the BSD | 1509
186 | 3DSM | Run 3DSM and Setup | 1510
186 | 3DSM | Save 3DSM Project | 1511
186 | 3DSM | Store 3DSM Project | 1512
85 | ZEN Automated Imaging | Install ZEN core EM.Exe (This is NOT ZEN Core) | 1513
85 | ZEN Automated Imaging | ZEN core Software Version | 1514
85 | ZEN Automated Imaging | Setup Exchange Position | 1515
85 | ZEN Automated Imaging | Capture Image | 1516
10670 | Fit EO PCB (Standard) | Confirm Standar EO PCB (356133-9004-000) | 1517
10670 | Fit EO PCB (Standard) | EO PCB S/No | 1518
10670 | Fit EO PCB (Standard) | Fit EO PCB | 1519
10670 | Fit EO PCB (Standard) | Confirm Standard EO PCB (356133-9004-000) | 1520
10670 | Fit EO PCB (Standard) | EO PCB S/No | 1521
10670 | Fit EO PCB (Standard) | Fit EO PCB | 1522
1318 | Fit EO PCB (Lithography) | Confirm Standard EO PCB (356133-9004-000) | 1523
1318 | Fit EO PCB (Lithography) | EO PCB S/No | 1524
1318 | Fit EO PCB (Lithography) | Fit EO PCB | 1525
1318 | Fit EO PCB (Lithography) | Confirm Lithography EO PCB (356133-9004-000) | 1526
1318 | Fit EO PCB (Lithography) | EO PCB S/No | 1527
1318 | Fit EO PCB (Lithography) | Fit EO PCB | 1528
11080 | EBBU | Vacuum Stability Test | 1529
11080 | EBBU | Slew Rate <2.5us | 1530
11080 | EBBU | Measure Specimen Current into a FC and then blank the beam with the EBBU. Confirm that measured Specimen current is zero. | 1531
11081 | EBBP | Fit Test tool BB Electronics | 1532
11081 | EBBP | Vacuum Stability Test | 1533
11081 | EBBP | Slew Rate <2.5us | 1534
11081 | EBBP | Measure Specimen Current into a FC and then blank the beam with the EBBU. Confirm that measured Specimen current is zero. | 1535
11081 | EBBP | Remove Test tool BB Electronics | 1536
1555 | Beamsleeve | Leak check | 1537
1555 | Beamsleeve | Pneumatics OK | 1538
1555 | Beamsleeve | Diode Alignment OK | 1539
1555 | Beamsleeve | Good Repeatability | 1540
1555 | Beamsleeve | Bias Voltage Working | 1541
1555 | Beamsleeve | Vac Degradation < 3x10^-5 mbar | 1542
1555 | Beamsleeve | NanoVP Operation Working over full range | 1543
1555 | Beamsleeve | XVP Operation working over full range | 1544
10868 | Atlas 5 | Connect Silver Box and Cables | 1545
10868 | Atlas 5 | Download latest Version Of Software | 1546
10868 | Atlas 5 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info | 1547
10868 | Atlas 5 | Request Licence File from Fibics | 1548
10868 | Atlas 5 | Install Atlas Software | 1549
10868 | Atlas 5 | Install Licence File | 1550
10868 | Atlas 5 | Calibrate SEM | 1551
10868 | Atlas 5 | Obtain Mosaic | 1552
10868 | Atlas 5 | Set ATLAS Rotation Offset | 1553
10868 | Atlas 5 | Save ATLAS Licence Document | 1554
1552 | SE | Collector Working | 1555
1552 | SE | Scintillator Working | 1556
1318 | Fit EO PCB (Lithography) | Confirm Lithography EO PCB (356133-9216-000) | 1557
1318 | Fit EO PCB (Lithography) | EO PCB S/No | 1558
1318 | Fit EO PCB (Lithography) | Fit EO PCB | 1559
1178 | SE detector light leakage check (XB) | SE Detector has no light leak. | 1560
785 | SE Detectors GS460-F | Collector1 Working | 1561
785 | SE Detectors GS460-F | Scintillator1 Working | 1562
785 | SE Detectors GS460-F | Collector2 Working | 1563
785 | SE Detectors GS460-F | Scintillator2 Working | 1564
121 | SESI | Collector Working | 1565
11135 | Gatan 3 View Unpacking | Unpack Gatan + check against packing list (Photograph each layer before removing) | 1566
11135 | Gatan 3 View Unpacking | Set up stage disable Macro in SmartSEM + assign icon on toolbar | 1567
11135 | Gatan 3 View Unpacking | Remove transit clamps from Gatan stage and fit tie rods | 1568
11135 | Gatan 3 View Unpacking | Adjust stage rails so both Klein and Gatan stages can pump without assistance | 1569
11135 | Gatan 3 View Unpacking | If Nav Cam fitted please contact Peter Lewis to have in – line connection made up | 1570
11132 | Gatan 3 View Setup | Plasma clean Gatan stage (20W 2Hrs) | 1571
11132 | Gatan 3 View Setup | Fit On Point detector and diode | 1572
11132 | Gatan 3 View Setup | Remove Mc Affee software from Gatan PC | 1573
11132 | Gatan 3 View Setup | Disable Firewalls on Gatan and Zeiss PC | 1574
11132 | Gatan 3 View Setup | Install and configure remote client on Gatan PC | 1575
11132 | Gatan 3 View Setup | Configure ‘R configure‘ server on Zeiss PC | 1576
11132 | Gatan 3 View Setup | Ping IP adresses | 1577
11132 | Gatan 3 View Setup | Configure Gatan settings | 1578
11133 | Gatan 3 View Testing | Align BSD and obtain image | 1579
11133 | Gatan 3 View Testing | Check SCM with Gatan Stage | 1580
11133 | Gatan 3 View Testing | Ripple X Spec 3nm (600Kx Mag 15KV WD2) | 1581
11133 | Gatan 3 View Testing | Ripple Y Spec 3nm (600Kx Mag 15KV WD2) | 1582
11133 | Gatan 3 View Testing | Resolution 1KV Spec 2.1nm | 1583
11133 | Gatan 3 View Testing | Resolution 15KV Spec 1.2nm | 1584
11133 | Gatan 3 View Testing | Contamination | 1585
11133 | Gatan 3 View Testing | Pump Down Time | 1586
11133 | Gatan 3 View Testing | Ultimate Vacuum | 1587
11133 | Gatan 3 View Testing | Focal CC fitted to 3View stage | 1588
11133 | Gatan 3 View Testing | Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar | 1589
11133 | Gatan 3 View Testing | Charging paper test carried out and images saved. | 1590
11134 | Gatan 3 View Packing | Assemble Gatan vessel pumping accessories. | 1591
11134 | Gatan 3 View Packing | Re-package system | 1592
11134 | Gatan 3 View Packing | ManEng/ Senior Eng/ Production Manger Sign off | 1593
1553 | VPSE | VP Mode Macro selects VPSE when going to VP | 1594
1553 | VPSE | Collector Bias working | 1595
1553 | VPSE | Collector Bias breakdown occurs at >290v @ 30Pa. | 1596
1553 | VPSE | Take Functional Image. | 1597
1550 | Inlens | Function check | 1598
11084 | C2D F2 | VP Mode Macro selects C2D when going to VP | 1599
11084 | C2D F2 | Black Level calibration | 1600
11084 | C2D F2 | Functional Checks ok | 1601
11084 | C2D F2 | Take Functional Image | 1602
11085 | SCD | Black Level Calibration | 1603
11085 | SCD | Functional Check | 1604
11085 | SCD | Take Image | 1605
1318 | Fit EO PCB (Lithography) | Confirm Lithography EO PCB (356133-9216-000) | 1606
1318 | Fit EO PCB (Lithography) | EO PCB S/No | 1607
1318 | Fit EO PCB (Lithography) | Fit EO PCB | 1608
1551 | InlensDUO / EsB | Grid working | 1609
1551 | InlensDUO / EsB | Check Cut Off | 1610
904 | EVO10 Stage Kit | Stage Kit | 1611
447 | AsB-1 | Configure ASB in Administrator | 1612
447 | AsB-1 | Input Offset Cal | 1613
447 | AsB-1 | Black Level Cal | 1614
447 | AsB-1 | Diode Segment Test | 1615
447 | AsB-1 | Diode Mirror Test | 1616
447 | AsB-1 | All Gain Ranges working | 1617
447 | AsB-1 | Check all channels in Invert mode | 1618
448 | AsB-4 | Configure ASB in Administrator | 1619
448 | AsB-4 | Input Offset Cal | 1620
448 | AsB-4 | Black Level Cal | 1621
448 | AsB-4 | Diode Segment Test | 1622
448 | AsB-4 | Diode Mirror Test | 1623
448 | AsB-4 | All Gain Ranges working | 1624
448 | AsB-4 | Check all channels in Invert mode | 1625
449 | BSD-1 | Configure BSD in Administrator | 1626
449 | BSD-1 | Input Offset Cal | 1627
449 | BSD-1 | Black Level Cal | 1628
449 | BSD-1 | Mechanical Alignment | 1629
449 | BSD-1 | Diode Segment Test | 1630
449 | BSD-1 | Diode Mirror Test | 1631
449 | BSD-1 | Vac Degradation < 3x10-5 mbar | 1632
449 | BSD-1 | All Gain Ranges working | 1633
449 | BSD-1 | Check all channels in Invert mode | 1634
450 | BSD-1 VP | Configure VP BSD in Administrator | 1635
450 | BSD-1 VP | Input Offset Cal | 1636
450 | BSD-1 VP | Black Level Cal | 1637
450 | BSD-1 VP | Diode Segment Test | 1638
450 | BSD-1 VP | Diode Mirror Test | 1639
450 | BSD-1 VP | All Gain Ranges working | 1640
450 | BSD-1 VP | Check all channels in Invert mode | 1641
451 | BSD-4 | Configure BSD in Administrator | 1642
451 | BSD-4 | Input Offset Cal | 1643
451 | BSD-4 | Black Level Cal | 1644
451 | BSD-4 | Mechanical Alignment | 1645
451 | BSD-4 | Diode Segment Test | 1646
451 | BSD-4 | Diode Mirror Test | 1647
451 | BSD-4 | Vac Degradation < 3x10-5 mbar | 1648
451 | BSD-4 | All Gain Ranges working | 1649
451 | BSD-4 | Check all channels in Invert mode | 1650
739 | Sense BSD | Record Diode Serial Number: | 1651
739 | Sense BSD | Mechanical Alignment | 1652
739 | Sense BSD | Configure BSD in Administrator | 1653
739 | Sense BSD | Input Offset Cal (Calibration Wizard) | 1654
739 | Sense BSD | Input Offset Cal (ASAP Test Panel) | 1655
739 | Sense BSD | Check BIAS operation | 1656
739 | Sense BSD | Signal To Noise Test (Result >12%) | 1657
739 | Sense BSD | Store Signal To Noise Result (Screen-shot) | 1658
739 | Sense BSD | Set Up Diode Protection Plate | 1659
739 | Sense BSD | Vac Degradation < 3x10-5 mbar | 1660
739 | Sense BSD | Fit Detector Mechanics Cover | 1661
1596 | Customer licenses | Check that the licences installed in the Zeiss Licence Management Tool are correct to the Sales Order and that they functio | 1662
452 | aSTEM 1 | Check each individual segment in normal and inverse modes (Save all images to aSTEM folder) | 1663
452 | aSTEM 1 | Analyze Segments | 1664
452 | aSTEM 1 | Gain Ranges | 1665
452 | aSTEM 1 | Pneumatics | 1666
452 | aSTEM 1 | Alignment | 1667
452 | aSTEM 1 | Stage Safe Position | 1668
452 | aSTEM 1 | Vac Degradation < 3x10-5 mbar | 1669
452 | aSTEM 1 | CZ Detector Cal | 1670
452 | aSTEM 1 | Store Holder Position 1 | 1671
10670 | Fit EO PCB (Standard) | Confirm Standard EO PCB (356133-9004-000) | 1672
10670 | Fit EO PCB (Standard) | EO PCB S/No | 1673
10670 | Fit EO PCB (Standard) | Fit EO PCB | 1674
453 | aSTEM 4 | Check each individual segment in normal and inverse modes (Save all images to aSTEM folder) | 1675
453 | aSTEM 4 | Analyze Segments | 1676
453 | aSTEM 4 | Gain Ranges | 1677
453 | aSTEM 4 | Pneumatics | 1678
453 | aSTEM 4 | Alignment | 1679
453 | aSTEM 4 | Stage Safe Position | 1680
453 | aSTEM 4 | Vac Degradation < 3x10-5 mbar | 1681
453 | aSTEM 4 | CZ Detector Cal | 1682
453 | aSTEM 4 | Store Holder Position 1 | 1683
1554 | CL | Obtain an image using the CL sample.(See WI for detials) | 1684
1556 | Control Panel | Functionality | 1685
1556 | Control Panel | Language | 1686
1563 | Joystick | All Axes Funtion | 1687
1563 | Joystick | Break Button Function | 1688
1557 | Chamberscope 1 (rear) | CCD | 1689
1557 | Chamberscope 1 (rear) | Focus | 1690
1558 | Chamberscope 2 (side) | CCD | 1691
1558 | Chamberscope 2 (side) | Focus | 1692
1559 | Airlock 80/200 mm | Macros loaded/tested | 1693
1559 | Airlock 80/200 mm | $exchange set up | 1694
1559 | Airlock 80/200 mm | Rod Position Correctly Adjusted | 1695
1559 | Airlock 80/200 mm | Front Panel Control all Functional | 1696
1561 | Tandem Decel | EHT Cal - Note: The 'Stage Bias mirror Calib' values should be < 30 V | 1697
1561 | Tandem Decel | Spark Alarm | 1698
1561 | Tandem Decel | Resolution Preperation | 1699
1561 | Tandem Decel | Acquire a customer proof image @ 1kV landing energy 5kV bias on gold on carbon (Resolution not required) | 1700
1560 | Quiet Mode | Disable/Enable | 1701
1560 | Quiet Mode | Save Vac Trace | 1702
1562 | Plasma Cleaner | Run Clean Cycle | 1703
1564 | Charge Compensator | Align Needle | 1704
1564 | Charge Compensator | Pneumatics ok | 1705
1564 | Charge Compensator | Pressure Test (CC Valve fully open) | 1706
1564 | Charge Compensator | Take image of paper | 1707
1565 | Charge Compensator Cleaning Option | Sample Cleaning (O2) | 1708
11023 | Zen | ZEN Version | 1709
11023 | Zen | Correct Version of Windows 10 Installed | 1710
11023 | Zen | Install ZEN | 1711
11023 | Zen | Install Licence | 1712
11023 | Zen | Install Remote API Client | 1713
11023 | Zen | Functional Test | 1714
11023 | Zen | Save Snapped Image | 1715
11023 | Zen | Store Snapped Image | 1716
11023 | Zen | Store Entitlement Certificate | 1717
342 | Temporary Test Points | Carry out any relevant additional tests – Refer to Man Eng if unsure. | 1718
1571 | Probe current | Measured value = | 1719
11325 | SmartStitch | Installation | 1720
11325 | SmartStitch | Functional Test | 1721
11325 | SmartStitch | Store SmartStitch Functional Image | 1722
11325 | SmartStitch | SmartStitch Folder is present on Desktop | 1723
11391 | Connectors | In-Lens and ESB | 1724
11391 | Connectors | EO Dynamic PCB | 1725
11391 | Connectors | EO Static PCB | 1726
11391 | Connectors | Scan Gen PCB (ASAP or SVSEM) | 1727
11391 | Connectors | Check that all column connections are securely fastened. | 1728
11391 | Connectors | Ensure torque of gun earth is 20NM | 1729
11392 | Objective lens re-work checks | Heater & Thermocouple cables plugged into AC unit | 1730
11392 | Objective lens re-work checks | Column mounting bolts all secured | 1731
11392 | Objective lens re-work checks | IGP bracket screws all secured | 1732
11392 | Objective lens re-work checks | All column cabling cable tied neatly | 1733
11396 | Objective lens re-work checks | Heater & Thermocouple cables plugged into AC unit | 1734
11396 | Objective lens re-work checks | Column mounting bolts all secured | 1735
11396 | Objective lens re-work checks | IGP bracket screws all secured | 1736
11396 | Objective lens re-work checks | All column cabling cable tied neatly | 1737
1244 | Fit System Warning Labels | All Warning labels Fitted. | 1738
10964 | Fit Column MU Metal Shield | Column Mechanical Alignment is good | 1739
10964 | Fit Column MU Metal Shield | MU Metal Shield fitted | 1740
343 | Clad Up | Ensure Cladding geometry is correct and all parts aligned. | 1741
343 | Clad Up | Ensure the suppliers Cladding list for both kits is put into the customers pack. | 1742
343 | Clad Up | Fit safety stickers to cladding (Sticker language should match any additional user manuals listed on order). | 1743
344 | Ripple | Actual X-Axis | 1744
344 | Ripple | Store X ripple shots and results | 1745
344 | Ripple | Actual Y-Axis | 1746
344 | Ripple | Store Y ripple shots and results | 1747
11013 | Ripple XB | Actual X-Axis | 1748
11013 | Ripple XB | Store X ripple shots and results | 1749
11013 | Ripple XB | Actual Y-Axis | 1750
11013 | Ripple XB | Store Y ripple shots and results | 1751
11047 | Ripple Raith | Actual X-Axis | 1752
11047 | Ripple Raith | Store X ripple shots and results | 1753
11047 | Ripple Raith | Actual Y-Axis | 1754
11047 | Ripple Raith | Store Y ripple shots and results | 1755
480 | Run Tablefiller | Mu metal shield is fitted | 1756
480 | Run Tablefiller | DC Tablefiller successfully completed | 1757
481 | VP Table filler Entries | VP Tablefiller Entries successfully completed. | 1758
345 | Contamination | Refer to Work Instruction for column conditions | 1759
345 | Contamination | Store contamination shot | 1760
10971 | Check The Probe current | Use Probe I Cal Graph | 1761
346 | Resolution | 1kV = | 1762
346 | Resolution | Store 1KV Resolution Images and Pate image | 1763
346 | Resolution | 15kV = | 1764
346 | Resolution | Store 1KV Resolution Images and Pate image | 1765
1570 | System vacuum (Chamber vacuum) | Reading (≤2.00e-06) = | 1766
347 | Customer Images | 1kV Inlens | 1767
347 | Customer Images | Store 1kV Inlens Customer Image | 1768
347 | Customer Images | 15kV Inlens | 1769
347 | Customer Images | Store 15kV Inlens Customer Image | 1770
11341 | Customer Images (ESB) | 1kV ESB | 1771
11341 | Customer Images (ESB) | Store 1kV ESB Customer Image | 1772
11049 | Dynamic EHT Test Macro | No EHT Trips or Errors During Test | 1773
11049 | Dynamic EHT Test Macro | Store Dynamic EHT gun trace | 1774
11353 | Customer Images (Raith) | 1kV Inlens | 1775
11353 | Customer Images (Raith) | Store 1kV Inlens Customer Image | 1776
11353 | Customer Images (Raith) | 15kV Inlens | 1777
11353 | Customer Images (Raith) | Store 15kV Inlens Customer Image | 1778
11353 | Customer Images (Raith) | 200V Inlens | 1779
11353 | Customer Images (Raith) | Store 200V Inlens Customer Image | 1780
11353 | Customer Images (Raith) | 30KV SE2 | 1781
11353 | Customer Images (Raith) | Store 30kV SE2 Customer Image | 1782
348 | Basic Interlock tests | Chamber Door Interlock | 1783
348 | Basic Interlock tests | Service Switch | 1784
348 | Basic Interlock tests | Lens De-activation | 1785
348 | Basic Interlock tests | Detector De-activation | 1786
348 | Basic Interlock tests | EHT + IPPSU De-activation | 1787
11256 | SCS Complete | Check SCS Folder/Sales Order | 1788
454 | Airlock Interlock Tests | Interlock tested and working | 1789
1586 | Gun vacuum (UHV vacuum) | Reading (≤5.00e-09 mbar) = | 1790
1619 | BSD 1 VP Detector | Pneumatics functional | 1791
1619 | BSD 1 VP Detector | SmartSEM Control tested | 1792
1619 | BSD 1 VP Detector | Images saved | 1793
349 | Earth Continuity Tests. | Record SEAWARD PRIMETEST 100 SN: | 1794
349 | Earth Continuity Tests. | Calibration Due Date: | 1795
349 | Earth Continuity Tests. | Tester verification test; Confirm earth short circuit = (<0.03Ω) | 1796
349 | Earth Continuity Tests. | Test all earth points as shown in procedure. (Max = 0.1Ω) Record max resistance shown | 1797
349 | Earth Continuity Tests. | Confirm correct SN label applied to product: (CSA Compliant / Non-CSA compliant) | 1798
350 | System Vacuum Check | Actual | 1799
351 | Final Gun Vacuum Check | Actual | 1800
354 | Install Customer Licences | Delete the Production Test Licences | 1801
354 | Install Customer Licences | Obtain the customer Entitlement Certificate | 1802
354 | Install Customer Licences | Download into the Engineer folder | 1803
354 | Install Customer Licences | Install the Customer Licences | 1804
354 | Install Customer Licences | Copy Contents of Image Folder to the Product Folder | 1805
354 | Install Customer Licences | Copy EMServer.db to the Product Folder | 1806
354 | Install Customer Licences | Print out the Customer Entitlement Certicate - To be sent to the customer | 1807
355 | Check S/N Label | Ensure the S/N & P/N match the License generator report / Sales order. | 1808
358 | Final Logbook checks: Consistency of data | Confirm all deviations have been cleared on OST - The Andon lamp should be green | 1809
358 | Final Logbook checks: Consistency of data | All verification & Safety tests done and within threshold limits defined | 1810
358 | Final Logbook checks: Consistency of data | Sales Order is latest version | 1811
358 | Final Logbook checks: Consistency of data | All items on Sales order accounted for | 1812
706 | ESB SN Check | Confirm ESB SN matches | 1813
356 | Hardware Dongle Check | Check System And Dongle S/No. Match In Service Panel | 1814
356 | Hardware Dongle Check | Un-tick Omit Dongle In Factory Panel | 1815
356 | Hardware Dongle Check | Check Dongle fits in USB HUB (“ICY BOX”) and replace if needed. | 1816
923 | Pre QA Clean Up | Remove all old / invalid data and images from the PC. | 1817
923 | Pre QA Clean Up | Ensure all image folders are labelled accordingly. | 1818
923 | Pre QA Clean Up | Ensure all 'return to stock' or 'reject parts have been processed. | 1819
701 | EM Server Screening | Confirm no repeated EM Server screening errors have been flagged | 1820
1169 | Raith Probe current | Measured value = | 1821
1169 | Raith Probe current | Nominal Extractor Voltage (If >4.2kv please consult MEng) | 1822
1169 | Raith Probe current | Nominal Extractor Voltage above Datasheet value (If >500v please consult MEng) | 1823
1587 | Probe current (Gem 560) | Measured value = | 1824
11112 | High current Probe Current Raith | Measured value = | 1825
1588 | Maximum Probe Current (20nA) | Measured value = | 1826
1588 | Maximum Probe Current (20nA) | Save annotated image of Max Current | 1827
1589 | Maximum Probe Current (100nA) | Measured value = | 1828
707 | Maximum Probe Current (Raith) | Measured value = | 1829
964 | Max Probe Current (DC) | Actual | 1830
1590 | Cathode Hours | Filament Age | 1831
1591 | Magnification Calibration | Take and save images at Mag 200, 2kX and 35kX. | 1832
1591 | Magnification Calibration | Mag is within +/- 1% for each. | 1833
1591 | Magnification Calibration | Orthognality is within +/- 0.1 deg. | 1834
1592 | Magnification Calibration | Take and save images at Mag 200, 3kX and 35kX. | 1835
1592 | Magnification Calibration | Mag is within +/- 2% for each. | 1836
1592 | Magnification Calibration | Orthognality is within +/- 0.1 deg | 1837
1593 | Column adjustment (Gemini SC) | (10kV, Central Aperture) - Aperture and Guns <60% | 1838
1593 | Column adjustment (Gemini SC) | (10kV, Central Aperture) - Stigs <20% | 1839
1593 | Column adjustment (Gemini SC) | (10kV, Outer Apertures) - Stigs <30% | 1840
1593 | Column adjustment (Gemini SC) | (1kV, Central Aperture) - Aperture and Guns <60% | 1841
1593 | Column adjustment (Gemini SC) | (1kV, Central Aperture) - Stigs <20% | 1842
1593 | Column adjustment (Gemini SC) | (1kV, Outer Apertures) - Stigs <30% | 1843
1593 | Column adjustment (Gemini SC) | <10% user align wobble required on all apertures | 1844
1136 | Column adjustment (Raith) | (10kV, Central Aperture) - Aperture and Guns <50% | 1845
1136 | Column adjustment (Raith) | (10kV, Central Aperture) - Stigs <20% | 1846
1136 | Column adjustment (Raith) | (10kV, Outer Apertures) - Stigs <30% | 1847
1136 | Column adjustment (Raith) | (1kV, Central Aperture) - Aperture and Guns <50% | 1848
1136 | Column adjustment (Raith) | (1kV, Central Aperture) - Stigs <20% | 1849
1136 | Column adjustment (Raith) | (1kV, Outer Apertures) - Stigs <30% | 1850
1136 | Column adjustment (Raith) | <10% user align wobble required on all apertures | 1851
973 | Column adjustment (Low kv lens) | (10kV, Central Aperture) - Aperture and Guns <60% | 1852
973 | Column adjustment (Low kv lens) | (10kV, Central Aperture) - Stigs <35% | 1853
973 | Column adjustment (Low kv lens) | (10kV, Outer Apertures) - Stigs <45% | 1854
973 | Column adjustment (Low kv lens) | (1kV, Central Aperture) - Aperture and Guns <60% | 1855
973 | Column adjustment (Low kv lens) | (1kV, Central Aperture) - Stigs <20% | 1856
973 | Column adjustment (Low kv lens) | (1kV, Outer Apertures) - Stigs <30% | 1857
973 | Column adjustment (Low kv lens) | <10% user align wobble required on all apertures | 1858
11052 | Emission Image | Image with Condenser OFF | 1859
11052 | Emission Image | Image with Condenser ON | 1860
1594 | Table Filler I | Check table editor graph for outliers (Stigs, Apertures, Guns) | 1861
1594 | Table Filler I | Confirm only 2mm & 5mm WD entries stored in table | 1862
1594 | Table Filler I | Confirm only 1x 30kV entry in table | 1863
1594 | Table Filler I | Confirm widefiled entry (Gem 560 ONLY) | 1864
1594 | Table Filler I | Save PDF of table editor plots | 1865
975 | VP Table Filler | Confirm VP alignment values are present and look reasonable in table editor | 1866
974 | Table Filler II | Check table editor graph for grouping and outliers (Apertures, Guns) | 1867
974 | Table Filler II | Wall of Stigs - Confirm no outliers and uniform grouping | 1868
974 | Table Filler II | Review DC tablefiller images | 1869
974 | Table Filler II | Save PDF of table editor plots | 1870
10983 | Positioning Calibration | Position Measurement | 1871
10984 | C2 Calibration | Focus Deviation | 1872
10985 | C3 Calibration | Focus Deviation | 1873
86 | VP-operation (No Beamsleeve) | Chamber vaccuum tracks demand | 1874
1601 | Vent Cycle and Touch alarm | Ventilation of chamber possible. Confirm Gun does not switch off. | 1875
1601 | Vent Cycle and Touch alarm | While system is vented check touch alarm works. | 1876
1601 | Vent Cycle and Touch alarm | Inspect inside of chamber using UV light and goggles (ME Trolley). If contamination found please raise deviation and clean. | 1877
1616 | ESB Detector | Grid Control working | 1878
1616 | ESB Detector | ESB Concentricity in Spec | 1879
1617 | BSD 1 Detector | Pneumatics functional | 1880
1617 | BSD 1 Detector | SmartSEM Control tested | 1881
1617 | BSD 1 Detector | Images saved | 1882
10670 | Fit EO PCB (Standard) | Confirm Standard EO PCB (356133-9004-000) | 1883
10670 | Fit EO PCB (Standard) | EO PCB S/No | 1884
10670 | Fit EO PCB (Standard) | Fit EO PCB | 1885
10670 | Fit EO PCB (Standard) | All Connections Are Secured with Screws | 1886
10670 | Fit EO PCB (Standard) | Cables Are Neatly Routed and Secured | 1887
1318 | Fit EO PCB (Lithography) | Confirm Lithography EO PCB (356133-9216-000) | 1888
1318 | Fit EO PCB (Lithography) | EO PCB S/No | 1889
1318 | Fit EO PCB (Lithography) | Fit EO PCB | 1890
1318 | Fit EO PCB (Lithography) | All Connections Are Secured with Screws | 1891
1318 | Fit EO PCB (Lithography) | Cables Are Neatly Routed and Secured | 1892
11079 | ATLAS | Atlas Hardware Installed | 1893
11079 | ATLAS | Atlas Software Installed | 1894
11079 | ATLAS | Atlas Licence Saved To SmartSEM USB Stick | 1895
1183 | SmartStitch | Check SmartStitch can be run | 1896
1183 | SmartStitch | SmartStitch folder is present on the Desktop | 1897
1576 | VPSE | Function checked in VP mode, signal intensity changes with VP pressure. | 1898
1576 | VPSE | Signal intensity changes with collector voltage. | 1899
11054 | InLens | Function check | 1900
11086 | C2D | Function checked in VP mode, signal intensity changes with VP pressure. | 1901
11086 | C2D | Signal intensity changes with collector voltage. | 1902
11055 | Signal Mixing | Mix SE2 Signal with Inlens Signal | 1903
1620 | aSTEM 1 Detector | Pneumatics functional | 1904
1620 | aSTEM 1 Detector | SmartSEM Control tested | 1905
1620 | aSTEM 1 Detector | Images saved | 1906
1621 | aSTEM 4 Detector | Pneumatics functional | 1907
1621 | aSTEM 4 Detector | SmartSEM Control tested | 1908
1621 | aSTEM 4 Detector | Images saved | 1909
1622 | ASB 1 Detector | SmartSEM Control tested | 1910
1622 | ASB 1 Detector | Images saved | 1911
1623 | ASB 4 Detector | SmartSEM Control tested | 1912
1623 | ASB 4 Detector | Images saved | 1913
1625 | CCD/ 2nd CCD | Function of CCD cameras/ focus. | 1914
1625 | CCD/ 2nd CCD | Confirm that camera toggle function works. | 1915
11078 | Nav Cam -Stage | Stage Moves To Correct Position When Arm Is Opened | 1916
11078 | Nav Cam -Stage | Image Is Captured | 1917
11078 | Nav Cam -Stage | Functional Check | 1918
11382 | Nav Cam - A/Lock | Image Is Captured | 1919
11382 | Nav Cam - A/Lock | Functional Check | 1920
1626 | Internal SCM | Function of Internal Specimen Current Monitor checked. | 1921
1626 | Internal SCM | Confirm when Beam Blanked measured Specimen Current <= 2pa | 1922
667 | Plasma Cleaner | Confirm Plasma Cleaner is working by running a short sample clean | 1923
1627 | Tandem Decel | Confrirm the 'Stage Bias mirror Calib' values are < 30 V | 1924
1627 | Tandem Decel | Confirm that the saved Resolution image is acceptable | 1925
11056 | Touch Alarm | Touch Alarm Functions OK | 1926
11077 | ZEN | ZEN Software Installed | 1927
1116 | Stage Axis Cycle Tests | Run Axis Cycle Test (min - 1Hr) | 1928
1116 | Stage Axis Cycle Tests | Stage Axis Cycle Test with 0.5Kg Weight for 10 Mins | 1929
940 | Ultimate Gun Vacuum | Vacuum should be ≤9.00e-10 mbar. Please consult Man Eng if this cannot be achieved. | 1930
10988 | Gun Run-up | Confirm “EHT ON“ Mode (No Interlock) | 1931
10988 | Gun Run-up | Runup Gun to correct conditions | 1932
11057 | Open X Interface | UPS Open X Interface (X121) is Functioning | 1933
11057 | Open X Interface | Function Checked | 1934
11057 | Open X Interface | Operation of UPS switch , V3 shuts, System Vents and System Shuts Down to Red Button | 1935
10796 | Fit Chamber Valves HV | Fit Vent Valve under chamber. | 1936
11098 | PC Clean Up | Set up Desktop icons (both accounts) | 1937
11098 | PC Clean Up | Remove ASAP Screening Software | 1938
11098 | PC Clean Up | Clear Internet History | 1939
11098 | PC Clean Up | Delete file in download folder | 1940
1585 | Wiring | Check wiring is neat and tidy and properly secured. | 1941
1585 | Wiring | Check cable ties have been cut back sufficiently and no sharp edges remain. | 1942
941 | XB Cable Routing | Ensure CCD Cables are correctly routed (Not Far RHS of cable rack). | 1943
941 | XB Cable Routing | Ensure the HV cable is corrctly ruted (inside of Ion Pump Bracket) | 1944
10779 | SCS Check | Special Fitted | 1945
10779 | SCS Check | If Fitted, Checked by SCS Team Member Name | 1946
362 | Backup Files (SC) | EM Server.db | 1947
362 | Backup Files (SC) | Em Server.Log | 1948
362 | Backup Files (SC) | MyBackup.smartbackup | 1949
362 | Backup Files (SC) | L500.ini | 1950
362 | Backup Files (SC) | License file.xml (N/A for Crossbeam) | 1951
362 | Backup Files (SC) | Registry file.reg | 1952
362 | Backup Files (SC) | CalibSC.dat | 1953
11024 | Backup Files (DC) | EM Server.db | 1954
11024 | Backup Files (DC) | Em Server.Log | 1955
11024 | Backup Files (DC) | MyBackup.smartbackup | 1956
11024 | Backup Files (DC) | L500.ini | 1957
11024 | Backup Files (DC) | License file.xml (N/A for Crossbeam) | 1958
11024 | Backup Files (DC) | Registry file.reg | 1959
11024 | Backup Files (DC) | CalibDC.dat | 1960
11024 | Backup Files (DC) | Export Tablefiller | 1961
11058 | Backup Files (Raith) | Em Server.db | 1962
11058 | Backup Files (Raith) | Em Server.log Text File | 1963
11058 | Backup Files (Raith) | Mybackup.Smartbackup | 1964
11058 | Backup Files (Raith) | Image Folder | 1965
11058 | Backup Files (Raith) | L500.ini | 1966
11058 | Backup Files (Raith) | Dynamic EHT.MN2 | 1967
11058 | Backup Files (Raith) | Probe Stability.MN2 | 1968
11058 | Backup Files (Raith) | License file.xml | 1969
11058 | Backup Files (Raith) | Registry file.reg | 1970
11058 | Backup Files (Raith) | System Checker Full Report | 1971
11058 | Backup Files (Raith) | Export Alignment Data From Table Editor(.MERLINCDB) | 1972
363 | Copy files to Customer Support folder | Customer Images | 1973
363 | Copy files to Customer Support folder | License file | 1974
363 | Copy files to Customer Support folder | Em Server.mdb | 1975
363 | Copy files to Customer Support folder | Saved Registry | 1976
363 | Copy files to Customer Support folder | L500.ini | 1977
363 | Copy files to Customer Support folder | Calibxx.dat | 1978
3360 | Cathode Values - Datasheet | Filament Serial number | 1979
3360 | Cathode Values - Datasheet | Filament heating Current | 1980
3360 | Cathode Values - Datasheet | Extrator Power at 0.2 mA/sr | 1981
3360 | Cathode Values - Datasheet | Total Current at 0.2 mA/sr | 1982
11059 | Raith OEM Customer Data Files | Copy- 1KV.Tif (Tiff customer image) | 1983
11059 | Raith OEM Customer Data Files | Copy- 15KV.Tif (Tiff customer image) | 1984
11059 | Raith OEM Customer Data Files | Copy- 30KV.Tif (Tiff customer image) | 1985
11059 | Raith OEM Customer Data Files | Copy- 200V.Tif (Tiff customer image) | 1986
11059 | Raith OEM Customer Data Files | Copy- All Ripple images (Tiff Files) | 1987
11059 | Raith OEM Customer Data Files | Copy- All Ripple images (JPG Files) | 1988
11059 | Raith OEM Customer Data Files | Copy- 9876543.xml (XML Smartsem Licence) | 1989
11059 | Raith OEM Customer Data Files | Copy- Merlin Compact XX-XX S/O (Licence Rep File) | 1990
11059 | Raith OEM Customer Data Files | Copy- EmServer.accdb (Database) | 1991
11059 | Raith OEM Customer Data Files | Copy- EmServer.txt (Server Log File) | 1992
11059 | Raith OEM Customer Data Files | Copy- L500.ini (SmartSEM Config) | 1993
11059 | Raith OEM Customer Data Files | Copy- xx-xx.reg (Export of Registry File) | 1994
11059 | Raith OEM Customer Data Files | Copy- Mybackup.smartbackup (SmartSEM Backup) | 1995
11059 | Raith OEM Customer Data Files | Copy- Logbook (scanned)-PDF | 1996
11059 | Raith OEM Customer Data Files | Copy- DynamicEHT (mn2 logfile) | 1997
11059 | Raith OEM Customer Data Files | Copy- Probe Conditioning Stability (mn2 logfile) | 1998
11059 | Raith OEM Customer Data Files | Copy- 1KV Res.jpg (PATE Result) | 1999
11059 | Raith OEM Customer Data Files | Copy- 1KV Res.Tif (Tiff Res Image) | 2000
11059 | Raith OEM Customer Data Files | Copy- 15KV.jpg (PATE Results) | 2001
11059 | Raith OEM Customer Data Files | Copy- 15KV.Tif (Tiff Res Image) | 2002
11059 | Raith OEM Customer Data Files | Emission-Condenser OFF image from QA (Tiff Image) | 2003
11059 | Raith OEM Customer Data Files | Emission-Condenser ON image from QA (Tiff Image) | 2004
11059 | Raith OEM Customer Data Files | Merlin Compact User Manual (English) (PDF Manual) | 2005
11059 | Raith OEM Customer Data Files | Merlin Compact User Manual (German) (PDF Manual) | 2006
11059 | Raith OEM Customer Data Files | Merlin Compact Assembly (PDF Manual) | 2007
889 | DJEL OEM Customer Data Files | Copy- WO.xml (XML Smartsem Licence) | 2008
889 | DJEL OEM Customer Data Files | Copy- S/N W/O (Licence Rep File) | 2009
889 | DJEL OEM Customer Data Files | Copy- EmServer.db (Database) | 2010
889 | DJEL OEM Customer Data Files | Copy- EmServer.log (Server Log File) | 2011
889 | DJEL OEM Customer Data Files | Copy- L500.ini (SmartSEM Config) | 2012
889 | DJEL OEM Customer Data Files | Copy- Registery S/N.reg (Export of Registry File) | 2013
889 | DJEL OEM Customer Data Files | Copy- Mybackup.smartbackup (SmartSEM Backup) | 2014
889 | DJEL OEM Customer Data Files | Copy - Optiprobe Calibration.Log | 2015
364 | Update the Production Database on the O: drive | Copy all sub directories and files from the Image directory | 2016
411 | Delete Images and Files | Images and files from the user directory deleted. | 2017
411 | Delete Images and Files | Download Folder Images Deleted | 2018
411 | Delete Images and Files | Empty Recycle bin. | 2019
1137 | Disable Predictive Service Software | Confirm Predictive Service is Disabled | 2020
365 | Load Customer Language Option | Language | 2021
366 | Record Normal Nominal values | Nominal Fil I: | 2022
366 | Record Normal Nominal values | Nominal extractor V: | 2023
367 | Record Imaging Nominal values | High Res Fil I: | 2024
367 | Record Imaging Nominal values | High Res Extractor V: | 2025
368 | Record Analytic Nominal values | Analytic Fil I | 2026
368 | Record Analytic Nominal values | Analytic Extractor V: | 2027
413 | Disable Expert Gun Mode | Go to 'Select Tools and Preferences' and set 'Expert Gun Mode = No' | 2028
205 | Enable FIB | Run the ‘Enable FIB_MultiGIS_ UniGIS’ file | 2029
11068 | PC Clean Up | Remove all Test Macro Shortcuts | 2030
11068 | PC Clean Up | Remove all Test Program Shortcuts | 2031
11068 | PC Clean Up | Set up Desktop icons (both accounts) | 2032
11068 | PC Clean Up | Reset WSUS Settings | 2033
11068 | PC Clean Up | Remove ASAP Screening Software | 2034
11068 | PC Clean Up | Clear Internet History | 2035
11068 | PC Clean Up | Delete file in download folder | 2036
11337 | Customer Documentation | Create Directory C:\Support\Manuals | 2037
11337 | Customer Documentation | Create Shortcut to above Directory on Desktops, called ‘Carl Zeiss User Documentation’. | 2038
11338 | Customer Documentation | C:\Support\Manuals | 2039
11338 | Customer Documentation | Customer USB documentation folder | 2040
11339 | Customer Documentation | Copy SmartSEM instruction manual from Customer USB to C:\Support\Manuals | 2041
891 | Remove MAC Address | Remove MAC address (Internet) | 2042
891 | Remove MAC Address | Remove MAC address (MIC-NET) | 2043
1567 | Pre Shutdown Checks | Remove the sample holder from the system. | 2044
1567 | Pre Shutdown Checks | Remove Docking Brackets and fit Transit clamps. | 2045
1567 | Pre Shutdown Checks | Airfeet deflated and free to move. | 2046
1567 | Pre Shutdown Checks | Backing lines re-configured so that PVSV valve is attached to flexi hose. Secure valve to frame. | 2047
1567 | Pre Shutdown Checks | Confirm that the 3 'o' rings are fitted to the flexi hose backing line | 2048
1567 | Pre Shutdown Checks | If applicable, ensure that the A/L push rod is removed and blank fitted in its place. Ensure A/L is pumped down before powering down system. | 2049
1567 | Pre Shutdown Checks | If applicable, ensure that the Beamsleeve funnel has been removed from the detector | 2050
1567 | Pre Shutdown Checks | Select 'Partial vent on standby' in Smartsem. | 2051
11408 | Pre Shutdown Checks (Raith) | Remove the sample holder from the system. | 2052
11408 | Pre Shutdown Checks (Raith) | Remove Test Tool SE2 and fit to warehouse stand. Fit SE2 blank plate to chamber and SHV connector to Spellman (00000-0302-367). | 2053
11408 | Pre Shutdown Checks (Raith) | Confirm that Raith top mounted SE2 detector is mounted to a warehouse stand and that a protective cap is fitted over the light guide. | 2054
11408 | Pre Shutdown Checks (Raith) | Set Stage position to Z=41mm & X & Y =65mm. | 2055
11408 | Pre Shutdown Checks (Raith) | Remove SCM and pack for shipment (include screws) | 2056
11408 | Pre Shutdown Checks (Raith) | Remove Docking Brackets and fit Transit clamps. | 2057
11408 | Pre Shutdown Checks (Raith) | Airfeet deflated and free to move. | 2058
11408 | Pre Shutdown Checks (Raith) | Backing lines re-configured so that PVSV valve is attached to flexi hose. Secure valve to frame. | 2059
11408 | Pre Shutdown Checks (Raith) | Confirm that the 3 'o' rings are fitted to the flexi hose backing line | 2060
11408 | Pre Shutdown Checks (Raith) | Select 'Partial vent on standby' in Smartsem. | 2061
1568 | System Shutdown Gemini | Remove docking brackets and fit transit brackets | 2062
1568 | System Shutdown Gemini | Power system down. | 2063
1568 | System Shutdown Gemini | Blow water out of the cooling system. | 2064
1568 | System Shutdown Gemini | Put thread lock on pendulum screws and transit block Screws. | 2065
1568 | System Shutdown Gemini | Chamber secured to vibration plate | 2066
1568 | System Shutdown Gemini | Chamber Sercuring Screws Tightened to 25NM | 2067
1568 | System Shutdown Gemini | Chamber Sercuring Screws Tightness Checked By Name: | 2068
1568 | System Shutdown Gemini | Secure Airfeet with foam blocks. | 2069
1568 | System Shutdown Gemini | Fit foam blocks around column cladding. | 2070
1568 | System Shutdown Gemini | Secure Turbo pump with foam blocks. | 2071
1568 | System Shutdown Gemini | Fit Covers to all unused connectors. | 2072
1568 | System Shutdown Gemini | Remove Cekon Plug From Mains Cable | 2073
1568 | System Shutdown Gemini | Fit Live and Neutral Labels to Mains Cable | 2074
1568 | System Shutdown Gemini | Mains and Earth cable stowed on Mains unit. | 2075
1568 | System Shutdown Gemini | Fit shipping Eyes to transit blocks. | 2076
1568 | System Shutdown Gemini | Confirm all column cladding is secure and screws Loctited in place. | 2077
1568 | System Shutdown Gemini | All doors closed and secured. | 2078
1568 | System Shutdown Gemini | All plinth panels removed and stored in transit trolley. | 2079
1568 | System Shutdown Gemini | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel). | 2080
1568 | System Shutdown Gemini | Fit M8 screw to secure stage door to chamber. | 2081
1568 | System Shutdown Gemini | Fit dust covers and secure / remove cables. | 2082
1568 | System Shutdown Gemini | Prepare column cladding transport jig . | 2083
1568 | System Shutdown Gemini | Print shipping stickers . | 2084
1233 | System Shutdown Crossbeam 550L | Remove docking brackets and fit transit brackets | 2085
1233 | System Shutdown Crossbeam 550L | Power system down. | 2086
1233 | System Shutdown Crossbeam 550L | Blow water out of the cooling system. | 2087
1233 | System Shutdown Crossbeam 550L | Put thread lock on pendulum screws and transit block Screws. | 2088
1233 | System Shutdown Crossbeam 550L | Chamber secured to vibration plate (N/A to XB 550L) | 2089
1233 | System Shutdown Crossbeam 550L | Chamber securing screws tightened to 25nm (N/A to XB 550L) | 2090
1233 | System Shutdown Crossbeam 550L | Chamber securing screws tightness checked. Type your initials (For XB 550L type N/A) | 2091
1233 | System Shutdown Crossbeam 550L | Secure Airfeet with foam blocks. | 2092
1233 | System Shutdown Crossbeam 550L | Secure Turbo pump with foam blocks. | 2093
1233 | System Shutdown Crossbeam 550L | Fit Covers to all unused connectors. | 2094
1233 | System Shutdown Crossbeam 550L | Mains and Earth cable stowed on Mains unit. | 2095
1233 | System Shutdown Crossbeam 550L | Fit shipping Eyes to transit blocks. | 2096
1233 | System Shutdown Crossbeam 550L | All doors closed and secured. | 2097
1233 | System Shutdown Crossbeam 550L | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel). | 2098
1233 | System Shutdown Crossbeam 550L | Print shipping stickers. | 2099
1233 | System Shutdown Crossbeam 550L | Fit 8x M6x55mm cap head screws to stage door | 2100
1174 | Generate the following Documents (polo) | X-Ray Certificate | 2101
1174 | Generate the following Documents (polo) | DENKA Data Sheet | 2102
1174 | Generate the following Documents (polo) | Options Removed packing List | 2103
1174 | Generate the following Documents (polo) | Options Retained packing List | 2104
1174 | Generate the following Documents (polo) | Standard Packing List | 2105
887 | System Shutdown DJEL | Remove docking brackets and fit transit brackets | 2106
887 | System Shutdown DJEL | Vent the Chamber. | 2107
887 | System Shutdown DJEL | Power system down. | 2108
887 | System Shutdown DJEL | Blow water out of the cooling system. | 2109
887 | System Shutdown DJEL | Remove On / Off switch panel | 2110
887 | System Shutdown DJEL | Print and fit DJEL S/N Label | 2111
887 | System Shutdown DJEL | Remove column cables only | 2112
887 | System Shutdown DJEL | Remove column water pipes only | 2113
887 | System Shutdown DJEL | Remove ASAP, EOS, EOD and Vacuum PCBS from plinth and pack seperately in static bag and bubble wrap. | 2114
887 | System Shutdown DJEL | Remove AC unit complete with mains cable | 2115
887 | System Shutdown DJEL | Remove Cekon Plug From Mains Cable | 2116
887 | System Shutdown DJEL | Remove PSU unit | 2117
887 | System Shutdown DJEL | Remove EHT Set | 2118
887 | System Shutdown DJEL | Remove IGP PSU | 2119
887 | System Shutdown DJEL | Remove EBBU Electonics from column and fit shorting plug | 2120
887 | System Shutdown DJEL | Remove customer SE2 detector (460F only) | 2121
887 | System Shutdown DJEL | Fit customer SE2 Detectors (2 off) to transport vessels (460 F only). | 2122
887 | System Shutdown DJEL | Remove BSD4 complete from chamber and fit protective cover over Diode (460F only). | 2123
887 | System Shutdown DJEL | Pack customer accessories as shown WI (460F Only) | 2124
887 | System Shutdown DJEL | Remove Water flow and Temperature sensor assembly | 2125
887 | System Shutdown DJEL | Remove column / support frame complete and attach to column storage frame | 2126
887 | System Shutdown DJEL | Fit blank plate to column bypass pipe so that lower column can be shipped at vacuum (460F Only). | 2127
887 | System Shutdown DJEL | Pump down lower column as shown in WI. | 2128
887 | System Shutdown DJEL | Chamber secured to vibration plate | 2129
887 | System Shutdown DJEL | Chamber Securing Screws Tightened to 25NM | 2130
887 | System Shutdown DJEL | Secure the pendulum feet using foam blocks and strap - Only required if test tool is being shipped to Bollore | 2131
887 | System Shutdown DJEL | Print shipping stickers. | 2132
11123 | System Shutdown Crossbeam | Remove docking brackets and fit transit brackets | 2133
11123 | System Shutdown Crossbeam | Power system down. | 2134
11123 | System Shutdown Crossbeam | Blow water out of the cooling system. | 2135
11123 | System Shutdown Crossbeam | Put thread lock on pendulum screws and transit block Screws. | 2136
11123 | System Shutdown Crossbeam | Chamber secured to vibration plate (N/A to XB 550L) | 2137
11123 | System Shutdown Crossbeam | Chamber securing screws tightened to 25nm (N/A to XB 550L) | 2138
11123 | System Shutdown Crossbeam | Chamber securing screws tightness checked. Type your initials (For XB 550L type N/A) | 2139
11123 | System Shutdown Crossbeam | Secure Airfeet with foam blocks. | 2140
11123 | System Shutdown Crossbeam | Secure Turbo pump with foam blocks. | 2141
11123 | System Shutdown Crossbeam | Fit Covers to all unused connectors. | 2142
11123 | System Shutdown Crossbeam | Mains and Earth cable stowed on Mains unit. | 2143
11123 | System Shutdown Crossbeam | Fit shipping Eyes to transit blocks. | 2144
11123 | System Shutdown Crossbeam | All doors closed and secured. | 2145
11123 | System Shutdown Crossbeam | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel). | 2146
11123 | System Shutdown Crossbeam | Print shipping stickers. | 2147
11123 | System Shutdown Crossbeam | Fit 6x M8x16mm cap head screws to stage door | 2148
11407 | System Shutdown Raith | Remove docking brackets and fit transit brackets | 2149
11407 | System Shutdown Raith | Power system down. | 2150
11407 | System Shutdown Raith | Blow water out of the cooling system. | 2151
11407 | System Shutdown Raith | Chamber secured to vibration plate | 2152
11407 | System Shutdown Raith | Chamber Sercuring Screws Tightened to 25NM | 2153
11407 | System Shutdown Raith | Chamber Sercuring Screws Tightness Checked By Name: | 2154
11407 | System Shutdown Raith | Check that the Stage bearing blocks screws are secure. | 2155
11407 | System Shutdown Raith | Remove test tool stage | 2156
11407 | System Shutdown Raith | Remove test tool stage PCB | 2157
11407 | System Shutdown Raith | Secure Airfeet with foam blocks. | 2158
11407 | System Shutdown Raith | Secure Turbo pump with foam blocks. | 2159
11407 | System Shutdown Raith | Confirm there is no cooling fluid laying in the floor of the plinth | 2160
11407 | System Shutdown Raith | Fit Covers to all unused connectors. | 2161
11407 | System Shutdown Raith | Remove Cekon Plug From Mains Cable | 2162
11407 | System Shutdown Raith | Fit Live and Neutral Labels to Mains Cable | 2163
11407 | System Shutdown Raith | Mains and Earth cable stowed on Mains unit. | 2164
11407 | System Shutdown Raith | Fit shipping Eyes to transit blocks. | 2165
11407 | System Shutdown Raith | Put thread lock on pendulum screws and transit block Screws. | 2166
11407 | System Shutdown Raith | All doors closed and secured. | 2167
11407 | System Shutdown Raith | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel). | 2168
11407 | System Shutdown Raith | Print shipping stickers. | 2169
371 | Generate the following Documents (Gemini) | X-Ray Certificate | 2170
371 | Generate the following Documents (Gemini) | Quality Certificate | 2171
371 | Generate the following Documents (Gemini) | Customer Entitlement Certificate | 2172
371 | Generate the following Documents (Gemini) | Quality Feedback Form | 2173
371 | Generate the following Documents (Gemini) | DENKA Data Sheet | 2174
371 | Generate the following Documents (Gemini) | Options Removed packing List | 2175
371 | Generate the following Documents (Gemini) | Options Retained packing List | 2176
371 | Generate the following Documents (Gemini) | Standard Packing List | 2177
371 | Generate the following Documents (Gemini) | PC Inferface assignment | 2178
371 | Generate the following Documents (Gemini) | Declaration of conformity | 2179
371 | Generate the following Documents (Gemini) | Installation report | 2180
371 | Generate the following Documents (Gemini) | Acceptance test record | 2181
371 | Generate the following Documents (Gemini) | Licence Generator report | 2182
11060 | Generate the Following Documents (Raith) | X-Ray Certificate ---------------------------------------------------------------------------------- | 2183
11060 | Generate the Following Documents (Raith) | Quality Certificate --------------------------------------------------------------------------------- | 2184
11060 | Generate the Following Documents (Raith) | Quality Feedback Form | 2185
11060 | Generate the Following Documents (Raith) | DENKA Data Sheet | 2186
11060 | Generate the Following Documents (Raith) | Options Retained packing List | 2187
11060 | Generate the Following Documents (Raith) | Standard Packing List | 2188
11060 | Generate the Following Documents (Raith) | Declaration of Conformity | 2189
11060 | Generate the Following Documents (Raith) | Raith Installation Report | 2190
11060 | Generate the Following Documents (Raith) | Software Licence Agreement | 2191
1118 | EVO15/ Sigma Stage Kit | Stage Kit | 2192
1118 | EVO15/ Sigma Stage Kit | 27mm Spacer Block Fitted | 2193
888 | Generate the Following Documents (DJEL) | X-Ray Certificate ---------------------------------------------------------------------------------- | 2194
888 | Generate the Following Documents (DJEL) | Quality Certificate --------------------------------------------------------------------------------- | 2195
888 | Generate the Following Documents (DJEL) | DENKA Data Sheet | 2196
888 | Generate the Following Documents (DJEL) | Standard Packing List | 2197
11099 | Generate the following Documents | X-Ray Certificate | 2198
11099 | Generate the following Documents | Quality Certificate | 2199
11099 | Generate the following Documents | SVSEM FIB EXT Datasheet | 2200
11099 | Generate the following Documents | DENKA Data Sheet | 2201
11099 | Generate the following Documents | Options Removed packing List | 2202
11099 | Generate the following Documents | Options Retained packing List | 2203
11099 | Generate the following Documents | Standard Packing List | 2204
11099 | Generate the following Documents | Send all supplier paperwork to Oberkochen | 2205
195 | Check 2 Off Lindy Mini DP to DP cables Supplied - | Lindy Cables Supplied | 2206
1637 | Customer documentation check | X-Ray Certificate | 2207
1637 | Customer documentation check | Quality Certificate | 2208
1637 | Customer documentation check | Quality Feedback Form | 2209
1637 | Customer documentation check | Denka Data Sheet | 2210
1637 | Customer documentation check | Options Removed Packing list | 2211
1637 | Customer documentation check | Options Retained Packing list | 2212
1637 | Customer documentation check | Standard Packing List | 2213
1637 | Customer documentation check | PC Interface assignment | 2214
1637 | Customer documentation check | Declaration of conformity | 2215
1637 | Customer documentation check | Installation report | 2216
1637 | Customer documentation check | Acceptance test record | 2217
1637 | Customer documentation check | License Generator Report | 2218
1637 | Customer documentation check | Rapid Quality Feedback | 2219
1292 | Customer documentation check (POLO +) | X-Ray Certificate | 2220
1292 | Customer documentation check (POLO +) | Denka Data Sheet | 2221
1292 | Customer documentation check (POLO +) | Options Removed Packing list | 2222
1292 | Customer documentation check (POLO +) | Options Retained Packing list | 2223
1292 | Customer documentation check (POLO +) | Standard Packing List | 2224
11065 | Customer Documentation Check | X-Ray Certificate | 2225
11065 | Customer Documentation Check | Quality Certificate | 2226
11065 | Customer Documentation Check | Quality Feedback Form | 2227
11065 | Customer Documentation Check | DENKA Data Sheet | 2228
11065 | Customer Documentation Check | Options Retained packing List | 2229
11065 | Customer Documentation Check | Standard Packing List | 2230
11065 | Customer Documentation Check | Declaration of Conformity | 2231
11065 | Customer Documentation Check | Raith Installation Report | 2232
11065 | Customer Documentation Check | Software Licence Agreement | 2233
1135 | Print shipping labels and idenitfy system material | Print shipment labels | 2234
1135 | Print shipping labels and idenitfy system material | Label material boxes | 2235
1135 | Print shipping labels and idenitfy system material | Place left over stickers in document holder | 2236
959 | Declaration of conformity (Sigma 300) | Please print 'Declaration of Conformity' from URL link | 2237
956 | Declaration of conformity (Sigma 360) | Please print 'Declaration of Conformity' from URL link | 2238
958 | Declaration of conformity (Sigma 500) | Please print 'Declaration of Conformity' from URL link | 2239
957 | Declaration of conformity (Sigma 560) | Please print 'Declaration of Conformity' from URL link | 2240
11101 | SCS Check | Check SCS Folder/Sales Order | 2241
11101 | SCS Check | SCS Number (If applicable) | 2242
3242 | PC Installation | Select PC type | 2243
3242 | PC Installation | PC S/No. | 2244
3242 | PC Installation | Is SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, | 2245
3242 | PC Installation | Hard Drive 1 Size | 2246
3242 | PC Installation | Hard Drive 2 Size | 2247
3242 | PC Installation | RAM | 2248
3242 | PC Installation | Processor | 2249
3242 | PC Installation | Operating System | 2250
3242 | PC Installation | Change Desktop Settings | 2251
3242 | PC Installation | Install SmartSEM and record Version | 2252
3242 | PC Installation | Check for Service Packs and install them | 2253
3242 | PC Installation | Upgrade ASAP Firmware (check Service Pack for latest version) and API | 2254
3242 | PC Installation | Copy the SmartSEM sotware into a Folder in C:Support and Name it SmartSEM X.x (i.e 8.2) Software | 2255
3242 | PC Installation | Installation of Zen Core EM if Hasap is fitted | 2256
3242 | PC Installation | Set BIOS AC Power loss to ‘ON’ | 2257
3242 | PC Installation | Set the Time and Date on the PC to current UK | 2258
3242 | PC Installation | Attach Cable Markers to USB Cables | 2259
708 | PC MAC Address | Record MAC address (Internet not MIC-NET) | 2260
1168 | Install Production Test Licence | Download Licence from SharePoint | 2261
1168 | Install Production Test Licence | Delete the Licence in SharePoint | 2262
1168 | Install Production Test Licence | Install the licence using the Zeiss License Management Tool. | 2263
1168 | Install Production Test Licence | Installation of Zen licence if Hasap is fitted using the Zeiss activation tool | 2264
3391 | Aperture alignment | 10kV All apertures - Confirm wobble can be achieved in user align window <20% | 2265
1294 | Setup the Support folder | Select the Support folder in the root dir (c:) | 2266
1294 | Setup the Support folder | Create a Data folder within the Support folder | 2267
1294 | Setup the Support folder | In the Data folder create the following new sub-folders- Images, Licences, Log files, SmartSEM Software. | 2268
3243 | Install Customer Language Option | Language | 2269
11271 | Install Predictive Service Software | Predictive Service S/W installed on Customer PC. | 2270
11271 | Install Predictive Service Software | Do Not Activate | 2271
11271 | Install Predictive Service Software | Copy the Predictive Service.json file into C:Support/Data/Licence folder. | 2272
3245 | Beam Blanker | Check Cap is Fitted | 2273
3245 | Beam Blanker | Check Earth Cable is Fitted | 2274
3247 | Vibration System | Set Air Suspension Height | 2275
3247 | Vibration System | Check Correct Operation | 2276
3246 | Connect Plinth Services | Air Pressure O/P = 6bar | 2277
3246 | Connect Plinth Services | Water | 2278
3246 | Connect Plinth Services | Mains | 2279
3248 | Heater Cables And Box Configuration | Plate | 2280
3248 | Heater Cables And Box Configuration | Screws | 2281
3248 | Heater Cables And Box Configuration | Spacers | 2282
3249 | Power Up System | Check That AC Loss Policy Depend On, is Set To Power On | 2283
3249 | Power Up System | Disable Windows Fast Start UP | 2284
3250 | Review customer order | Sales order is latest revision | 2285
3250 | Review customer order | All accessories are accounted for and are fitted to correct ports | 2286
3250 | Review customer order | Identify any un-described items on sales order | 2287
3252 | EO P.C.B Versions | F/W Version | 2288
3252 | EO P.C.B Versions | Hardware Version | 2289
3253 | EHT Set Versions | F/W Version | 2290
3253 | EHT Set Versions | Hardware Version | 2291
3255 | Stage Versions | Controller Version | 2292
3255 | Stage Versions | Axes Version | 2293
3255 | Stage Versions | Hardware Version | 2294
3254 | ASAP / SVSEM Versions | F/W Version | 2295
3254 | ASAP / SVSEM Versions | Hardware Version | 2296
3257 | SmartSEM Configuration | In SEM status set 'Mains Sync' to off | 2297
1286 | Set EO Temp Limit to 40 | Set EO Temp Limit to 40 deg C | 2298
1048 | Confirm 4-eye check by additional operator | Checked Carried Our By: | 2299
1220 | ASAP Type Check | Check if ASAP or h-ASAP | 2300
1220 | ASAP Type Check | h-ASAP Serial No (if applicable, if not applicable enter NA) | 2301
10690 | Fit ASAP | ASAP Serial Number | 2302
3261 | Smart SEM Bake Out | Opening of column manual valve/s required to achieve gun vacuum | 2303
3261 | Smart SEM Bake Out | Bake Out Hours | 2304
3265 | Check Serial No. on Datasheet & Column | Serial No. | 2305
3270 | Check Specimen Current @ 10kV, WD5, Condenser Off | Data Sheet Vext | 2306
3270 | Check Specimen Current @ 10kV, WD5, Condenser Off | Vext | 2307
3270 | Check Specimen Current @ 10kV, WD5, Condenser Off | Speciment Current | 2308
3271 | Check Specimen Current @ 10kV, WD5, Condenser On | Data Sheet Vext | 2309
3271 | Check Specimen Current @ 10kV, WD5, Condenser On | Vext | 2310
3271 | Check Specimen Current @ 10kV, WD5, Condenser On | Speciment Current | 2311
1225 | Hasap scanning calibration using Zen | carry out scanning calibration (Hasap only) | 2312
1126 | Orth Mag 1 and Mag 2 Proof images. | Store Orth Picture Mag Range 1 | 2313
1126 | Orth Mag 1 and Mag 2 Proof images. | Store Orth Picture Mag Range 2 | 2314
1287 | Stage X and Y Macro | X/Y Macro Run For 30 Minutes | 2315
6927 | Stage Axis Cycle Test | Macro run time | 2316
6928 | Stage Repeatability Macro | Select Stage Type | 2317
6928 | Stage Repeatability Macro | Macro Result 1 ($[StageRepeatability]) | 2318
6928 | Stage Repeatability Macro | Macro Result 2 ($[StageRepeatability]) | 2319
6928 | Stage Repeatability Macro | Macro Result 3 ($[StageRepeatability]) | 2320
6928 | Stage Repeatability Macro | Image of Macro Results | 2321
6929 | Stage Centre Wizard | X | 2322
6929 | Stage Centre Wizard | Y | 2323
6931 | Stage Drift Program | Drift | 2324
6932 | Holder Calibration | Spacer Height Setup | 2325
6932 | Holder Calibration | Holder Height | 2326
6932 | Holder Calibration | Holder Rotation | 2327
10691 | Fit EO PCB | EO PCB Serial Number | 2328
10692 | Fit Stage PCB | Stage PCB Serial Number | 2329
11398 | Install PSU's | Fit Vac/ASAP PSU | 2330
11398 | Install PSU's | Vac/ASAP PSU Serial Number | 2331
11398 | Install PSU's | Connect Turbo Controller PCB Power (P2) | 2332
11398 | Install PSU's | Fit EO/Stage PCB | 2333
11398 | Install PSU's | EO/Stage PSU Serial Number | 2334
10701 | Mains Box | EHT Cable PL5 is Plugged Into PL5 | 2335
10701 | Mains Box | All Connectors Are Securely Fitted | 2336
10701 | Mains Box | Filter Earth Link Fitted | 2337
10701 | Mains Box | Cables Are Neatly Routed And Secured | 2338
10768 | CAN Bus Test | Vacuum PCB Test Passed | 2339
10768 | CAN Bus Test | EO PCB Test Passed | 2340
10768 | CAN Bus Test | Stage PCB Test Passed | 2341
11247 | Production Documentation | Print Plinth Seperators | 2342
11247 | Production Documentation | Sort Documentation with Seperators | 2343
10775 | All Paperwork Present | Picking List | 2344
10775 | All Paperwork Present | Turbo Pump | 2345
10775 | All Paperwork Present | Turbo Pump Controller | 2346
10775 | All Paperwork Present | Chamber | 2347
10775 | All Paperwork Present | Plinth | 2348
10775 | All Paperwork Present | EO PCB | 2349
10775 | All Paperwork Present | Stage PCB | 2350
10775 | All Paperwork Present | Stage Logbook | 2351
10781 | Chamber & Table Top | Chamber is Free From Defects | 2352
10781 | Chamber & Table Top | Chamber Meshes Are Fitted | 2353
10781 | Chamber & Table Top | Chamber Pumping Pipe Is Fitted Centrally To The Table Top | 2354
10781 | Chamber & Table Top | Table Top Is Free From Defects | 2355
10782 | Plinth | Plinth Base Is Clean and Free From Debris | 2356
10783 | Bay | Clean & Tidy Assembly Bay | 2357
6933 | Touch Alarm Test | Stage stops when touch alarm is initiated or F12 is pressed | 2358
3278 | C2D Detector function tests | Black Level Cal | 2359
3278 | C2D Detector function tests | C2D Bias Target Check | 2360
3278 | C2D Detector function tests | Gain Range Check | 2361
3278 | C2D Detector function tests | Auto Level Check | 2362
3278 | C2D Detector function tests | Auto Range Check | 2363
3278 | C2D Detector function tests | Function Image | 2364
11044 | Pre aligment checks | Set Aperture sizes in alignment panel | 2365
11044 | Pre aligment checks | Set Aperture default Gun and Aperture alignment figures | 2366
11044 | Pre aligment checks | Re-align aperture 1 | 2367
1143 | Calibrate Stigmator Balance | Stigmators Balanced On All Apertures (Normal Current) | 2368
1143 | Calibrate Stigmator Balance | Stigmators Balanced On Aperture 5 (High Current ) - X Upper Value | 2369
1143 | Calibrate Stigmator Balance | Stigmators Balanced On Aperture 5 (High Current ) - X Lower Value | 2370
1143 | Calibrate Stigmator Balance | Stigmators Balanced On Aperture 5 (High Current ) - Y Upper Value | 2371
1143 | Calibrate Stigmator Balance | Stigmators Balanced On Aperture 5 (High Current ) - Y Lower Value | 2372
1143 | Calibrate Stigmator Balance | Stigmators Balanced On Aperture 6 (High Current ) - X Upper Value | 2373
1143 | Calibrate Stigmator Balance | Stigmators Balanced On Aperture 6 (High Current ) - X Lower Value | 2374
1143 | Calibrate Stigmator Balance | Stigmators Balanced On Aperture 6 (High Current ) - Y Upper Value | 2375
1143 | Calibrate Stigmator Balance | Stigmators Balanced On Aperture 6 (High Current ) - Y Lower Value | 2376
1143 | Calibrate Stigmator Balance | Stigmators Balanced On Aperture 7 (High Current ) - X Upper Value | 2377
1143 | Calibrate Stigmator Balance | Stigmators Balanced On Aperture 7 (High Current ) - X Lower Value | 2378
1143 | Calibrate Stigmator Balance | Stigmators Balanced On Aperture 7 (High Current ) - Y Upper Value | 2379
1143 | Calibrate Stigmator Balance | Stigmators Balanced On Aperture 7 (High Current ) - Y Lower Value | 2380
1143 | Calibrate Stigmator Balance | Stigmators Balanced On The Remaining Apertures (High Current ) | 2381
3283 | Calibrate Max. Extractor Voltage | EXT Voltage | 2382
3284 | Run OPTI Probe | Max specimen current | 2383
3287 | Aperture Alignment Check (WD 5mm 20% max) | 1kV X | 2384
3287 | Aperture Alignment Check (WD 5mm 20% max) | 1kV Y | 2385
3287 | Aperture Alignment Check (WD 5mm 20% max) | 10kV X | 2386
3287 | Aperture Alignment Check (WD 5mm 20% max) | 10kV Y | 2387
3287 | Aperture Alignment Check (WD 5mm 20% max) | 20kV X | 2388
3287 | Aperture Alignment Check (WD 5mm 20% max) | 20kV Y | 2389
11259 | Input High Res Tip Data into Gun Parameters | Fil I | 2390
11259 | Input High Res Tip Data into Gun Parameters | Ext V | 2391
3288 | Stigmator Figures (WD 5mm 20% max) | 1kV X | 2392
3288 | Stigmator Figures (WD 5mm 20% max) | 1kV Y | 2393
3288 | Stigmator Figures (WD 5mm 20% max) | 10kV X | 2394
3288 | Stigmator Figures (WD 5mm 20% max) | 10kV Y | 2395
3288 | Stigmator Figures (WD 5mm 20% max) | 20kV X | 2396
3288 | Stigmator Figures (WD 5mm 20% max) | 20kV Y | 2397
3414 | Check Column, Chamber, and Plinth for X-ray Leakage. | Accelerating Voltage=30KV | 2398
3414 | Check Column, Chamber, and Plinth for X-ray Leakage. | Specimen Current | 2399
3414 | Check Column, Chamber, and Plinth for X-ray Leakage. | Monitor Model | 2400
3414 | Check Column, Chamber, and Plinth for X-ray Leakage. | Serial Number | 2401
3414 | Check Column, Chamber, and Plinth for X-ray Leakage. | Cal. Expiry Date | 2402
3414 | Check Column, Chamber, and Plinth for X-ray Leakage. | No X-Ray Leaks Detected | 2403
3415 | X-Ray Test | Serial No. (or N/A) | 2404
3415 | X-Ray Test | No X-Ray Leakage detected. | 2405
11010 | EHT Earth Check | Ensure EHT Earth Connection is Torqued to 3Nm | 2406
1176 | Sigma560 - Cartesian Stage | 15kv 0.9nm | 2407
1176 | Sigma560 - Cartesian Stage | Store 15KV Resolution Images and Pate image for 2 x consecutive images | 2408
1176 | Sigma560 - Cartesian Stage | 1kv 1.3nm | 2409
1176 | Sigma560 - Cartesian Stage | Store 1KV Resolution Images and Pate image for 2 x consecutive images | 2410
1176 | Sigma560 - Cartesian Stage | 500V 2nm | 2411
1176 | Sigma560 - Cartesian Stage | Store 500V Resolution Images and Pate image for 2 x consecutive images | 2412
3291 | Ripple (Target $[MaxRipple]nm) | X = | 2413
3291 | Ripple (Target $[MaxRipple]nm) | Y = | 2414
11425 | 360 Ressoution Measurements | 15kv 0.9nm | 2415
11425 | 360 Ressoution Measurements | Store 15KV Resolution Images and Pate image for 2 x consecutive images | 2416
11425 | 360 Ressoution Measurements | 1kv 1.3nm | 2417
11425 | 360 Ressoution Measurements | Store 1KV Resolution Images and Pate image for 2 x consecutive images | 2418
11425 | 360 Ressoution Measurements | 500V 2nm | 2419
11425 | 360 Ressoution Measurements | Store 500V Resolution Images and Pate image for 2 x consecutive images | 2420
11426 | 560 Ressoution Measurements | 15kv 0.7nm | 2421
11426 | 560 Ressoution Measurements | Store 15KV Resolution Images and Pate image for 2 x consecutive images | 2422
11426 | 560 Ressoution Measurements | 1kv 1.2nm | 2423
11426 | 560 Ressoution Measurements | Store 1KV Resolution Images and Pate image for 2 x consecutive images | 2424
11426 | 560 Ressoution Measurements | 500V 1.9nm | 2425
11426 | 560 Ressoution Measurements | Store 500V Resolution Images and Pate image for 2 x consecutive images | 2426
3292 | Standard Resolution (2 x consecutive images) | 1kV (Spec: $[Resolution1kV]nm) | 2427
3292 | Standard Resolution (2 x consecutive images) | Store 1KV Resolution Images and Pate image for 2 x consecutive images | 2428
3292 | Standard Resolution (2 x consecutive images) | 15kV (Spec: $[Resolution15kV]nm) | 2429
3292 | Standard Resolution (2 x consecutive images) | Store 15KV Resolution Images and Pate image for 2 x consecutive images | 2430
1279 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 2431
1279 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 2432
1279 | ZEN Core EM | Install Customer's ZEN Licence | 2433
1279 | ZEN Core EM | Setup SmartSEM Toggle Button | 2434
1279 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 2435
1279 | ZEN Core EM | Mains Sync Set To OFF | 2436
1279 | ZEN Core EM | Choose your ZEN core EM runtime profile | 2437
1279 | ZEN Core EM | Create User Supervisor EM | 2438
1279 | ZEN Core EM | Create User Supervisor LM | 2439
1279 | ZEN Core EM | Create User Supervisor Volutome (only applicable for Volutome systems) | 2440
1279 | ZEN Core EM | Place all users as “Member of” Supervisor group | 2441
1279 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 2442
1279 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 2443
1279 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 2444
1279 | ZEN Core EM | Configue $NavCam (Capital N and C) | 2445
1279 | ZEN Core EM | Configue $Exchange (capital E) | 2446
1279 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 2447
1279 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 2448
1279 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 2449
1279 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 2450
1279 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 2451
1279 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 2452
1245 | HP Stage Mapping | Stage mapping completed successfuly | 2453
950 | Insitu stage | Fit the Insitu substage | 2454
950 | Insitu stage | Test the Insitu substage | 2455
950 | Insitu stage | System vac <2.0e-06mbar | 2456
950 | Insitu stage | X-ray test | 2457
950 | Insitu stage | Repackage all hardware excluding the chamber feed through in Deben packaging | 2458
11082 | CL Detector | Functional Check | 2459
11082 | CL Detector | Function Image | 2460
3296 | Inlens Detector | Function Test | 2461
10867 | Additional Hard Disk & Cloning Software | Install Hard Disk | 2462
10867 | Additional Hard Disk & Cloning Software | Format Hard Disk (NTFS) | 2463
10867 | Additional Hard Disk & Cloning Software | Hard Disk Is Seen In File Explorer | 2464
10867 | Additional Hard Disk & Cloning Software | Copy A File To The Hard Disk, Then Delete It | 2465
10867 | Additional Hard Disk & Cloning Software | Install Cloning Software | 2466
3297 | Inlens-Duo Detector | BSE Collector Voltage Test | 2467
3297 | Inlens-Duo Detector | BSE Macro | 2468
3297 | Inlens-Duo Detector | BSE 0V Bias Image | 2469
3297 | Inlens-Duo Detector | BSE 620V Bias Image | 2470
3298 | SE2 Detector | Detector Alignment Check | 2471
3298 | SE2 Detector | Collector Bias Check | 2472
3298 | SE2 Detector | Function Test | 2473
10702 | Vacuum PCB | All Connections Are Secured with Screws | 2474
10702 | Vacuum PCB | Cables Are Neatly Routed and Secured | 2475
10703 | EO PCB | All Connections Are Secured with Screws | 2476
10703 | EO PCB | Cables Are Neatly Routed and Secured | 2477
10704 | Stage PCB | All Connections Are Secured with Screws | 2478
10704 | Stage PCB | Cables Are Neatly Routed and Secured | 2479
10705 | ASAP | All Connections Are Secured with Screws | 2480
10705 | ASAP | Cables Are Neatly Routed and Secured | 2481
949 | HTBSD | Fit the HT-BSD | 2482
949 | HTBSD | Software setup of the HTBSD | 2483
949 | HTBSD | Test the HTBSD | 2484
949 | HTBSD | X-ray check HT-BSD | 2485
949 | HTBSD | System vac in spec (<2.0e-06mbar) | 2486
949 | HTBSD | Remove detector before shipping | 2487
924 | EFT | Fit to Stage Port | 2488
917 | Volutome | 1. Unpack Volutome and check all parts on the packing list are present (packing list link can be found in attachments) | 2489
917 | Volutome | 1. Upload data from USB stick to the O drive product folder | 2490
917 | Volutome | 1. Remove transit clamps from Volutome stage | 2491
917 | Volutome | 1. Ensure both safety switches are fitted to the Volutome door | 2492
917 | Volutome | 2. Ensure hardware has been configured according to work instructions | 2493
917 | Volutome | 2. Fit Volutome stage | 2494
917 | Volutome | 2. Check protection circuit for Piezo is fitted (349567-9013-000) | 2495
917 | Volutome | 2. Connect Volutome stage and controller to the SEM | 2496
917 | Volutome | 2. Fit Stemi holder bracket to Volutome stage door (STEMI adjustment bracket needs to fit / disconnect without force from door) | 2497
917 | Volutome | 3. Ethernet Connection Setup | 2498
917 | Volutome | 4. Install the latest version of Zen - Ensure to configure it for Volutome (Refer to Zen installation WI) | 2499
917 | Volutome | 4. Ensure Athena inspector manager has all green lights, Physik microtome and Physik stage are present - save a screenshot in the Volutome folder | 2500
917 | Volutome | 4. Install PI Micromove | 2501
917 | Volutome | 4. Run getversions.exe and save text file in the Volutome folder | 2502
917 | Volutome | 4. Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator | 2503
917 | Volutome | 4. Set up the Volutome workbench in Zen | 2504
917 | Volutome | 5. PID parameters configured | 2505
917 | Volutome | 6. Record VolumeBSD serial number | 2506
917 | Volutome | 6. Record Volutome stage serial number | 2507
917 | Volutome | 6. Record VolumeBSD amplifier serial number | 2508
917 | Volutome | 6. Record VolumeBSD mechanics serial number | 2509
917 | Volutome | 6. Record C-885 PI controller serial number | 2510
917 | Volutome | 6. Record E-727 PI controller serial number | 2511
917 | Volutome | 7. Volutome 'internal LED light' function check | 2512
917 | Volutome | 8. Volutome 'External cold light' function check | 2513
917 | Volutome | 9. Adjust Volutome stage centre (Crosshairs should be within 4 central squares) | 2514
917 | Volutome | 10. Install Nav cam in-line connections (If applicable) | 2515
917 | Volutome | 11. Perform the microtome and stage functionality checks | 2516
917 | Volutome | 11. Sigma cartesian stage angle correction correct | 2517
917 | Volutome | 12. Fit and align the VolumeBSD diode | 2518
917 | Volutome | 12. Check VolumeBSD alignment repeatability | 2519
917 | Volutome | 12. Check all gain ranges of the VolumeBSD (normal/inverted) | 2520
917 | Volutome | 12. Perform Signal/noise ratio check | 2521
917 | Volutome | 12. Ensure BSD diode protector is fitted and adjusted in chamber | 2522
917 | Volutome | 12. Install cable tie to loop BSD signal cable | 2523
917 | Volutome | 12. Test BSD functionality in Zen - Insert, retract and gain ranges all functional | 2524
917 | Volutome | 13. Check SCM with Volutome stage | 2525
917 | Volutome | 14. Oscillation check - Check there is no visible oscillation when imaging with a pixel size of 5nm | 2526
917 | Volutome | 15. Perform X-ray check | 2527
917 | Volutome | 16. Pump Down Time | 2528
917 | Volutome | 16. Ultimate Vacuum | 2529
917 | Volutome | 17. Focal CC fitted to Volutome stage | 2530
917 | Volutome | 17. Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar | 2531
917 | Volutome | 17. Remove carbon needle from focal CC after testing and place in the engineers toolbox container | 2532
917 | Volutome | 18. Generate live image from Axiocam | 2533
917 | Volutome | 19. Dummy experiment completed with no erros reported in Zen - Results saved in Volutome folder | 2534
917 | Volutome | 20. Re-package Volutome and check it's completeness | 2535
917 | Volutome | 21. Check functionality of all additional software installed e.g Atlas | 2536
917 | Volutome | 22. Ensure Arivis Software Licence is available (1 year or permanent) | 2537
917 | Volutome | 23. ManEng/ Senior Eng/ Production Manger Sign off | 2538
783 | VPBSD1 | Setup | 2539
783 | VPBSD1 | Calibration | 2540
783 | VPBSD1 | Fit beamsleeve funnel | 2541
783 | VPBSD1 | Mechanical alignment of beamsleeve | 2542
783 | VPBSD1 | set rise and fall height of beamsleeve | 2543
783 | VPBSD1 | Segment test | 2544
783 | VPBSD1 | Diode mirror test | 2545
783 | VPBSD1 | Removing diode contamination with plasma cleaner | 2546
212 | Xplore 30EDS | Language Set to English (United States | 2547
212 | Xplore 30EDS | Ethernet Connection Setup | 2548
212 | Xplore 30EDS | Install Aztec Software | 2549
212 | Xplore 30EDS | AZtec Software Version | 2550
212 | Xplore 30EDS | Install Licences | 2551
212 | Xplore 30EDS | Run AZtec Assist+ | 2552
212 | Xplore 30EDS | Connected to Test Tool Module | 2553
115 | Oxford EDS Detector Installation and Test | Setup Ethernet Connection | 2554
115 | Oxford EDS Detector Installation and Test | Install Aztec Software | 2555
115 | Oxford EDS Detector Installation and Test | Aztec Software Version | 2556
115 | Oxford EDS Detector Installation and Test | Licence Installation | 2557
115 | Oxford EDS Detector Installation and Test | Run Aztec Assist+ | 2558
115 | Oxford EDS Detector Installation and Test | Install Detector | 2559
115 | Oxford EDS Detector Installation and Test | Imaging Setup | 2560
115 | Oxford EDS Detector Installation and Test | Stage Calibration Setup | 2561
115 | Oxford EDS Detector Installation and Test | Calibrate Detector | 2562
115 | Oxford EDS Detector Installation and Test | Save Aztec Sign Off PDF and Calibration Folder | 2563
67 | ZEN Automated Imaging | Install ZEN core EM.Exe | 2564
67 | ZEN Automated Imaging | ZEN core Software Version | 2565
67 | ZEN Automated Imaging | Setup Exchange Position | 2566
67 | ZEN Automated Imaging | Capture Image | 2567
89 | Raman | Test Valves Using Raman Test Interlock Box | 2568
89 | Raman | Install SemRamanStage Software | 2569
89 | Raman | Run SEM Raman Correlation | 2570
89 | Raman | Save Calibration File as Raman | 2571
89 | Raman | second chamberscope fitted | 2572
89 | Raman | Run Smart Stage Mapping | 2573
3306 | STEM / ASTEM | Check Segments | 2574
3306 | STEM / ASTEM | Check all Gain Ranges | 2575
3306 | STEM / ASTEM | Proof Images | 2576
1122 | Quality sign off. | Quality Sign Off Carried Out By: | 2577
10879 | SmartTouch | SmartTouch Software Version | 2578
10879 | SmartTouch | Install SmartTouch Software | 2579
10879 | SmartTouch | Install Easy VP Aperture | 2580
10879 | SmartTouch | Capture Sample Image Using the Nav Cam | 2581
10879 | SmartTouch | Select Sample Holder and Sample Type | 2582
10879 | SmartTouch | Set Yellow Line To 8.5mm WD | 2583
10879 | SmartTouch | Sample Setup Complete | 2584
10879 | SmartTouch | In the Image Tab, Check Sliders Function | 2585
10879 | SmartTouch | Save Image - Photomode 1, Scan speed 8 | 2586
10879 | SmartTouch | Save Image - Photomode 2, Single Image | 2587
10879 | SmartTouch | Save Image - Photomode 3, Scan Speed 2, Multiple Frames | 2588
10879 | SmartTouch | Check Auto Saturation (W Filament Only) | 2589
10879 | SmartTouch | Check Auto Stigmation | 2590
10879 | SmartTouch | Check Manual Stigmation | 2591
10879 | SmartTouch | Check Annotate | 2592
10879 | SmartTouch | Check All Detectors Are Presaent | 2593
10879 | SmartTouch | Select the Automate tab & Create Project as Zeiss | 2594
10879 | SmartTouch | Setup 3 Point Calibration (Nav Cam ) | 2595
10879 | SmartTouch | Automation Run Completed | 2596
10879 | SmartTouch | Select Browse Tab and Check Functions | 2597
10870 | Camera Stand | Camera Model | 2598
10870 | Camera Stand | Install Camera Software | 2599
3299 | Chamberscope | Edit Chamberscope Macro | 2600
3299 | Chamberscope | Set Chamberscope Focus | 2601
3299 | Chamberscope | Chamberscope 2 - Select Port or Not Fitted | 2602
3299 | Chamberscope | Chamberscope 2 - Check Camera 2 for Focus | 2603
3299 | Chamberscope | Chamberscope 2 - Functional Check-Black/White + Colour Mode | 2604
3299 | Chamberscope | Chamberscope 2 - Run wxPropview to Remove Defective Pixels | 2605
3299 | Chamberscope | Chamberscope 2 - Add Icon to Toolbar | 2606
3300 | Styled Joystick | Function Test | 2607
3300 | Styled Joystick | Check Break Button | 2608
217 | Joystick | Function Test | 2609
3301 | Control Panel | Function Test | 2610
3301 | Control Panel | Edit Exchange/Resume Macros | 2611
3302 | Plasma Cleaner | Run ‘Quick Sample Clean’ | 2612
3303 | Quiet Mode | Function Test | 2613
3303 | Quiet Mode | Disable Confirmation Check | 2614
3303 | Quiet Mode | Turbo Load Check | 2615
3303 | Quiet Mode | Save .CSV of quite mode cycle time | 2616
3304 | BSD | BSD Diode Pre-test | 2617
3304 | BSD | Check each individual segment in normal and inverse modes | 2618
3304 | BSD | Analyze Segments | 2619
3304 | BSD | Gain Ranges | 2620
3305 | ASB | BSD Diode Pre-test | 2621
3305 | ASB | Check each individual segment in normal and inverse modes | 2622
3305 | ASB | Analyze Segments | 2623
3305 | ASB | Gain Ranges | 2624
11074 | BSD 1 /BSD 4 / Sense-BSD/VP BSD | Check each individual segment in normal and inverse modes | 2625
11074 | BSD 1 /BSD 4 / Sense-BSD/VP BSD | Analyze Segments | 2626
11074 | BSD 1 /BSD 4 / Sense-BSD/VP BSD | Gain Ranges | 2627
11074 | BSD 1 /BSD 4 / Sense-BSD/VP BSD | Pneumatics | 2628
11074 | BSD 1 /BSD 4 / Sense-BSD/VP BSD | Alignment | 2629
11074 | BSD 1 /BSD 4 / Sense-BSD/VP BSD | Vac Degradation < 3x10-5 mbar | 2630
11074 | BSD 1 /BSD 4 / Sense-BSD/VP BSD | CZ Detector Cal | 2631
11074 | BSD 1 /BSD 4 / Sense-BSD/VP BSD | Z Adjustment | 2632
10706 | Plinth | Cables Are Neatly Routed and Secured | 2633
10706 | Plinth | Plinth Is free From Debris | 2634
10706 | Plinth | Table Top Is Free From Damage | 2635
10706 | Plinth | Table Top Is Secured | 2636
10706 | Plinth | Clean and Tidy Assembly Bay | 2637
762 | Material Availability | Are there any shortages? | 2638
10750 | Nav Cam | SPP 360 Completed | 2639
10485 | Overnight Stage Macro ( DO Not Do On Hittech Stages) | Minimum of 8 Hours (Cartesian Stage) | 2640
1328 | SCS Check | Check SCS Folder/Sales Order | 2641
1328 | SCS Check | SCS Number (If applicable) | 2642
10927 | YAG BSD | Install YAG BSD Lightpipe | 2643
10927 | YAG BSD | Alignment | 2644
10927 | YAG BSD | Select YAG BSD in the Administrator | 2645
10927 | YAG BSD | 30KV Function Image | 2646
10927 | YAG BSD | Obtain a Image at 3KV | 2647
3307 | Airlock | Set Rod Retract Switch | 2648
3307 | Airlock | Edit Airlock Macro | 2649
3307 | Airlock | Set Up Stage Exchange | 2650
3307 | Airlock | Airlock Function Test | 2651
11011 | EBBU | Function Test | 2652
11012 | EBBP | Fit Test tool BB Electronis | 2653
11012 | EBBP | Function Test | 2654
11012 | EBBP | Remove Test tool BB Electronics | 2655
11347 | Gatan 3 View Unpacking | Unpack Gatan + check against packing list (Photograph each layer before removing) | 2656
11347 | Gatan 3 View Unpacking | Set up stage disable Macro in SmartSEM + assign icon on toolbar | 2657
11347 | Gatan 3 View Unpacking | Remove transit clamps from Gatan stage and fit tie rods | 2658
11347 | Gatan 3 View Unpacking | Adjust stage rails so both Klein and Gatan stages can pump without assistance | 2659
11347 | Gatan 3 View Unpacking | If Nav Cam fitted please contact Peter Lewis to have in – line connection made up | 2660
11348 | Gatan 3 View Setup | Plasma clean Gatan stage (20W 2Hrs) | 2661
11348 | Gatan 3 View Setup | Fit On Point detector and diode | 2662
11348 | Gatan 3 View Setup | Remove Mc Affee software from Gatan PC | 2663
11348 | Gatan 3 View Setup | Disable Firewalls on Gatan and Zeiss PC | 2664
11348 | Gatan 3 View Setup | Install and configure remote client on Gatan PC | 2665
11348 | Gatan 3 View Setup | Configure ‘R configure‘ server on Zeiss PC | 2666
11348 | Gatan 3 View Setup | Ping IP adresses | 2667
11348 | Gatan 3 View Setup | Configure Gatan settings | 2668
11349 | Gatan 3 View Testing | Align BSD and obtain image | 2669
11349 | Gatan 3 View Testing | Check SCM with Gatan Stage | 2670
11349 | Gatan 3 View Testing | Ripple X Spec 3nm (600Kx Mag 15KV WD2) | 2671
11349 | Gatan 3 View Testing | Ripple Y Spec 3nm (600Kx Mag 15KV WD2) | 2672
11349 | Gatan 3 View Testing | Resolution 1KV Spec 2.1nm | 2673
11349 | Gatan 3 View Testing | Resolution 15KV Spec 1.2nm | 2674
11349 | Gatan 3 View Testing | Contamination | 2675
11349 | Gatan 3 View Testing | Pump Down Time | 2676
11349 | Gatan 3 View Testing | Ultimate Vacuum | 2677
11349 | Gatan 3 View Testing | Focal CC fitted to 3View stage | 2678
11349 | Gatan 3 View Testing | Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar | 2679
11349 | Gatan 3 View Testing | Charging paper test carried out and images saved. | 2680
11350 | Gatan 3 View Packing | Assemble Gatan vessel pumping accessories. | 2681
11350 | Gatan 3 View Packing | Re-package system | 2682
11350 | Gatan 3 View Packing | ManEng/ Senior Eng/ Production Manger Sign off | 2683
3349 | System vacuum (Chamber vacuum) | Reading (< 3.00e-06 mbar) = | 2684
3309 | Atlas 5 | Connect Silver Box and Cables | 2685
3309 | Atlas 5 | Download latest Version Of Software (see link in docs) | 2686
3309 | Atlas 5 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info | 2687
3309 | Atlas 5 | Request Licence File from Fibics | 2688
3309 | Atlas 5 | Install Atlas Software | 2689
3309 | Atlas 5 | Install Licence File | 2690
3309 | Atlas 5 | Calibrate SEM | 2691
3309 | Atlas 5 | Obtain Mosaic | 2692
3309 | Atlas 5 | Set ATLAS Rotation Offset | 2693
3309 | Atlas 5 | Save ATLAS Licence Document | 2694
3313 | Smart Stitch | Installation | 2695
3313 | Smart Stitch | Functional Test | 2696
3313 | Smart Stitch | Store SmartStitch Functional Image | 2697
3313 | Smart Stitch | SmartStitch Folder is present on Desktop | 2698
3314 | SmartEDX | Type of Detector | 2699
3314 | SmartEDX | Detector Serial number | 2700
3314 | SmartEDX | DDP Module Serial Number | 2701
3314 | SmartEDX | Fit Flange & Mount Detector | 2702
3314 | SmartEDX | Install EDS PCBs into PC | 2703
3314 | SmartEDX | Installation of DPP Module & Cables | 2704
3314 | SmartEDX | Configure PC Network & Test | 2705
3314 | SmartEDX | Disable Windows10 Fast Startup | 2706
3314 | SmartEDX | Install ApexKey.edx Licences | 2707
3314 | SmartEDX | Install Genesis.edx Licences (if Applicable) | 2708
3314 | SmartEDX | Fit Dongle | 2709
3314 | SmartEDX | APEX Software Version | 2710
3314 | SmartEDX | Install APEX Software | 2711
3314 | SmartEDX | Install Remoting | 2712
3314 | SmartEDX | Setup RConfigure | 2713
3314 | SmartEDX | Run API Test | 2714
3314 | SmartEDX | Configure imaging – Run SGShell | 2715
3314 | SmartEDX | Calibrate Detector – Run eDPPShell | 2716
3314 | SmartEDX | Save Calibration Report | 2717
3314 | SmartEDX | Detector Resolution - ≥ 129 eV | 2718
3314 | SmartEDX | Save calibration data on exit | 2719
3314 | SmartEDX | Run APEX | 2720
3314 | SmartEDX | Check Spectrum | 2721
3314 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected | 2722
3314 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected | 2723
3314 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected | 2724
3314 | SmartEDX | Check Mapping | 2725
3314 | SmartEDX | Check Line Scanning | 2726
3314 | SmartEDX | Save Aquired Spectrum, Mapping and Line Scanning Images to Image Folder | 2727
3314 | SmartEDX | Save Calibration Data to the Licence Folderon the SmartSem USB stick. | 2728
3314 | SmartEDX | Save Calibration Data to the Image Folder | 2729
10873 | Bruker EDS Detector Install | Number of Detectors | 2730
10873 | Bruker EDS Detector Install | Detector Serial Number(s) | 2731
10873 | Bruker EDS Detector Install | Detector Resolution | 2732
10873 | Bruker EDS Detector Install | Detector Area | 2733
10873 | Bruker EDS Detector Install | Detector Installed | 2734
10873 | Bruker EDS Detector Install | In/Out Stops Set | 2735
10873 | Bruker EDS Detector Install | Isolation Check of Detector & Column >2MΩ | 2736
10873 | Bruker EDS Detector Install | Install Remoting & Setup Network Connection | 2737
10873 | Bruker EDS Detector Install | Bruker Esprit Software Version | 2738
10873 | Bruker EDS Detector Install | Install Bruker Esprit Software | 2739
10873 | Bruker EDS Detector Install | Bruker EDS Manager Version | 2740
10873 | Bruker EDS Detector Install | Install EDS Manager | 2741
10873 | Bruker EDS Detector Install | Install Bruker Licences | 2742
10873 | Bruker EDS Detector Install | Setup Pulse Processor | 2743
10873 | Bruker EDS Detector Install | Check Standard Timing Cards Fitted | 2744
10873 | Bruker EDS Detector Install | List Optional Timing Cards Fitted | 2745
10873 | Bruker EDS Detector Install | Install Detector Data | 2746
10873 | Bruker EDS Detector Install | Check Microscope driver | 2747
10873 | Bruker EDS Detector Install | Setup SEM Conditions | 2748
10873 | Bruker EDS Detector Install | Determine Analytical Working Distance | 2749
10873 | Bruker EDS Detector Install | Calibrate Detector | 2750
10873 | Bruker EDS Detector Install | Calibrate Image in Espirit | 2751
3350 | Gun vacuum (UHV vacuum) | Reading (< 5.00e-09mbar) = | 2752
3351 | VP-operation | Chamber vacuum can be changed continuously between 10 and 133 PA (Beam On) | 2753
214 | Photograph Oxford Sales Number Label | Take Photo of Label | 2754
214 | Photograph Oxford Sales Number Label | Oxford Sales Order | 2755
3316 | SmartPI | Correct PC - 347823-9024-660 Installed | 2756
3316 | SmartPI | Print Out SmartPI Check List (Link in Documents) | 2757
3316 | SmartPI | Install SmartSEM Licences | 2758
3316 | SmartPI | Install MIL Redistribution | 2759
3316 | SmartPI | Fit MIL Dongle | 2760
3316 | SmartPI | SmartPI Software Version | 2761
3316 | SmartPI | Install SmartPI Software | 2762
3316 | SmartPI | Calibrate I/O-SCAN card | 2763
3316 | SmartPI | Setup & Run Application | 2764
3316 | SmartPI | Check Results & Image Montage | 2765
10707 | Turbo Pump Assembly | Turbo Stack Assembled | 2766
10707 | Turbo Pump Assembly | Turbo M8 Bolts Torqued to 7Nm | 2767
10710 | Fit Chamber Assembly | Check Stage Rail Grub Screws are Tight | 2768
10710 | Fit Chamber Assembly | Torque Chamber Bolts to 35Nm | 2769
10710 | Fit Chamber Assembly | Record Chamber SN | 2770
10710 | Fit Chamber Assembly | Check Bellows is concentric to the hole in the desktop. | 2771
10711 | Clean Chamber - CARBOFLO | Internal Surfaces | 2772
10711 | Clean Chamber - CARBOFLO | Tapped Holes | 2773
11405 | Fit Electrical Feed Through | Select Number of Ways | 2774
10754 | System Cabling | All Cables Are Neatly Routed | 2775
10754 | System Cabling | All Plugs Are Secured Correctly | 2776
10754 | System Cabling | All Earth Cables Are Tightly Secured | 2777
10874 | Mineralogic | Print Out Mineralogic Check List (Link in Documents) | 2778
10874 | Mineralogic | Correct PC - Installed | 2779
10874 | Mineralogic | Install SmartSEM Licences | 2780
10874 | Mineralogic | Instal MIL Redistribution | 2781
10874 | Mineralogic | Fit MIL Dongle | 2782
10874 | Mineralogic | Install Mineralogic Software | 2783
10874 | Mineralogic | Install MAC Standard Recipe Files | 2784
10874 | Mineralogic | Calibrate I/O-SCAN card | 2785
10874 | Mineralogic | MAC Standard & Stage Postion Set | 2786
10874 | Mineralogic | Adjust Recipe for the System | 2787
10874 | Mineralogic | Run 20 min Rapid MAC STD Measure | 2788
10874 | Mineralogic | Check Results & Image Montage - 20 min Rapid MAC STD Measure | 2789
10874 | Mineralogic | Run 90 min High Mag Rapid MAC STD Measure | 2790
10874 | Mineralogic | Check Results and Compare with 20 min Run | 2791
10874 | Mineralogic | 10 Hour Overnight Run | 2792
10874 | Mineralogic | Check Results and Compare with 20 and 90 min Run | 2793
10883 | Zen | ZEN Version | 2794
10883 | Zen | Correct Version of Windows 10 Installed | 2795
10883 | Zen | Install ZEN | 2796
10883 | Zen | Install Licence | 2797
10883 | Zen | Install Remote API Client | 2798
10883 | Zen | Functional Test | 2799
10883 | Zen | Save Snapped Image | 2800
10883 | Zen | Store Snapped Image | 2801
10883 | Zen | Store Entitlement Certificate | 2802
114 | C2D/C2D F2 | Black Level Cal | 2803
114 | C2D/C2D F2 | C2D Bias Target Check | 2804
114 | C2D/C2D F2 | Gain Range Check | 2805
114 | C2D/C2D F2 | Auto Level Check | 2806
114 | C2D/C2D F2 | Auto Range Check | 2807
114 | C2D/C2D F2 | Function Image | 2808
114 | C2D/C2D F2 | Store C2D Functional Image | 2809
3317 | System Vacuum Check | Pressure (<3.00e-6mbar) | 2810
11102 | SCS Complete | Check SCS Folder/Sales Order | 2811
3318 | Final Gun Vacuum Check | Pressure (<5.00e-9 mbar) | 2812
3319 | System Pump Down Time | Time (<5 min) | 2813
3320 | Copy into support Folder | 1KV Resolution Image | 2814
3320 | Copy into support Folder | 15KV Resolution Image | 2815
3320 | Copy into support Folder | Licence File | 2816
3320 | Copy into support Folder | EM Server ACCDB | 2817
3320 | Copy into support Folder | EM Server Screening .TXT | 2818
3321 | Install Customer Licences | Delete the Production Test Licences | 2819
3321 | Install Customer Licences | Copy EMServer.db to the C: Support/Data/Log Files folder | 2820
3321 | Install Customer Licences | Obtain the customer Entitlement Certificate | 2821
3321 | Install Customer Licences | Download into the C:Support/Data/Licence folder | 2822
3321 | Install Customer Licences | Print out the Customer Entitlement Certificate - To be sent to the customer | 2823
3321 | Install Customer Licences | Load customer licences and check against order | 2824
3321 | Install Customer Licences | Test customer licenses | 2825
3326 | Final Logbook checks: Consistency of data | All verification & Safety tests done and within threshold limits defined | 2826
3326 | Final Logbook checks: Consistency of data | Sales Order is latest version | 2827
3326 | Final Logbook checks: Consistency of data | All items on Sales order accounted for | 2828
3327 | Copy files to server | Em Server.accdb | 2829
3327 | Copy files to server | Em Server.txt Log | 2830
3327 | Copy files to server | Optiprobe.txt Log | 2831
3327 | Copy files to server | Em Server Screening.txt | 2832
3327 | Copy files to server | All Images | 2833
3328 | Dual Channel | Two detectors can be used simultaneously | 2834
3329 | Dual Mag | Two Mags can be achieved simultaneously | 2835
3330 | Centre Feature | Feature can be centred at high and low mag | 2836
3331 | Fisheye Mode | Ensure mode works without distortions or rotation. | 2837
3332 | Advanced Measurement | 5+ annotation | 2838
3333 | Gamma + LUT | Able to edit LUT | 2839
3334 | Optiprobe | Optiprobe can be enabled and specimen current varied | 2840
3335 | Smart Browse | Application is installed and can be opened | 2841
3336 | Plasma Cleaner | Plasma Clean can be enabled | 2842
3342 | Cladding | No pipes are touching cladding. | 2843
3342 | Cladding | Cladding is secured correctly and not able to move. | 2844
3342 | Cladding | Safety stickers are applied correctly and are the correct language. | 2845
6934 | Specs | PC S/N | 2846
6934 | Specs | SmartSEM Version | 2847
6934 | Specs | Confirm Stage, EO and Vac FW versions are latest release | 2848
3344 | Hard Drives | Images in TIFF format can be saved on disk and Hard disk and reloaded without loss of information | 2849
3345 | Operating system / Smartsem | EHT set remains powered on yellow button | 2850
3345 | Operating system / Smartsem | Check SmartSEM is in the folder C:Support/Data/SmartSEM Software | 2851
3346 | Customer licenses | Ensure all customer licences are loaded in the Zeiss Licence Management Tool and functions correctly | 2852
3346 | Customer licenses | Check thst the SmartSEM 8 folder is present in the support folder | 2853
1297 | Support Folder - Check Customer Entitlement Certificate and EM Server.db | The Customer Entitlement Certificate is present in C:Support/Data/Licence folder | 2854
1297 | Support Folder - Check Customer Entitlement Certificate and EM Server.db | The EM Server.db is present in C:Support/Data/Log Files folder | 2855
3358 | Probe current | Reading (220 +/- 10pA) | 2856
3358 | Probe current | Confirm Ext V is running <500V over data sheet @220pA | 2857
3358 | Probe current | Beam Blank - Confirm reading goes under-range when beam blanked. | 2858
3359 | Probe current | Reading (300 +/- 10pA) | 2859
3359 | Probe current | Beam Blank - Confirm reading goes under-range when beam blanked. | 2860
10865 | Cathode Values-Actual | Filament heating Current | 2861
10865 | Cathode Values-Actual | Extrator Power at 0.2 mA/sr | 2862
10865 | Cathode Values-Actual | Total Current at 0.2 mA/sr | 2863
3361 | Maximum Specimen Current | Measured Value | 2864
10885 | X-Ray Check | Check That all detectors are listed under and signed off in X-Ray checks, If any extra detectors have been fitted perform X-Ray checks on those detectors | 2865
10885 | X-Ray Check | Sign XRay Certificate | 2866
3363 | InLens-Detector | Confirm no noise can be seen on fastest scan speeds | 2867
3364 | SE-Detector | Collector - Changing the collector bias voltage bias results in a change of the signal intensity | 2868
3364 | SE-Detector | Scintillator - Confirm switching scintillator on/off results in a change of the signal intensity | 2869
1177 | VPBSD | Pneumatics and segments functional | 2870
1177 | VPBSD | SmartSEM Control tested | 2871
1177 | VPBSD | Images saved | 2872
3366 | BSD-Detector | Confirm there is no rotation of the BSD Diode when inserting the rod. | 2873
3366 | BSD-Detector | Confirm is smooth action when rod inserted | 2874
3366 | BSD-Detector | Check each gain ranges and functionality of each segment | 2875
3366 | BSD-Detector | Check quality of BSD images | 2876
11075 | BSD 1, BSD 4 Detector, BSD 1 VP | Pneumatics functional | 2877
11075 | BSD 1, BSD 4 Detector, BSD 1 VP | SmartSEM Control tested | 2878
11075 | BSD 1, BSD 4 Detector, BSD 1 VP | Images saved | 2879
10911 | C2D Detector | C2D Is Selected When Going Into VP Mode | 2880
10911 | C2D Detector | Image Brightness Changes When Changing Bias | 2881
10911 | C2D Detector | Breakdown Of Bias Occurs at 100% | 2882
10914 | InLensDuo | Confirm a signal intesity change can be seen when changing bias voltage. | 2883
10914 | InLensDuo | No breakdown of the detector can be seen at max bias voltage | 2884
10912 | Quiet Mode | Check correct operation of Quiet Mode. Deselect Quiet Mode and ensure rotary pump starts | 2885
10912 | Quiet Mode | Load previously saved .CSV file to confirm acceptable cycle times | 2886
3369 | CCTV-Camera | Focused chamber-scope image can be obtained at predefined contrast and brightness | 2887
3370 | CCTV-Camera 2 | Focused chamber-scope image can be obtained at predefined contrast and brightness | 2888
10919 | Nav Cam | Stage Moves To Correct Position When Arm Is Opened | 2889
10919 | Nav Cam | Image Is Captured | 2890
10919 | Nav Cam | Functional Check | 2891
3374 | VIDEO INPUT | Video Input Kit Fitted | 2892
3376 | Control Panel | All controls function correctly | 2893
3376 | Control Panel | Ensure keyboard language is correct | 2894
3377 | EBBU | Beam is blanked when the EBBU is on | 2895
3377 | EBBU | Remove EBBU Electronics and fit Blanking Cap before shipping | 2896
10909 | Camera Stand | Software Loaded | 2897
10910 | Additional Hard Disk & Cloning Software | Hard Disk Is Seen In File Manager | 2898
10910 | Additional Hard Disk & Cloning Software | Cloning Software Installed | 2899
10913 | SmartTouch | Run Up & Obtain An Image | 2900
10913 | SmartTouch | Confirm Image Stitching is acceptable | 2901
10915 | Plasma cleaner | Run quick sample clean and confirm no errors | 2902
10916 | Airlock | Sample can load without any issues | 2903
10916 | Airlock | Exchange and Resume Macros function correctly from conrol panel / airlock panel | 2904
10918 | ATLAS 5 | Atlas Hardware Installed | 2905
10918 | ATLAS 5 | Atlas Software Installed | 2906
10918 | ATLAS 5 | Atlas Licence Saved To SmartSEM USB Stick | 2907
10921 | SmartEDX | Check Aquire Spectrum image | 2908
10921 | SmartEDX | Check Line Scanning | 2909
10921 | SmartEDX | Check Mapping | 2910
10921 | SmartEDX | Check EDX Icon Is Present In All User Toolbars | 2911
10921 | SmartEDX | EDAX Calibration File Saved To SmartSEM USB Stick | 2912
10922 | 3DSM | Load project image and confirm ok | 2913
10923 | EDS Detector | Confirm oxford software is loaded | 2914
10923 | EDS Detector | Check Licenses | 2915
10924 | SmartPI | SmartPI Software Installed and Tested | 2916
10925 | Mineralogic | Mineralogic Software Installed and Tested | 2917
3380 | Stigmator | Stigmator value X (< 30%) | 2918
3380 | Stigmator | Stigmator value Y (< 30%) | 2919
1180 | Run Calibration Extraction Untility | Download the CEU folder and copy to the Desktop. | 2920
1180 | Run Calibration Extraction Untility | Install RClientinstaller.exe (if not already installed) | 2921
1180 | Run Calibration Extraction Untility | Run SEM_Calibration Extraction Tool. | 2922
1180 | Run Calibration Extraction Untility | Store the SEM_CEU.txt and Sigma***-**-**.csv files | 2923
1180 | Run Calibration Extraction Untility | copy ceu files to the C:Support/Data/Log files Folder. | 2924
1180 | Run Calibration Extraction Untility | If the RClientinstaller has been installed for the the CEU only, then uninstall | 2925
1298 | Copy the Support folder to the (D:) Drive | Copy the Support folder and its sub folders to the (D:) drive on the PC | 2926
1298 | Copy the Support folder to the (D:) Drive | Verify that all the folders and their contents have been copied | 2927
3396 | SCS Check | Special Fitted | 2928
3396 | SCS Check | If Fitted, Checked by SCS Team Member | 2929
3397 | Update Server Files | Em Server.mdb | 2930
3397 | Update Server Files | Em Server.txt Log | 2931
3397 | Update Server Files | Optiprobe.txt Log | 2932
3397 | Update Server Files | All Images | 2933
3398 | PC Clean Up | Remove test macro files | 2934
3398 | PC Clean Up | Delete Contents of Log Folder | 2935
3398 | PC Clean Up | Download Folder Images Deleted | 2936
3398 | PC Clean Up | Remove images from image gallery | 2937
3398 | PC Clean Up | Empty Recycle Bin | 2938
3398 | PC Clean Up | Clear Internet History | 2939
892 | Remove MAC Address | Remove MAC address (Internet) | 2940
892 | Remove MAC Address | Remove MAC address (MIC-NET) | 2941
931 | Earth Continuity Test | PAT Tester Serial Number | 2942
931 | Earth Continuity Test | PAT Tester Cal Expiry Date | 2943
931 | Earth Continuity Test | Earth resistance – measured value must be < 0.1 Ω | 2944
931 | Earth Continuity Test | Insulation resistance – measured value must be > 1.0M Ω | 2945
931 | Earth Continuity Test | Test Certificate Completed & Signed | 2946
969 | Fit Front Panel Logo | Clean Front Panel | 2947
969 | Fit Front Panel Logo | Prepare Logo | 2948
969 | Fit Front Panel Logo | Fitting Logo | 2949
3402 | Change Customer Language Option | Set Customer Language | 2950
3403 | Fit transit brackets | Fit foam between column cladding and chamber (356581-5085-000 can be used and modified as needed) | 2951
3403 | Fit transit brackets | Fit LH Bracket | 2952
3403 | Fit transit brackets | Fit RH Bracket | 2953
3403 | Fit transit brackets | Fit Rear Bracket | 2954
3403 | Fit transit brackets | Fit Front Bracket | 2955
3403 | Fit transit brackets | Checked By Name: | 2956
3405 | Copy Customer Information Sheets | X-Ray Leakage Check | 2957
3405 | Copy Customer Information Sheets | Quality Certificate | 2958
3405 | Copy Customer Information Sheets | Customer Entitlement Certificate | 2959
3405 | Copy Customer Information Sheets | Gun/Aperture Align Sheet | 2960
3405 | Copy Customer Information Sheets | Denka Data Sheet | 2961
3405 | Copy Customer Information Sheets | Test Accessories Packing list | 2962
3405 | Copy Customer Information Sheets | Standard Packing List | 2963
3405 | Copy Customer Information Sheets | Items retained on Microscope Packing List | 2964
3405 | Copy Customer Information Sheets | Items removed from Microscope Packing List | 2965
3405 | Copy Customer Information Sheets | China RoHS Certificate | 2966
194 | Check 2 Off Lindy Mini DP to DP cables Supplied - | Lindy Cables Supplied | 2967
3411 | Check Software Licences | Licensed software on Sales Order matches Customer Entitlement Certificate | 2968
10660 | Record Power Supply S/No | Serial Number | 2969
10662 | Pre Assembly QA | Check All Crimps Are Correctly Located And Secure. | 2970
10662 | Pre Assembly QA | Test PSU Ensure All LED’s Light Up Correctly | 2971
10662 | Pre Assembly QA | To be Checked By Another Engineer - Name: | 2972
1111 | Sigma Cartesian Travel | Confirm X Travel (0 - 125mm) | 2973
1111 | Sigma Cartesian Travel | Confirm Y Travel (0 - 125mm) | 2974
1111 | Sigma Cartesian Travel | Confirm Z Travel (0 - 50mm) | 2975
1111 | Sigma Cartesian Travel | Check Tilt Travel -10° to +90° | 2976
1073 | Confirm 4-eye check by additional operator | Checked Carried Our By: | 2977
1079 | Wiring, Micro switches and Cable form Fitment. (Set Micro-switch Positions) | Z Microswitch Position Set | 2978
1079 | Wiring, Micro switches and Cable form Fitment. (Set Micro-switch Positions) | T Microswitch Position Set | 2979
1079 | Wiring, Micro switches and Cable form Fitment. (Set Micro-switch Positions) | R Microswitch Position Set | 2980
1082 | Confirm 4-eye check by additional operator | Checked Carried Our By: | 2981
1095 | Confirm 4-eye check by additional operator | Checked Carried Our By: | 2982
1130 | Fit ZTR Motors and record SN in the box provided below: | Z Motor Serial Number: | 2983
1130 | Fit ZTR Motors and record SN in the box provided below: | T Motor Serial Number: | 2984
1130 | Fit ZTR Motors and record SN in the box provided below: | R Motor Serial Number: | 2985
1107 | Confirm 4-eye check by additional operator | Checked Carried Our By: | 2986
1109 | Micro Switch Checks | Check Set Point X Axis Micro Switch | 2987
1109 | Micro Switch Checks | Check Set Point Y Axis Micro Switch | 2988
1109 | Micro Switch Checks | Check Set Point Z Axis Micro Switch | 2989
1109 | Micro Switch Checks | Check Set point T Axis (-10.8º) | 2990
1110 | Stage Initialisation | Check Stage Initialises Correctly | 2991
1121 | Stage Type & Serial Number | Stage Serial Number: | 2992
11117 | SCS Check | Check SCS Folder/Sales Order | 2993
11117 | SCS Check | SCS Number (If applicable) | 2994
10663 | Inspect Damper Plate | Damper Plate is Sigma Version (Additional airlock hole) | 2995
10663 | Inspect Damper Plate | Damper Plate is Located Correctly (Centred) | 2996
10664 | Inspect Table Top | Table Top is Correct Version (Airlock punchout hole present) | 2997
10664 | Inspect Table Top | Table Top is Free from Damage | 2998
10666 | Fit Mains Box & Turbo Controller | Fit Turbo Controller To Mains Box | 2999
10666 | Fit Mains Box & Turbo Controller | Turbo Controller Serial Number | 3000
10666 | Fit Mains Box & Turbo Controller | Fit Mains Box | 3001
10666 | Fit Mains Box & Turbo Controller | Add connector block to earth | 3002
10668 | Fit VAC PCB | VAC PCB S/No | 3003
10668 | Fit VAC PCB | All Connections Are Secured with Screws | 3004
10668 | Fit VAC PCB | Cables Are Neatly Routed and Secured | 3005
10669 | Fit H-ASAP | H-ASAP S/No | 3006
10669 | Fit H-ASAP | All Connections Are Secured with Screws | 3007
10669 | Fit H-ASAP | Cables Are Neatly Routed and Secured | 3008
10671 | Fit Stage PCB (Klein) | Stage PCB S/No | 3009
10671 | Fit Stage PCB (Klein) | Fit Stage Light Barrier PCB | 3010
10671 | Fit Stage PCB (Klein) | All Connections Are Secured with Screws | 3011
10671 | Fit Stage PCB (Klein) | Cables Are Neatly Routed and Secured | 3012
1319 | Fit Stage PCB (Cartesian / HiTech | Stage PCB S/No | 3013
1319 | Fit Stage PCB (Cartesian / HiTech | All Connections Are Secured with Screws | 3014
1319 | Fit Stage PCB (Cartesian / HiTech | Cables Are Neatly Routed and Secured | 3015
11105 | Record Plinth S/No | Serial Number | 3016
10674 | Fit additional Cables | Fit D-Sub connectors to back of plinth | 3017
10674 | Fit additional Cables | Ensure stage cable ferrite does not clash when closing plinth door | 3018
10674 | Fit additional Cables | Fit Stage Earth Cable (Leave coiled) | 3019
10674 | Fit additional Cables | Fit Turbo Earth Cable (Leave coiled) | 3020
10675 | Install Main Loom | Specimen current test | 3021
10712 | Clean Chamber - IPA | Internal Surfaces | 3022
10712 | Clean Chamber - IPA | Tapped Holes | 3023
10755 | Detectors and Accessories | Fitted To Correct Ports | 3024
10755 | Detectors and Accessories | Detectors Connected To Correct ASAP Port And Labelled | 3025
10756 | Bay | Clean and tidy Assembly Bay | 3026
1163 | SCS Sign off | Does system have an SCS? please select | 3027
1163 | SCS Sign off | If yes, Product Manager to enter initials. If no please type N/A: | 3028
10773 | Vacuum Test | No Leaks Detected (Prior to running up Turbo) | 3029
10773 | Vacuum Test | No Leaks Detected (With Turbo running) when in high sensitivity | 3030
10772 | Stage Links | Stage Links Set Correctly | 3031
1147 | Fit Stage Assembly (only applicable if Stage not already fitted at TAKT 4) | Stage Serial Number | 3032
1147 | Fit Stage Assembly (only applicable if Stage not already fitted at TAKT 4) | Shim Stage Door | 3033
1147 | Fit Stage Assembly (only applicable if Stage not already fitted at TAKT 4) | Set Stage Clamp Height | 3034
10665 | Fit Damper Plate Metal Work | Earth Plate | 3035
10665 | Fit Damper Plate Metal Work | HV Cable Plate | 3036
10665 | Fit Damper Plate Metal Work | Bellows | 3037
1320 | Plinth Checks | Cables Are Neatly Routed and Secured | 3038
1320 | Plinth Checks | Plinth Is free From Debris | 3039
1320 | Plinth Checks | Table Top Is Free From Damage | 3040
1320 | Plinth Checks | Table Top Is Secured | 3041
1320 | Plinth Checks | EHT Set Is Secured | 3042
1320 | Plinth Checks | Clean and Tidy Assembly Bay | 3043
754 | Material Availability | Are there any shortages? | 3044
10793 | Assemble Turbo Stack | Turbo Stack Assembled | 3045
10793 | Assemble Turbo Stack | Turbo M8 Bolts Torqued to 7Nm | 3046
10808 | Fit Turbo Pump Assembly | Turbo Pump Serial Number | 3047
10794 | Chamber Checks | Chamber Is Free From Paint Chips / Scratches | 3048
10794 | Chamber Checks | Sealing faces are free from defects (Stage sealing face, Intermediate ring) | 3049
10794 | Chamber Checks | Confirm vented screws fitted to chamber | 3050
1321 | Fit Chamber Valves VP | Fit Vent Valve under chamber. | 3051
1321 | Fit Chamber Valves VP | Fit Block Valve. | 3052
1321 | Fit Chamber Valves VP | Fit Pirani elbow | 3053
1321 | Fit Chamber Valves VP | Fit Seconf Vent Valve | 3054
10803 | Fit TIV | TIV Serial Number | 3055
10803 | Fit TIV | Function Check | 3056
10803 | Fit TIV | Fit TIV | 3057
10797 | Fit Chamber Assembly | Check Stage Rail Grub Screws are Tight | 3058
10797 | Fit Chamber Assembly | Fit Blue Gasket | 3059
10797 | Fit Chamber Assembly | Apply Loctit 638 and Torque Chamber Bolts To 10Nm | 3060
10797 | Fit Chamber Assembly | Record Chamber SN | 3061
10798 | Clean Chamber | Clean Internal Chamber Surfaces with Carbaflow and lint free cloth. | 3062
10798 | Clean Chamber | Clean Internal Chamber Surfaces with IPA and lint free cloth. | 3063
1323 | Fit Pirani Gauge | Fit Pirani Gauge | 3064
1323 | Fit Pirani Gauge | Fit pumping pipe (Black valve to turbo) | 3065
11386 | Bake Start | Time = | 3066
11386 | Bake Start | Date = | 3067
638 | Plasma Clean Chamber | Run 3 hour plasma clean | 3068
11388 | Bake Finish | Time = | 3069
11388 | Bake Finish | Date = | 3070
10811 | Vacuum System | All Vacuum Clamps are Tight | 3071
10811 | Vacuum System | All Valves Are Fitted Corrctly | 3072
10811 | Vacuum System | All Valves Are correctly Connected | 3073
10812 | Damping System | Check Correct Routing Of Air Pipes | 3074
10812 | Damping System | Check Correct Orientation Of Air Valves | 3075
10813 | Chamber | Check Torque Of Chamber Bolts | 3076
10813 | Chamber | Chamber Free From Any Detects | 3077
10813 | Chamber | Screws Fitted To Internal Holes | 3078
10814 | Table Top | Free From Any Defects | 3079
10814 | Table Top | Table Top Secured | 3080
10816 | Fit Column | Correct Chamber Bung Fitted | 3081
416 | Vacuum Test | No Leaks Detected (Prior to running up Turbo) | 3082
416 | Vacuum Test | No Leaks Detected (With Turbo running) when in high sensitivity | 3083
416 | Vacuum Test | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng | 3084
10819 | Fit LHS Column Cabling | Ion Pump Cables | 3085
10819 | Fit LHS Column Cabling | Column X1 Lens cable | 3086
10819 | Fit LHS Column Cabling | Objective Lens Cable | 3087
10819 | Fit LHS Column Cabling | Scan Cable | 3088
10819 | Fit LHS Column Cabling | Penning Cable | 3089
10821 | Fit RHS Column Cabling | Heater Cables | 3090
10821 | Fit RHS Column Cabling | Ion Pump Supply Cables | 3091
10821 | Fit RHS Column Cabling | EHT Cables | 3092
10821 | Fit RHS Column Cabling | Liner Cable | 3093
10821 | Fit RHS Column Cabling | Column Earth Cable | 3094
10822 | Fit Ion Pump PSU | Bracket | 3095
10822 | Fit Ion Pump PSU | Supply | 3096
10822 | Fit Ion Pump PSU | Heater Blocks | 3097
10823 | Fit EHt Set | 24V Supply Kit | 3098
10823 | Fit EHt Set | Spellman EHT Set | 3099
10823 | Fit EHt Set | EHT Set Serial Number | 3100
10823 | Fit EHt Set | Earth Connection Torqued to 3Nm | 3101
113 | Torque Gun assembly earth connection | Safety – Torque the earth stud to 15 N.m | 3102
113 | Torque Gun assembly earth connection | Torque the M8 Zinc plated earth nuts to 15 N.m | 3103
10824 | System Cabling | All Cables Are Neatly Routed | 3104
10824 | System Cabling | All Plugs Are Secured Correctly | 3105
10824 | System Cabling | All Earth Cables Are Tightly Secured | 3106
10824 | System Cabling | Verify the EHT gun assembly earth is torqued p to 15Nm. | 3107
757 | Material Availability | Are there any shortages? | 3108
10826 | Fit Stage Assembly | Select Stage Type | 3109
10826 | Fit Stage Assembly | Serial Number | 3110
10826 | Fit Stage Assembly | Shim Stage Door | 3111
10826 | Fit Stage Assembly | Set Stage Clamp Height - Using Jig | 3112
10827 | Fit Stage Cabling | Stage Cable | 3113
10827 | Fit Stage Cabling | Chamberscope | 3114
10827 | Fit Stage Cabling | Specimen Touch Cable | 3115
11406 | Fit Electrical Feed Through | Select Number of Ways | 3116
11400 | Chamberscope 1 | Fit Chamberscope | 3117
10829 | Install BSD Push/Pull Assembly | Fit Adaptor Plate to Chamber | 3118
10829 | Install BSD Push/Pull Assembly | Fit Amplifier | 3119
10829 | Install BSD Push/Pull Assembly | Check and Setup Push Pull Assembly | 3120
10829 | Install BSD Push/Pull Assembly | Fit Push Pull Assembly | 3121
10716 | Fit TIV | TIV Serial Number | 3122
10716 | Fit TIV | Function Check | 3123
10716 | Fit TIV | Fit TIV | 3124
10719 | Connect Air Pipes | Select Vacuum Type | 3125
10719 | Connect Air Pipes | Luftfilter Fitted | 3126
10720 | Fit Turbo Pump Assembly | Turbo Pump Serial Number | 3127
10720 | Fit Turbo Pump Assembly | Fit Turbo Earth Cable | 3128
10720 | Fit Turbo Pump Assembly | Qualification Marks Checked | 3129
1184 | Turbo Pump Cable Fitted | Turbo cable fitted Correctly | 3130
10722 | Vacuum System | All Vacuum Clamps are Tight | 3131
10722 | Vacuum System | All Valves Are Fitted Corrctly | 3132
10722 | Vacuum System | All Valves Are correctly Connected | 3133
10723 | Damping System | Check Correct Routing Of Air Pipes | 3134
10723 | Damping System | Check Correct Orientation Of Air Valves | 3135
10724 | Chamber | Check Torque Of Chamber Bolts | 3136
10724 | Chamber | Chamber Free From Any Detects | 3137
10725 | Table Top | Free From Any Defects | 3138
10725 | Table Top | Table Top Secured | 3139
11194 | Fit SE Detector | Check that the Assembly is the Correct Part Number | 3140
11194 | Fit SE Detector | Serial Number (or NA) | 3141
11194 | Fit SE Detector | Correct Collector Beak Fitted | 3142
11194 | Fit SE Detector | Fit Detector Assembly | 3143
11194 | Fit SE Detector | Apply Emersion Oil to PM Tube Face | 3144
11194 | Fit SE Detector | Fit PM Head Amp Assembly | 3145
11268 | Install SmartEDX Detector | Detector Type | 3146
11268 | Install SmartEDX Detector | Detector Serial Number | 3147
10835 | Install Nav Cam | SPP 360 Completed | 3148
10836 | Install Nav Cam (Airlock) | SPP 360 Completed | 3149
10838 | Instal Specimen Current Detector | Install SCD Amplifier | 3150
10838 | Instal Specimen Current Detector | Fit SCD Relay PCB - Sigma500 Only | 3151
1145 | Airlock FW Update (MATLAB) | Install Airlock FW Version 2.0.0 | 3152
1162 | SCS Sign off | Does system have an SCS? please select | 3153
1162 | SCS Sign off | If yes, Product Manager to enter initials. If no please type N/A: | 3154
10844 | System Cabling | All Cables Are Neatly Routed | 3155
10844 | System Cabling | All Plugs Are Secured Correctly | 3156
10844 | System Cabling | All Earth Cables Are Tightly Secured | 3157
10844 | System Cabling | Specimen Current Test | 3158
10845 | Detectors & Accessories | Fitted To Correct Ports | 3159
10845 | Detectors & Accessories | Detectors Connected To Correct ASAP Port And Labelled | 3160
10855 | Vacuum Test | No Leaks Detected (Prior to running up Turbo) | 3161
10855 | Vacuum Test | No Leaks Detected (With Turbo running) when in high sensitivity | 3162
10854 | Stage Links | Stage Links Set Correctly | 3163
10856 | Pressure Test-Air System | No Leaks Detected | 3164
11417 | Hi Pot Test | Connect safety check rig (354993-9205-000) and confirm no LED's lit. | 3165
11417 | Hi Pot Test | HAL SN: | 3166
11417 | Hi Pot Test | HAL Calibration Date: | 3167
11417 | Hi Pot Test | DC IR (Direct Current Insulation Resistance @500V) (>10.0MΩ) | 3168
11417 | Hi Pot Test | DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3169
11417 | Hi Pot Test | Confirm System Passed Hi-Pot Tests | 3170
1204 | Heater HiPot Tests HV | IGP Heaters (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3171
1204 | Heater HiPot Tests HV | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3172
1204 | Heater HiPot Tests HV | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3173
1203 | Heater HiPot Tests VP | IGP Heaters (IGP#1) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3174
1203 | Heater HiPot Tests VP | IGP Heaters (IGP#2) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3175
1203 | Heater HiPot Tests VP | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3176
1203 | Heater HiPot Tests VP | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3177
1203 | Heater HiPot Tests VP | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3178
11415 | Hi Pot Verification Test | Confirm test unit fails with shorting test box connected (354933-9205-000). | 3179
951 | PAT Test (Seaward PAT Tester) | Seaward Supernova SN: | 3180
951 | PAT Test (Seaward PAT Tester) | Seaward Calibration Date: | 3181
951 | PAT Test (Seaward PAT Tester) | PE Test - Confirm all test points (<0.1Ω) | 3182
951 | PAT Test (Seaward PAT Tester) | R INS - ( > 0.5MΩ) | 3183
951 | PAT Test (Seaward PAT Tester) | I HLSF - ( < 6000µA) | 3184
951 | PAT Test (Seaward PAT Tester) | Confirm EN 61010 Passed | 3185
951 | PAT Test (Seaward PAT Tester) | Confirm Visual Inspection Criterea Passed | 3186
621 | PAT Tester Verification Test (Seaward Supernova + PAT Checkbox) | PAT Checkbox Calibration Date | 3187
621 | PAT Tester Verification Test (Seaward Supernova + PAT Checkbox) | Earth Bond Test TP1 (0.1Ω - 0.07Ω) | 3188
621 | PAT Tester Verification Test (Seaward Supernova + PAT Checkbox) | Earth Bond Test TP2 (0.55Ω - 0.52Ω) | 3189
621 | PAT Tester Verification Test (Seaward Supernova + PAT Checkbox) | Insulation Test (2.0MΩ - 1.9MΩ) | 3190
621 | PAT Tester Verification Test (Seaward Supernova + PAT Checkbox) | Earth Leakage Test (0.9mA - 0.7mA) | 3191
929 | Print and Fit S/N Label | Print and fit the relevant S/N label to the system. See WI for details. | 3192
930 | China RoHS Compliance Label | Fit China RoHS Compliance Label | 3193
10857 | Pressure Test-Water System | No Leaks Detected | 3194
10849 | Power Up To Standby (Yellow Button) | Rotary Pump Pumping | 3195
10849 | Power Up To Standby (Yellow Button) | Penning Gauge LED is On | 3196
10849 | Power Up To Standby (Yellow Button) | Light Is Lit On EHT Set | 3197
10849 | Power Up To Standby (Yellow Button) | The Distance Between The Vac Cable Loom AND The Plinth Base Is Greater Than 30mm | 3198
10850 | CAN Bus Test | Vacuum PCB Test Passed | 3199
10850 | CAN Bus Test | EO PCB Test Passed | 3200
10850 | CAN Bus Test | Stage PCB Test Passed | 3201
10853 | Power Supply Set Up (ON LOAD) | Vac 5.15V (P1) | 3202
10853 | Power Supply Set Up (ON LOAD) | Vac 5.00V (P10) | 3203
10853 | Power Supply Set Up (ON LOAD) | Stage 5.15V | 3204
10853 | Power Supply Set Up (ON LOAD) | Stage 21.5V | 3205
10853 | Power Supply Set Up (ON LOAD) | Stage -21.5V | 3206
10853 | Power Supply Set Up (ON LOAD) | EO 5.15V | 3207
10851 | Run Up SmartSEM | ASAP FW/HW Present | 3208
10851 | Run Up SmartSEM | Stage FW/HW Present | 3209
10851 | Run Up SmartSEM | Vacuum FW/HW Present | 3210
10851 | Run Up SmartSEM | EO PCB FW/HW Present | 3211
1232 | SmartSEM conditions | System Vacuum Reading - Confirm system vacuum reads < 1x10-5mbar | 3212
1232 | SmartSEM conditions | Gun Vacuum Reading OK - Confirm gun vacuum <3x10-6 mbar (If vacuum showing no improvment please contact ME) | 3213
1232 | SmartSEM conditions | Initialise Stage - Confirm stage initialises without errors | 3214
1232 | SmartSEM conditions | Water flow test - Confirm SmartSEM reports 'Water loss error' when water switched off at services. | 3215
10852 | Run UP SmartSEM Service Centre | PSU Tests Passed | 3216
10852 | Run UP SmartSEM Service Centre | EO Tests Passed | 3217
11248 | Production Documentation | Print Plinth Seperators | 3218
11248 | Production Documentation | Sort Documentation with Seperators | 3219
10858 | All Paperwork Present | Picking List | 3220
10858 | All Paperwork Present | Turbo Pump | 3221
10858 | All Paperwork Present | Turbo Pump Controller | 3222
10858 | All Paperwork Present | Chamber | 3223
10858 | All Paperwork Present | Plinth | 3224
10858 | All Paperwork Present | EO PCB | 3225
10858 | All Paperwork Present | Stage PCB | 3226
10858 | All Paperwork Present | Stage Logbook | 3227
10859 | System Cabling | All Cables Are Neatly Routed | 3228
10859 | System Cabling | All Plugs Are Secured Correctly | 3229
10859 | System Cabling | All Earth Cables Are Tightly Secured | 3230
10859 | System Cabling | Check Rear Panel Cables Are Correctly Located | 3231
10859 | System Cabling | Check Opto Devices Are Not Damaged On All PCBs & Mains Box | 3232
10859 | System Cabling | Check Stage Board Link Configuration | 3233
10860 | Detector & Accessories | Fitted To Correct ports | 3234
10860 | Detector & Accessories | Connected to Correct ASAP Ports & Labeled | 3235
10861 | Air & Water System | Air Pipes Are Correctly Routed and Kink Free | 3236
10861 | Air & Water System | Pressure Regulators Are Fitted In THe Correct Orientation | 3237
10861 | Air & Water System | WaterPipes Are Correctly Routed And Free From Kinks | 3238
10862 | Chamber & Table Top | Chamber is Free From Defects | 3239
10862 | Chamber & Table Top | Chamber Meshes Are Fitted | 3240
10862 | Chamber & Table Top | Chamber Pumping Pipe Is Fitted Centrally To The Table Top | 3241
10862 | Chamber & Table Top | Table Top Is Free From Defects | 3242
10863 | Plinth | Plinth Base Is Clean and Free From Debris | 3243
10864 | Bay | Clean & Tidy Assembly Bay | 3244
966 | Confirm 4-eye check by additional operator | Checked Carried Our By: | 3245
842 | Wiring, Micro switches and Cable form Fitment. (Set Micro-switch Positions) | Z Microswitch Position Set | 3246
842 | Wiring, Micro switches and Cable form Fitment. (Set Micro-switch Positions) | T Microswitch Position Set | 3247
842 | Wiring, Micro switches and Cable form Fitment. (Set Micro-switch Positions) | R Microswitch Position Set | 3248
843 | Fit ZTR Motors and place the appro-priate sticker in the box provided be-low: | Z Motor SeriaL Number: | 3249
843 | Fit ZTR Motors and place the appro-priate sticker in the box provided be-low: | T Motor Serial Number: | 3250
843 | Fit ZTR Motors and place the appro-priate sticker in the box provided be-low: | R Motor Serial Number: | 3251
845 | Confirm 4-eye check by additional operator | Checked Carried Our By: | 3252
860 | Confirm 4-eye check by additional operator | Checked Carried Our By: | 3253
868 | Confirm 4-eye check by additional operator | Checked Carried Our By: | 3254
871 | X Axis Micro Switch | Check Set Point | 3255
870 | Y Axis Micro Switch | Check Set Point | 3256
873 | Z Axis Micro Switch | Check Set Point | 3257
874 | T Axis Micro Switch | Check Set point (-10.8º) | 3258
875 | Stage Initialisation | Check Stage Initialises Correctly | 3259
876 | X Travel | Select Stage Type and Check Travel | 3260
877 | Y Travel | Select Stage Type and Check Travel | 3261
878 | Z Travel | Select Stage Type and Check Travel | 3262
879 | Tilt Travel | Check Tilt Travel -10° to +90° | 3263
880 | Stage Axis Cycle Test | Stage Axis Cycle Test with 0.5Kg Weight for 10 Mins | 3264
881 | Stage Axis Cycle Test (1 Hour Min) | Run Axis Cycle Test | 3265
882 | EV025 Stage Kit | Stage Kit | 3266
882 | EV025 Stage Kit | 27mm Spacer Block | 3267
882 | EV025 Stage Kit | 54mm Spacer Block | 3268
882 | EV025 Stage Kit | 81mm Spacer Block - Fitted | 3269
903 | EVO15/ Sigma Stage Kit | Stage Kit | 3270
903 | EVO15/ Sigma Stage Kit | 27mm Spacer Block Fitted | 3271
905 | Stage Type & Serial Number | Select Stage Type | 3272
905 | Stage Type & Serial Number | Stage Serial Number: | 3273
884 | Quality sign off. | Quality Sign Off Carried Out By: | 3274
896 | Build Air Water Coupling (DJEL) | Record the temp sensor probe resistance | 3275
743 | Material Availability | Are there any shortages? | 3276
11253 | SCS Check | Check SCS Folder/Sales Order | 3277
11253 | SCS Check | SCS Number (If applicable) | 3278
10 | Mount AC unit | Record AC unit S/N | 3279
10 | Mount AC unit | Qualification Marks Checked | 3280
10 | Mount AC unit | Fit mounting bracket to AC unit 'Angle Bracket' - 349500-8040-000 | 3281
10 | Mount AC unit | Set up mounting bracket height using Jig 356580-0174-000 | 3282
12 | Fit EHT Assembly | Serial No. | 3283
12 | Fit EHT Assembly | Version | 3284
12 | Fit EHT Assembly | Qualification Marks Checked | 3285
5 | Build Air and Water Coupling | Fit the pre-assembled Air and Water assembly | 3286
5 | Build Air and Water Coupling | Loctite 243 applied to water fitting nuts on back panel. | 3287
9 | Assemble Mount and Setup water flow switch | Set water flow pressure to 0.9L / min | 3288
6 | Fit Nitrogen Lines | Fit nitrogen hoses # 1,2,3,4,5,7,9 | 3289
6 | Fit Nitrogen Lines | Set regulator pressure to 0.03MPa | 3290
7 | Fit Stage PCB (5 Axis Stage) | Fit Stage PCB | 3291
1264 | Fit Stage PCB (6 Axis Stage) | Fit Stage PCB | 3292
8 | Install Ion Pump PSU 2 | Record IPPSU S/N | 3293
8 | Install Ion Pump PSU 2 | Set DIP switches | 3294
8 | Install Ion Pump PSU 2 | Qualification Marks Checked | 3295
1247 | Install Ion Pump PSU 1 | Fit IPPSU 1, Record S/N | 3296
1247 | Install Ion Pump PSU 1 | Set DIP switches | 3297
1247 | Install Ion Pump PSU 1 | Qualification Marks Checked | 3298
13 | Fit Dynamic EO board - Single Scan (346001-9110-315) | Confirm EO Dynamic PCB PN is '346001-9110-315' | 3299
13 | Fit Dynamic EO board - Single Scan (346001-9110-315) | Fit EO Dynamic PCB | 3300
1153 | Fit Dynamic EO board - Double Scan (346001-9110-305) | Confirm EO PCB PN is '346001-9110-305' | 3301
1153 | Fit Dynamic EO board - Double Scan (346001-9110-305) | Fit EO Dynamic PCB | 3302
14 | Fit Static EO board SC (346001-9100-410) | Confirm EO Static PCB PN is '346001-9100-410' | 3303
14 | Fit Static EO board SC (346001-9100-410) | Fit EO Static PCB | 3304
1154 | Fit Static EO board DC (346001-9100-400) | Confirm EO Static PCB PN is '346001-9100-400' | 3305
1154 | Fit Static EO board DC (346001-9100-400) | Fit EO Static PCB | 3306
1218 | Assemble and fit ASAP / SVSEM | Check if ASAP / H-ASAP / SVSEM | 3307
1218 | Assemble and fit ASAP / SVSEM | Record Serial No | 3308
1218 | Assemble and fit ASAP / SVSEM | Fit additional ASAP cards if provided (DFE / VBE Cards) | 3309
17 | Fit Earth Bonding Straps | Fit X4 braided earth straps to rear PCB's. (Vac, EOS, EOD, ASAP /SVSEM) | 3310
17 | Fit Earth Bonding Straps | Fit earth cable to DC unit. | 3311
11309 | Check Raith damping system version. | Correct version damping system supplied | 3312
20 | Install damping system slave | Install Compressed Air Line | 3313
20 | Install damping system slave | Set foot height | 3314
21 | Install damping system X142 | Set Valve Screw torque IN (3 turns) | 3315
21 | Install damping system X142 | Set Valve Screw torque OUT (3 turns) | 3316
21 | Install damping system X142 | Set foot height | 3317
21 | Install damping system X142 | Install Compressed Air Line | 3318
21 | Install damping system X142 | Install Control Cable | 3319
22 | Install damping system X141 | Set Valve Screw torque IN (1.5 turns) | 3320
22 | Install damping system X141 | Set Valve Screw torque OUT(3 turns) | 3321
22 | Install damping system X141 | Set foot height | 3322
22 | Install damping system X141 | Install Compressed Air Line | 3323
22 | Install damping system X141 | Install Control Cable | 3324
23 | Install damping system X140 | Set Valve Screw torque IN (1.5 turns) | 3325
23 | Install damping system X140 | Set Valve Screw torque OUT(3 turn) | 3326
23 | Install damping system X140 | Set foot height | 3327
23 | Install damping system X140 | Install Compressed Air Line | 3328
23 | Install damping system X140 | Install Control Cable | 3329
745 | Material Availability | Are there any shortages? | 3330
25 | Fit system cables - Earth terminal strip, route. Connect earth. | Torque all earth cables to fitted to earth terminal strip. | 3331
25 | Fit system cables - Earth terminal strip, route. Connect earth. | Torque the following earth connections; DC Unit, Stage PCB, Turbo pump. | 3332
1265 | Fit Stage Links (5 Axis Stage) | Set Stage links (348247-9001-500) | 3333
1265 | Fit Stage Links (5 Axis Stage) | Fill in Jumper settings label | 3334
1266 | Fit Stage Links (6 Axis Stage) | Set Stage links (348247-9002-120) | 3335
1266 | Fit Stage Links (6 Axis Stage) | Fill in Jumper settings label | 3336
1267 | Reconfigure light barrier (XB550L) | Reconfigure loght barries for XB-550L Stage. | 3337
30 | Fit System cable - EHT | Fit EHT Cables | 3338
30 | Fit System cable - EHT | Torque EHT Connections to 3Nm | 3339
10765 | Hi Pot Test (Mains Box) | Connect safety check rig (354993-9205-000) and confirm no LED's lit. | 3340
10765 | Hi Pot Test (Mains Box) | HAL SN: | 3341
10765 | Hi Pot Test (Mains Box) | HAL Calibration Date: | 3342
10765 | Hi Pot Test (Mains Box) | DC IR (Direct Current Insulation Resistance @500V) (>10.0MΩ) | 3343
10765 | Hi Pot Test (Mains Box) | DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3344
10765 | Hi Pot Test (Mains Box) | Confirm System Passed Hi-Pot Tests | 3345
1270 | Heater HiPot Test | HAL SN: | 3346
1270 | Heater HiPot Test | HAL Calibration Date: | 3347
1270 | Heater HiPot Test | Column Heater (354920-9179-000). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3348
1357 | Accesories Configuration | Port configuration Check | 3349
748 | Material Availability | Are there any shortages? | 3350
11389 | Preparing and Fitting Chamber | Fit chamber pumping pipe | 3351
11389 | Preparing and Fitting Chamber | Fit chamber to damper | 3352
11389 | Preparing and Fitting Chamber | Fit chamber bolts | 3353
11389 | Preparing and Fitting Chamber | Fit earthquake bolts | 3354
11389 | Preparing and Fitting Chamber | Record Chamber SN | 3355
668 | Preparing and Fitting Large Chamber (360-L, 460-L and 560-L) | Fit chamber to damper | 3356
668 | Preparing and Fitting Large Chamber (360-L, 460-L and 560-L) | Fit chamber bolts item 10 and 12 and Torque to 42Nm. | 3357
668 | Preparing and Fitting Large Chamber (360-L, 460-L and 560-L) | Fit chamber bolts item 13 and 14 and Torque to 21Nm. | 3358
668 | Preparing and Fitting Large Chamber (360-L, 460-L and 560-L) | Fit chamber pumping tube and Toruqe bolts item 7 to 12Nm. | 3359
668 | Preparing and Fitting Large Chamber (360-L, 460-L and 560-L) | Fit Turbo guard with non silver vent screw 000000-0539-927. | 3360
668 | Preparing and Fitting Large Chamber (360-L, 460-L and 560-L) | Record Chamber SN | 3361
1190 | Stage Baking Trial (Temp test point) | Was the Stage baked independently? (Chamber not baked) | 3362
37 | Fit Stage in chamber | Shim Stage | 3363
37 | Fit Stage in chamber | Check Mag Clamp (adjust position if necessary, see WI) | 3364
37 | Fit Stage in chamber | Fit Gun Interlock Key to Stage Door | 3365
37 | Fit Stage in chamber | Record Stage SN (Can be found on right hand side of XY platform) | 3366
39 | Assemble and fit Turbo Stack | Turbo Pump S/No. | 3367
39 | Assemble and fit Turbo Stack | Qualification Marks Checked | 3368
629 | Fit quiet mode Pirani Gauge | Ensure Pirani gauge is isolated from the roughing line | 3369
629 | Fit quiet mode Pirani Gauge | For the APG100-XM Remove the 2 metal clips, but for the APG200-XM Leave them in | 3370
42 | Fit flexi pipe to turbo pump | Fit flexi pipe to turbo | 3371
42 | Fit flexi pipe to turbo pump | Fit 3X isolation O-rings to flexi pipe | 3372
1254 | Fit Ducting Plate | Fit Ducting Plate & Earth Cables | 3373
1254 | Fit Ducting Plate | Fit Manifold | 3374
1254 | Fit Ducting Plate | Fit Air and Nitrogen Pipes | 3375
43 | Fit TIV Nitrogen pipe | Fit the Nitrogen vent pipe to TIV vent valve | 3376
1262 | Fit Ducting Plate | Fit Ducting Plate & Earth Cables | 3377
1262 | Fit Ducting Plate | Fit Manifold | 3378
1262 | Fit Ducting Plate | Fit Air and Nitrogen Pipes | 3379
46 | Clean Chamber | Carboflow Holes | 3380
46 | Clean Chamber | Carboflow Chamber | 3381
46 | Clean Chamber | IPA Holes | 3382
46 | Clean Chamber | IPA Chamber | 3383
11384 | Leak check - Vacuum Test | Warning -ONLY VENT using the electronic vent valve in the chamber | 3384
11384 | Leak check - Vacuum Test | No Leaks Detected (Prior to running up Turbo) | 3385
11384 | Leak check - Vacuum Test | No Leaks Detected (With Turbo running) when in high sensitivity | 3386
11016 | Bake Start | Time = | 3387
11016 | Bake Start | Date = | 3388
637 | Plasma Clean Chamber | 3 hour plasma clean completed | 3389
11017 | Bake Finish | Time = | 3390
11017 | Bake Finish | Date = | 3391
764 | Material Availability | Are there any shortages? | 3392
750 | Material Availability | Are there any shortages? | 3393
49 | Check Plinth Paperwork is complete for T3 | Check Plinth Paperwork is complete for Takt 3 | 3394
10992 | Fit In-Lens Multiplier | Apply 1 drop of immersion oil to PM glass | 3395
10993 | Fit ESB Multiplier | Apply 1 drop immersion oil to PM glass | 3396
1255 | Label up cables | In-Lens And ESB cable | 3397
1255 | Label up cables | IGP cables (IGP1 & IGP2) | 3398
1255 | Label up cables | SCM cable | 3399
1255 | Label up cables | CCD camera cables | 3400
1544 | Fit Airlock | Fit A/L Key to Stage Door | 3401
1544 | Fit Airlock | Check distance of 200mm door extended = 228mm | 3402
10959 | Leak Check - Vacuum test Column + Airlock | No Leaks Detected (With Turbo running) when in high sensitivity | 3403
10959 | Leak Check - Vacuum test Column + Airlock | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng | 3404
11274 | Fit Column Cables + Pipes | Fit and route all column cables | 3405
11274 | Fit Column Cables + Pipes | Fit and route objective water pipes | 3406
11274 | Fit Column Cables + Pipes | Safety – Torque the earth stud to 15 N.m | 3407
11274 | Fit Column Cables + Pipes | Torque both earth zinc plated M8 nuts to 15 N.m | 3408
11274 | Fit Column Cables + Pipes | Confirm HV cable is routed though plinth, and secured on top of plinth | 3409
423 | Routing of Raith liner cable | Confirm Liner cable routed as specified | 3410
1256 | SCM | Fit SCM | 3411
1257 | SCM + Nav Cam | Fit SCM | 3412
1257 | SCM + Nav Cam | Fit Nav Cam | 3413
50 | Fit camera | Fit Camera | 3414
58 | Accesories Configuration | Port configuration Check | 3415
189 | Sense BSD | Ensure the detector is fitted the correct way up | 3416
10991 | VPSE | Fit Indigo Filter | 3417
1259 | Route all Chamber / Stage / Detector Cables and pipes | Configure ASAP / SVSEM detectors ports as specified in procedure. | 3418
1261 | Nitrogen purge valve. | Fit Nitrogen purge valve to turbo. | 3419
1261 | Nitrogen purge valve. | Fit Nitrogen vent line. | 3420
1261 | Nitrogen purge valve. | Fit vent cable. | 3421
1161 | SCS Sign off | Does system have an SCS? please select | 3422
1161 | SCS Sign off | If yes, Product Manager to enter initials. If no please type N/A: | 3423
60 | Pressure Test Water Connections | No Pressure loss in 2mins | 3424
1252 | Chamber checks | Check interlock keys fitted to door | 3425
1252 | Chamber checks | Check vented screws fitted to inside of chamber ceiling | 3426
1252 | Chamber checks | Check magnet clamp set correctly | 3427
1252 | Chamber checks | Check door alignment | 3428
61 | Connect Services | Safety Earth | 3429
61 | Connect Services | Compressed Air (Set to 0.5MPa) | 3430
61 | Connect Services | PVP Valve. | 3431
61 | Connect Services | Power - 32A Connected to wall and switch ON. | 3432
61 | Connect Services | Rotary Pump IEC Lead | 3433
61 | Connect Services | Nitrogen (Set to 0.05 MPa inside plinth) | 3434
61 | Connect Services | Cooling Water - Confirm flow is set to 0.9LPM | 3435
61 | Connect Services | EHT Lens output resistance (Safety Test) ( < 2.2kΩ) | 3436
61 | Connect Services | Switch on system breaker and run system to green button | 3437
62 | Load VAC Firmware | Firmware (Mini Mon) | 3438
62 | Load VAC Firmware | PLDs (Xlinks) | 3439
63 | System VAC TEST Program | Confirm Vacuum System is configured correctly | 3440
63 | System VAC TEST Program | Confirm all vacuum modules are Green in Vac Test | 3441
63 | System VAC TEST Program | Confirm Turbo reached 1000Hz and penning reading | 3442
63 | System VAC TEST Program | Confirm IGP's haev switched on without trip (system vac better than 2x10-5mbar) | 3443
1253 | System Checker | Confirm all FW is correct (Update any FW not latest version) | 3444
64 | SmartSEM | Start SmartSEM and confirm there are no errors in EM server. | 3445
64 | SmartSEM | Initialise Stage can be initialised with no errors. | 3446
69 | Damping System Checks | Foot Height | 3447
69 | Damping System Checks | Pipe Configuration | 3448
69 | Damping System Checks | Sealing Plugs Removed | 3449
921 | Water Flow Tests | Record Water Flow Temperature (20ºC +/- 5ºC) | 3450
921 | Water Flow Tests | Water flow test - Confirm SmartSEM reports 'Water loss error' when water switched off at services. | 3451
68 | Leak Check - Vacuum test Full System | No Leaks Detected (With Turbo running) when in high sensitivity | 3452
1546 | Heater HiPot Tests HV | HAL SN: | 3453
1546 | Heater HiPot Tests HV | HAL Calibration Date: | 3454
1546 | Heater HiPot Tests HV | IGP Heaters (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3455
1546 | Heater HiPot Tests HV | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3456
1546 | Heater HiPot Tests HV | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3457
11411 | Hi Pot Verification Test (HAL 101 High Voltage Tester) | Confirm test unit fails with shorting test box connected (354933-9205-000). | 3458
625 | Heater HiPot Tests VP | HAL SN: | 3459
625 | Heater HiPot Tests VP | HAL Calibration Date: | 3460
625 | Heater HiPot Tests VP | IGP Heaters (IGP#1) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3461
625 | Heater HiPot Tests VP | IGP Heaters (IGP#2) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3462
625 | Heater HiPot Tests VP | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3463
625 | Heater HiPot Tests VP | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3464
625 | Heater HiPot Tests VP | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 3465
1545 | HiPot Verification Test (HAL 101 High Voltage Tester) | Confirm test unit fails using the High voltage probe, Hi-Pot Test Set (356180-9217-000) and shorting test box connected (354933-9205-000). | 3466
628 | PAT Test | Secutest SN: | 3467
628 | PAT Test | Secutest Calibration Date: | 3468
628 | PAT Test | PE Test - Confirm all test points (<0.1Ω) | 3469
628 | PAT Test | R INS - ( > 0.5MΩ) | 3470
628 | PAT Test | I HLSF - ( < 8000µA) | 3471
628 | PAT Test | Confirm EN 61010 Passed | 3472
628 | PAT Test | Confirm Visual Inspection Criterea Passed | 3473
1547 | PAT Tester Verification Test (Secustest SII PAT Tester + PAT Checkbox) | PAT Checkbox Calibration Date | 3474
1547 | PAT Tester Verification Test (Secustest SII PAT Tester + PAT Checkbox) | Earth Bond Test TP1 (0.07Ω - 0.09Ω) | 3475
1547 | PAT Tester Verification Test (Secustest SII PAT Tester + PAT Checkbox) | Earth Bond Test TP2 (0.50Ω - 0.53Ω) | 3476
1547 | PAT Tester Verification Test (Secustest SII PAT Tester + PAT Checkbox) | Insulation Test (2.0MΩ - 1.9MΩ) | 3477
1547 | PAT Tester Verification Test (Secustest SII PAT Tester + PAT Checkbox) | Earth Leakage Test (0.7mA - 0.9mA) | 3478
1236 | Print and Fit SN Label (NO CSA, NO SEMI) | Confirm SN label does not have any CSA markings. | 3479
1236 | Print and Fit SN Label (NO CSA, NO SEMI) | Confirm SN label does not have any SEMI markings. | 3480
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | Confirm SN label does not have any CSA markings. | 3481
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | Confirm SN label does have the correct SEMI marking. | 3482
1234 | Print and Fit S/N Label (CSA UL certification) | Confirm SN label does have the correct CSA marking. | 3483
1234 | Print and Fit S/N Label (CSA UL certification) | Confirm SN label does not have any SEMI marking. | 3484
94 | Fit HE Inlens Detector | Serial: | 3485
94 | Fit HE Inlens Detector | Remove metal protective cover from detector and inspect detector head | 3486
1235 | Print and Fit S/N Label ( SEMI S2 S8 certification + CSA UL certification) | Confirm SN label does have the correct CSA marking. | 3487
1235 | Print and Fit S/N Label ( SEMI S2 S8 certification + CSA UL certification) | Confirm SN label does have the correct SEMI marking. | 3488
710 | Record Final Gun Vacuum | Opening of column manual valves required to achieve gun vacuum | 3489
710 | Record Final Gun Vacuum | Actual IGP1 vacuum | 3490
222 | SCS Check | Check SCS Folder/Sales Order | 3491
222 | SCS Check | SCS Number (If applicable) | 3492
507 | Inspect and Clean Parts | Anode Inspected & cleaned in Ultra Sonic tank. | 3493
507 | Inspect and Clean Parts | Final Aperture Holder (Multi / Single) Inspected & cleaned in Ultra Sonic tank. | 3494
11237 | Ensure Open Anode (Non EBBU) | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed | 3495
11239 | Ensure Closed Anode (Non EBBU) | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered | 3496
11236 | Ensure Open Anode (EBBU) | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed. | 3497
1645 | Fit ESB Detector | Serial: | 3498
1645 | Fit ESB Detector | Remove metal protective cover from detector and inspect detector head | 3499
11264 | Ensure Closed Anode (EBBU) | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be observed. | 3500
11265 | EBBU Anode | Make sure there are 3 cut-outs on the edge of the anode. No particles are to be observed. | 3501
11263 | EBBU Anode | Make sure there are 3 cut-outs on the edge of the anode. No particles are to be observed | 3502
504 | Closed Multi Hole aperture holder observations | Select PC | 3503
504 | Closed Multi Hole aperture holder observations | Hold aperture holder, look for holes on the sides.No particles are to be observed | 3504
505 | Ventilated Multi Open Hole aperture holder observations | Hold aperture holder, look for holes on the sides.No particles are to be observed | 3505
508 | Assemble Final Aperture (Multi / Single) and Anode | Assemble Final Aperture (Multi / Single). Inspect using optical microscope. | 3506
508 | Assemble Final Aperture (Multi / Single) and Anode | Flame Anode aperture and assemble into Anode. | 3507
91 | Section Type | Housing Type | 3508
91 | Section Type | Inspect insulation disc using white light for cracks | 3509
11187 | Multi_TS_HRQA 100nA, Surface Overview and Concentricity | Check there is no engraved ring around the cluster of apertures, this should be on the underside. | 3510
11187 | Multi_TS_HRQA 100nA, Surface Overview and Concentricity | Multi_TS_HRQA - Passed inspection criteria, Image Saved | 3511
11277 | Multi_TS_HRQA 100nA, Surface Overview and Concentricity | Check there is no engraved ring around the cluster of apertures, this should be on the underside. | 3512
11277 | Multi_TS_HRQA 100nA, Surface Overview and Concentricity | Multi_TS_HRQA - Passed inspection criteria, Image Saved | 3513
11283 | Multi_TS_HRQA 20nA, Surface Overview and Concentricity | Check there is no engraved ring around the cluster of apertures, this should be on the underside. | 3514
11283 | Multi_TS_HRQA 20nA, Surface Overview and Concentricity | Multi_TS_HRQA - Passed inspection criteria, Image Saved | 3515
11284 | Multi_TS_HRQA 20nA, Surface Overview and Concentricity | Check there is no engraved ring around the cluster of apertures, this should be on the underside. | 3516
11284 | Multi_TS_HRQA 20nA, Surface Overview and Concentricity | Multi_TS_HRQA - Passed inspection criteria, Image Saved | 3517
11252 | Multi_TS_30um 100nA, 30um Centre Hole | Measure aperture size, 30um (+/- 2um) | 3518
11252 | Multi_TS_30um 100nA, 30um Centre Hole | Multi_TS_30um - Passed inspection criteria, Image Saved | 3519
11285 | Multi_TS_30um 20nA, 30um Centre Hole | Measure aperture size, 30um (+/- 2um) | 3520
11285 | Multi_TS_30um 20nA, 30um Centre Hole | Multi_TS_30um - Passed inspection criteria, Image Saved | 3521
1274 | Plasma Clean Parts | 20 min plasma cycle carried out - Ceramic Disc, Bean Can, Gehause Mont | 3522
11156 | Multi_TS_20um 100nA, 20um Centre Hole | Measure aperture size, 20um (+/- 2um) | 3523
11156 | Multi_TS_20um 100nA, 20um Centre Hole | Multi_TS_20um - Passed inspection criteria, Image Saved | 3524
11286 | Multi_TS_20um 20nA, 20um Centre Hole | Measure aperture size, 20um (+/- 2um) | 3525
11286 | Multi_TS_20um 20nA, 20um Centre Hole | Multi_TS_20um - Passed inspection criteria, Image Saved | 3526
11160 | Multi_TS_120um 20nA, 120um Large Aperture | Measure aperture size, 120um (+/- 2um) | 3527
11160 | Multi_TS_120um 20nA, 120um Large Aperture | Multi_TS_120um - Passed inspection criteria, Image Saved | 3528
11287 | Multi_TS_120um 20nA, 120um Large Aperture | Measure aperture size, 120um (+/- 2um) | 3529
11287 | Multi_TS_120um 20nA, 120um Large Aperture | Multi_TS_120um - Passed inspection criteria, Image Saved | 3530
11288 | Multi_TS_300um 100nA, 300um Large Aperture | Measure aperture size, 300um (+/- 10um) | 3531
11288 | Multi_TS_300um 100nA, 300um Large Aperture | Multi_TS_300um - Passed inspection criteria, Image Saved | 3532
11159 | Multi_TS_300um 100nA, 300um Large Aperture | Measure aperture size, 300um (+/- 10um) | 3533
11159 | Multi_TS_300um 100nA, 300um Large Aperture | Multi_TS_300um - Passed inspection criteria, Image Saved | 3534
92 | Fit the Objective Lens & Connector Brackets | Cooling Isolation Check (Open Circuit Coil – GND) | 3535
11161 | Multi_US_HRQA 100nA, Surface Overview | Check there is an engraved ring around the cluster of apertures, this should be on the underside | 3536
11161 | Multi_US_HRQA 100nA, Surface Overview | Multi_US_HRQA - Passed inspection criteria, Image Saved | 3537
215 | TTL Upgrade | Fit Valves and Pipework | 3538
215 | TTL Upgrade | Connect Valve Cable to PL3 on Vac PCB | 3539
414 | Vacuum Test | No Leaks Detected (Prior to running up Turbo) | 3540
414 | Vacuum Test | No Leaks Detected (With Turbo running) when in high sensitivity | 3541
414 | Vacuum Test | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng | 3542
10728 | Fit Column Cabling | Scan Cables | 3543
10728 | Fit Column Cabling | Lens Cables | 3544
10728 | Fit Column Cabling | EHT Cables | 3545
10728 | Fit Column Cabling | Penning Cable | 3546
11413 | PAT Test | Secutest SN: | 3547
11413 | PAT Test | Secutest Calibration Date: | 3548
11413 | PAT Test | PE Test - Confirm all test points (<0.1Ω) | 3549
11413 | PAT Test | R INS - ( > 0.5MΩ) | 3550
11413 | PAT Test | I HLSF - ( < 6000µA) | 3551
11413 | PAT Test | Confirm EN 61010 Passed | 3552
11413 | PAT Test | Confirm Visual Inspection Criterea Passed | 3553
11289 | Multi_US_HRQA 100nA, Surface Overview | Check there is an engraved ring around the cluster of apertures, this should be on the underside | 3554
11289 | Multi_US_HRQA 100nA, Surface Overview | Multi_US_HRQA - Passed inspection criteria, Image Saved | 3555
11290 | Multi_US_HRQA 20nA, Surface Overview 30um Centre Hole | Check there is an engraved ring around the cluster of apertures, this should be on the underside | 3556
11290 | Multi_US_HRQA 20nA, Surface Overview 30um Centre Hole | Multi_US_HRQA - Passed inspection criteria, Image Saved | 3557
11308 | Multi_US_HRQA 20nA, Surface Overview 20um Centre Hole | Check there is an engraved ring around the cluster of apertures, this should be on the underside | 3558
11308 | Multi_US_HRQA 20nA, Surface Overview 20um Centre Hole | Multi_US_HRQA - Passed inspection criteria, Image Saved | 3559
506 | Additional Raith Aperture Measurments | Measure aperture size, 7.5um (+/-2um) | 3560
506 | Additional Raith Aperture Measurments | Multi_TS_7.5um - Passed inspection criteria, Image Saved | 3561
506 | Additional Raith Aperture Measurments | Measure aperture size, 10um (+/-2um) | 3562
506 | Additional Raith Aperture Measurments | Multi_TS_10um - Passed inspection criteria, Image Saved | 3563
506 | Additional Raith Aperture Measurments | Measure aperture size, 15um (+/-2um) | 3564
506 | Additional Raith Aperture Measurments | Multi_TS_15um - Passed inspection criteria, Image Saved | 3565
506 | Additional Raith Aperture Measurments | Measure aperture size, 20um (+/-2um) | 3566
506 | Additional Raith Aperture Measurments | Multi_TS_20um - Passed inspection criteria, Image Saved | 3567
506 | Additional Raith Aperture Measurments | Measure aperture size, 60um (+/-2um) | 3568
506 | Additional Raith Aperture Measurments | Multi_TS_60um - Passed inspection criteria, Image Saved | 3569
11262 | Anode_TS_HRQA 100nA, Surface Overview and Concentricity | Anode_TS_HRQA - Passed inspection criteria, Image Saved | 3570
11317 | Anode_TS_HRQA 300nA, Surface Overview and Concentricity | Anode_TS_HRQA - Passed inspection criteria, Image Saved | 3571
96 | Column Separation Valve Checks | Pneumatic Operation | 3572
96 | Column Separation Valve Checks | Pin Touching when Retracted | 3573
11318 | Anode_TS_HRQA 20nA, Surface Overview and Concentricity | Anode_TS_HRQA - Passed inspection criteria, Image Saved | 3574
11319 | Anode_TS_HRQA 40nA, Surface Overview and Concentricity | Anode_TS_HRQA - Passed inspection criteria, Image Saved | 3575
11163 | Anode_TS_Bore1 100nA, First Bore Shot | Anode_TS_Bore1 - Passed inspection criteria, Image Saved | 3576
11326 | Anode_TS_Bore1 300nA, First Bore Shot | Anode_TS_Bore1 - Passed inspection criteria, Image Saved | 3577
11327 | Anode_TS_Bore1 20nA, First Bore Shot | Anode_TS_Bore1 - Passed inspection criteria, Image Saved | 3578
11328 | Anode_TS_Bore1 40nA, First Bore Shot | Anode_TS_Bore1 - Passed inspection criteria, Image Saved | 3579
1655 | UHV Chamber S/N | S/N = | 3580
1655 | UHV Chamber S/N | Please confirm UHV supplier | 3581
11164 | Anode_TS_Bore2 100nA, Second Bore Shot | Anode_TS_Bore2 - Passed inspection criteria, Image Saved | 3582
11329 | Anode_TS_Bore2 300nA, Second Bore Shot | Anode_TS_Bore2 - Passed inspection criteria, Image Saved | 3583
11330 | Anode_TS_Bore2 20nA, Second Bore Shot | Anode_TS_Bore2 - Passed inspection criteria, Image Saved | 3584
11331 | Anode_TS_Bore2 40nA, Second Bore Shot | Anode_TS_Bore2 - Passed inspection criteria, Image Saved | 3585
11165 | Anode_TS_40um 20nA, 40um Centre Hole | Measure aperture size, 40um (+/- 2um) | 3586
11165 | Anode_TS_40um 20nA, 40um Centre Hole | Anode_TS_40um - Passed inspection criteria, Image Saved | 3587
11166 | Anode_TS_55um 40nA, 55um Centre Hole | Measure aperture size, 55um (+/- 2um) | 3588
11166 | Anode_TS_55um 40nA, 55um Centre Hole | Anode_TS_55um - Passed inspection criteria, Image Saved | 3589
11167 | Anode_TS_90um 100nA, 90um Centre Hole | Measure aperture size, 90um (+/- 5um) | 3590
11167 | Anode_TS_90um 100nA, 90um Centre Hole | Anode_TS_90um - Passed inspection criteria, Image Saved | 3591
11168 | Anode_TS_150um 300nA, 150um Centre Hole | Measure aperture size, 150um (+/- 5um) | 3592
11168 | Anode_TS_150um 300nA, 150um Centre Hole | Anode_TS_90um - Passed inspection criteria, Image Saved | 3593
11169 | Anode_US_HRQA 100nA, Surface Overview | Anode_US_HRQA - Passed inspection criteria, Image Saved | 3594
11332 | Anode_US_HRQA 300nA, Surface Overview | Anode_US_HRQA - Passed inspection criteria, Image Saved | 3595
11334 | Anode_US_HRQA 40nA, Surface Overview | Anode_US_HRQA - Passed inspection criteria, Image Saved | 3596
11333 | Anode_US_HRQA 20nA, Surface Overview | Anode_US_HRQA - Passed inspection criteria, Image Saved | 3597
1166 | Enter measurement tools SN | Anode + Extractor measuring jig S/N = | 3598
1166 | Enter measurement tools SN | Gun head measuring jig S/N = | 3599
1166 | Enter measurement tools SN | Depth gauge S/N = | 3600
758 | SF_TS_HRQA, Surface Overview and Concentricity | SF_TS_HRQA - Passed inspection criteria, Image Saved | 3601
11173 | SF_TS_35um – SINGLE FILM APERTURE (SFA) Topside | Measure aperture size, 35um (+/- 2um) | 3602
11174 | Plat_TS_750um_HRQA, Surface Overview and 750um Centre Hole | Measure aperture size, 750um (+/- 15um) | 3603
11174 | Plat_TS_750um_HRQA, Surface Overview and 750um Centre Hole | Plat_TS_750um_HRQA - Passed inspection criteria, Image Saved | 3604
397 | SCS Check | Check SCS Folder/Sales Order | 3605
397 | SCS Check | SCS Number (If applicable) | 3606
223 | SCS Check | Check SCS Folder/Sales Order | 3607
223 | SCS Check | SCS Number (If applicable) | 3608
231 | SCS Check | Check SCS Folder/Sales Order | 3609
231 | SCS Check | SCS Number (If applicable) | 3610
1301 | Low kV Isolator Tracking | Record Low kV Isolator S/N from Lens Material Tracking Deviation (only applicable if Isolator has a S/N, if not then put NA. See attachment for example): | 3611
1301 | Low kV Isolator Tracking | Raise new Material Tracking Deviation (only applicable if Isolator has a S/N, if not then put NA): | 3612
1301 | Low kV Isolator Tracking | Record NEW Material Tracking Deviation Number (only applicable if Isolator has a S/N, if not then put NA): | 3613
377 | SCS Check | Check SCS Folder/Sales Order | 3614
377 | SCS Check | SCS Number (If applicable) | 3615
1651 | Detector Housing S/N | S/N = | 3616
95 | Fit Column Separation Valve | Operational repeatability test x 3 | 3617
11336 | Fit InLens Duo Detector | Serial: | 3618
11336 | Fit InLens Duo Detector | Remove metal protective cover from detector and inspect detector head | 3619
3295 | Fine positioning Inlens Detector | Serial: | 3620
3295 | Fine positioning Inlens Detector | Remove metal protective cover from detector and inspect detector head | 3621
1275 | Plasma Clean Parts | 20 min plasma cycle carried out - Ceramic Disc, Bean Can | 3622
1125 | Differential Pumping Aperture (Volutome) | Confirm differential aperture is long bore PN: 349520-0086-020 | 3623
102 | Differential Pumping Aperture Cleaning and Inspection | Aperture cleaned using IPA and 1mm brush | 3624
102 | Differential Pumping Aperture Cleaning and Inspection | Aperture run through ultra-sonic tank (in upright position) | 3625
102 | Differential Pumping Aperture Cleaning and Inspection | Aperture inspected and confimred clean. Eng 1 Initials: | 3626
102 | Differential Pumping Aperture Cleaning and Inspection | Aperture inspected and confimred clean. Eng 2 Initials: | 3627
1652 | Differential Pumping Stage S/N | S/N = | 3628
11242 | Inspect, Clean, & Fit The Liner Adapter | O-Rings TEM oiled | 3629
11354 | Thermocouple Test | UHV Thermocouple resistance is 110Ὡ (+/-10Ὡ) | 3630
144 | Mushroom Contact Spring Continuity Check | Continuity Check | 3631
149 | Closed Anode Pre-assembly | Confirm Configuration is correct for $[MaxCurrent]nA Column | 3632
148 | Open Anode Pre-assembly | Check Configuration is correct for $[MaxCurrent]nA Column. | 3633
150 | Closed Multihole Aperture Holder Pre-assembly | Check Configuration is correct for $[MaxCurrent]nA Column. | 3634
151 | Open Multihole Aperture Holder Pre-assembly | Check Configuration is correct for $[MaxCurrent]nA Column. | 3635
145 | Inspection of Liner tube | Checked By Assembly Engineer: | 3636
1660 | 2nd Inspection of Liner tube Different Eng, | Initials of 2nd Engineer confirming. | 3637
152 | Assemble Liner Tube | Aperture Orientation is correct for $[MaxCurrent]na Column | 3638
11385 | Fit Raith spray aperture | Confirm correct fitting of Raith spray aperture | 3639
157 | Fit Anode, Check Anode Parallelism | Postion A | 3640
157 | Fit Anode, Check Anode Parallelism | Postion B | 3641
157 | Fit Anode, Check Anode Parallelism | Postion C | 3642
157 | Fit Anode, Check Anode Parallelism | Postion D | 3643
157 | Fit Anode, Check Anode Parallelism | Max. Difference: | 3644
225 | Blanking cable check | Confirm that blanker cable routing is correct | 3645
142 | Plasma Clean UHV Chamber | Plasma cleaning completed | 3646
158 | Fit Extractor Assembly, Check Extractor Parallelsim | Position A | 3647
158 | Fit Extractor Assembly, Check Extractor Parallelsim | Position B | 3648
158 | Fit Extractor Assembly, Check Extractor Parallelsim | Position C | 3649
158 | Fit Extractor Assembly, Check Extractor Parallelsim | Position D | 3650
158 | Fit Extractor Assembly, Check Extractor Parallelsim | Max. Difference: | 3651
159 | Assemble Gun Head, Check Dummy Denka Parallelsim | Position A | 3652
159 | Assemble Gun Head, Check Dummy Denka Parallelsim | Position B | 3653
159 | Assemble Gun Head, Check Dummy Denka Parallelsim | Position C | 3654
159 | Assemble Gun Head, Check Dummy Denka Parallelsim | Position D | 3655
159 | Assemble Gun Head, Check Dummy Denka Parallelsim | Max. Difference: | 3656
160 | Denka Tip Shim Measurement | Denka Avg. X = | 3657
160 | Denka Tip Shim Measurement | Extractor Avg. Y = | 3658
160 | Denka Tip Shim Measurement | Denka to Ext. Distance: (Ext. D. Avg.) – (Denka H. Avg.) = | 3659
160 | Denka Tip Shim Measurement | Shim Value: Denka to Ext. Distance – 0.75 = | 3660
161 | Check Shimmed Dummy Denka Parallelsim | Position A | 3661
161 | Check Shimmed Dummy Denka Parallelsim | Position B | 3662
161 | Check Shimmed Dummy Denka Parallelsim | Position C | 3663
161 | Check Shimmed Dummy Denka Parallelsim | Position D | 3664
161 | Check Shimmed Dummy Denka Parallelsim | Max. Difference: | 3665
162 | Denka Tip Shim Re-Measurement | Tip Re-measured Height Avg: Z = | 3666
162 | Denka Tip Shim Re-Measurement | Actual Tip Height: Y - Z = | 3667
162 | Denka Tip Shim Re-Measurement | Tip Gap: Actual Tip Height – 0.75 = | 3668
162 | Denka Tip Shim Re-Measurement | Final Shim Value (Please ensure to always round up the shim value) = | 3669
162 | Denka Tip Shim Re-Measurement | Denka Torque Setting = | 3670
163 | Fit Cathode | Tip S/N | 3671
163 | Fit Cathode | 2nd Engineer to inspect live tip prior to fitting gun head (Eng Initials) | 3672
164 | Fit Gun Head | Replace the M8 nuts on the gun head ground with 2x Zinc equivalents | 3673
165 | Gun Head Checks | Filament OK (Green LED Lit) | 3674
165 | Gun Head Checks | Extractor - Suppressor ok, (No Red LED Lit) | 3675
165 | Gun Head Checks | Gun head driven to four corners (No Red LED Lit) | 3676
166 | Column Isolation Check | Liner In – Column GND (Open Circuit) | 3677
11144 | Fit Water Pipes | Vacuum Check (-15PSI @ 5 Mins) | 3678
11145 | Leak Check whole Column | Number of UHV Vent cycles during leak checking | 3679
11208 | All Heater cables isolation check | All cables Megger reads >1000Mohms at 1000V to ground | 3680
11266 | Bakeout cable routing | Flange Heater Earths routed correctly (where applicable) | 3681
11266 | Bakeout cable routing | Thermocouple cables routed correctly | 3682
168 | 250V Isolation Pre-fitment Checks | Driver A – Liner (>1000MΩ) | 3683
168 | 250V Isolation Pre-fitment Checks | Driver B – Liner (>1000MΩ) | 3684
168 | 250V Isolation Pre-fitment Checks | Driver A – Driver B (>1000MΩ) | 3685
11190 | 250V Isolation Pre-fitment Checks (Raith) | Driver A – Liner (>1000MΩ) | 3686
1129 | EBBU Isolator Pin Check | Ensure the EBBU pins do not protrude from the EBBU (See work instruction for details. | 3687
736 | If DC EBBU, Install Pumping adapter | Pumping adapter has been fitted before DC EBBU post fitment checks | 3688
171 | 250V Isolation Post-fitment Checks | Driver A – Liner (>1000MΩ) | 3689
171 | 250V Isolation Post-fitment Checks | Driver B – Liner (>1000MΩ) | 3690
171 | 250V Isolation Post-fitment Checks | Driver A – Driver B (>1000MΩ) | 3691
171 | 250V Isolation Post-fitment Checks | Driver A – GROUND (>1000MΩ) | 3692
171 | 250V Isolation Post-fitment Checks | Driver B – GROUND (>1000MΩ) | 3693
171 | 250V Isolation Post-fitment Checks | Liner – GROUND (>1000MΩ) | 3694
10821 | Fit RHS Column Cabling | Heater Cables | 3695
10821 | Fit RHS Column Cabling | Gun Cable | 3696
10821 | Fit RHS Column Cabling | Liner Cable | 3697
10821 | Fit RHS Column Cabling | Column Earth Cable | 3698
172 | EBBU Continuity Check | Driver A – Driver A In | 3699
172 | EBBU Continuity Check | Driver B – Driver B In | 3700
172 | EBBU Continuity Check | Anode – Liner In | 3701
11191 | EBBU Isolation Check (Raith) | Driver A In – Anode (>1000MΩ) | 3702
11191 | EBBU Isolation Check (Raith) | Driver A In – GND (>1000MΩ) | 3703
11191 | EBBU Isolation Check (Raith) | Anode - GND (>1000MΩ) | 3704
11192 | EBBU Continuity Check (Raith) | Liner In - Anode | 3705
11192 | EBBU Continuity Check (Raith) | Driver A - Driver A In | 3706
181 | Column Bake 48 Hours Ion Pumps off | Bake Time = | 3707
184 | Gun Conditioning Test when at final vacuum | Start Gun Vacuum Tracing | 3708
184 | Gun Conditioning Test when at final vacuum | Save Gun Vacuum Trace | 3709
185 | Final Gun Vaccum | IGP Pressure (≤ 9.0e-10 mbar) | 3710
11122 | Column Interventions | UHV Vent cycles - Gun conditioning | 3711
11122 | Column Interventions | UHV Vent cycles - Vacuum leak | 3712
1649 | Final Gun Vaccum - VP | IGP 1 Pressure (≤ 9.0e-10 mbar) | 3713
1649 | Final Gun Vaccum - VP | IGP 2 Pressure | 3714
11245 | Production and Supplier Paperwork | All Paperwork Sorted and Stapled Together | 3715
10816 | Fit Column | Correct Chamber Bung Fitted | 3716
416 | Vacuum Leak Test | No Leaks Detected (Prior to running up Turbo) | 3717
416 | Vacuum Leak Test | No Leaks Detected (With Turbo running) when in high sensitivity | 3718
416 | Vacuum Leak Test | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng | 3719
416 | Vacuum Leak Test | No Leaks Detected (With Turbo running) when in high sensitivity | 3720
416 | Vacuum Leak Test | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng | 3721
416 | Vacuum Leak Test | No Leaks Detected (With Turbo running) when in high sensitivity | 3722
416 | Vacuum Leak Test | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng | 3723
10818 | Fit Column Air & Water Pipes | Route air and water pipes | 3724
10818 | Fit Column Air & Water Pipes | No Leaks Detected (With Turbo running) when in high sensitivity | 3725
10818 | Fit Column Air & Water Pipes | No Leaks Detected (With Turbo running) when in high sensitivity | 3726
10818 | Fit Column Air & Water Pipes | Route air and water pipes | 3727
10818 | Fit Column Air & Water Pipes | Fit clamp | 3728
10818 | Fit Column Air & Water Pipes | No Leaks Detected (With Turbo running) when in high sensitivity | 3729
10818 | Fit Column Air & Water Pipes | Route air and water pipes | 3730
10818 | Fit Column Air & Water Pipes | Fit water pipe clamp | 3731
10818 | Fit Column Air & Water Pipes | No Leaks Detected (With Turbo running) when in high sensitivity | 3732
10818 | Fit Column Air & Water Pipes | Route air and water pipes | 3733
10818 | Fit Column Air & Water Pipes | Fit water pipe clamp | 3734
10819 | Fit LHS Column Cabling | Ion Pump Cables | 3735
10819 | Fit LHS Column Cabling | Column X1 Stig / Alignment Cable | 3736
10819 | Fit LHS Column Cabling | Objective Lens Cable | 3737
10819 | Fit LHS Column Cabling | Scan Cable | 3738
10819 | Fit LHS Column Cabling | Penning Cable | 3739
10819 | Fit LHS Column Cabling | Ion Pump Cables | 3740
10819 | Fit LHS Column Cabling | Column X1 Stig / Alignment Cable | 3741
10819 | Fit LHS Column Cabling | Objective Lens Cable | 3742
10819 | Fit LHS Column Cabling | Scan Cable | 3743
10821 | Fit RHS Column Cabling | Heater Cables | 3744
10821 | Fit RHS Column Cabling | EHT Cables | 3745
10821 | Fit RHS Column Cabling | Liner Cable | 3746
10821 | Fit RHS Column Cabling | Column Earth Cable | 3747
10822 | Fit Ion Pump PSU (HV) | Bracket | 3748
10822 | Fit Ion Pump PSU (HV) | Supply | 3749
10822 | Fit Ion Pump PSU (HV) | Heater Blocks | 3750
1325 | Fit Ion Pump PSU (VP) | Bracket | 3751
1325 | Fit Ion Pump PSU (VP) | Supply | 3752
1325 | Fit Ion Pump PSU (VP) | Heater Blocks | 3753
10822 | Fit Ion Pump PSU (HV) | Fit IPPSU Bracket | 3754
10822 | Fit Ion Pump PSU (HV) | Supply | 3755
10822 | Fit Ion Pump PSU (HV) | Heater Blocks | 3756
10822 | Fit Ion Pump PSU (HV) | Fit IPPSU Bracket | 3757
10822 | Fit Ion Pump PSU (HV) | Fit Power Supply | 3758
10822 | Fit Ion Pump PSU (HV) | Fit Heater Blocks | 3759
1325 | Fit Ion Pump PSU (VP) | Fit 2X IPPSU Bracket | 3760
1325 | Fit Ion Pump PSU (VP) | Fit Power Supply | 3761
1325 | Fit Ion Pump PSU (VP) | Fit Heater Blocks | 3762
1325 | Fit Ion Pump PSU (VP) | Fit 2X IPPSU Bracket | 3763
1325 | Fit Ion Pump PSU (VP) | Fit 2X Power Supply | 3764
1325 | Fit Ion Pump PSU (VP) | Fit Heater Blocks | 3765
1325 | Fit Ion Pump PSU (VP) | Fit 2X IPPSU Bracket | 3766
1325 | Fit Ion Pump PSU (VP) | Fit 2X Power Supply | 3767
1325 | Fit Ion Pump PSU (VP) | Fit Heater Blocks | 3768
10822 | Fit Ion Pump PSU (HV) | Fit IPPSU Bracket | 3769
10822 | Fit Ion Pump PSU (HV) | Fit Power Supply | 3770
10822 | Fit Ion Pump PSU (HV) | Fit Heater Blocks | 3771
10823 | Fit EHT Set | 24V Supply Kit | 3772
10823 | Fit EHT Set | Spellman EHT Set | 3773
10823 | Fit EHT Set | EHT Set Serial Number | 3774
10823 | Fit EHT Set | Earth Connection Torqued to 3Nm | 3775
10823 | Fit EHT Set | Fit 24V Supply Kit to EHT SET | 3776
10823 | Fit EHT Set | Fit Spellman EHT Set | 3777
10823 | Fit EHT Set | EHT Set Serial Number | 3778
10823 | Fit EHT Set | Earth Connection Torqued to 3Nm | 3779
10823 | Fit EHT Set | Fit 24V Supply Kit to EHT SET | 3780
10823 | Fit EHT Set | Fit Spellman EHT Set | 3781
10823 | Fit EHT Set | EHT Set Serial Number | 3782
10823 | Fit EHT Set | Earth Connection Torqued to 3Nm | 3783
10823 | Fit EHT Set | Connect gun cable | 3784
10823 | Fit EHT Set | Fit 24V Supply Kit to EHT SET | 3785
10823 | Fit EHT Set | Fit Spellman EHT Set | 3786
10823 | Fit EHT Set | EHT Set Serial Number | 3787
10823 | Fit EHT Set | Earth Connection Torqued to 3Nm | 3788
10823 | Fit EHT Set | Connect Gun Cable | 3789
10823 | Fit EHT Set | Connect Liner Cable | 3790
10730 | Fit LaB6 Cabling | Valves Cables | 3791
10730 | Fit LaB6 Cabling | Heater Cables | 3792
10730 | Fit LaB6 Cabling | Ion Pump Cable | 3793
10730 | Fit LaB6 Cabling | Ion Pump PSU Serial Number | 3794
10730 | Fit LaB6 Cabling | Fit Ion Pump PSU Bracket | 3795
10730 | Fit LaB6 Cabling | Fit Ion Pump PSU | 3796
10730 | Fit LaB6 Cabling | PSU Qualification Marks Checked | 3797
10731 | Fit EHT Set | EHT Set Type Fitted | 3798
10731 | Fit EHT Set | EHT Serial Number | 3799
10731 | Fit EHT Set | Revision/Mod Level | 3800
10731 | Fit EHT Set | Qualification Marks Checked | 3801
10731 | Fit EHT Set | Torque Earth Connection to 3Nm | 3802
10732 | System Cabling | All Cables Are Neatly Routed | 3803
10732 | System Cabling | All Plugs Are Secured Correctly | 3804
10732 | System Cabling | All Earth Cables Are Tightly Secured | 3805
10732 | System Cabling | EHT Set Is Secured | 3806
11412 | PAT Tester Verification Test (Seaward Supernova + PAT Checkbox) | PAT Checkbox Calibration Date | 3807
11412 | PAT Tester Verification Test (Seaward Supernova + PAT Checkbox) | Earth Bond Test TP1 (0.1Ω - 0.07Ω) | 3808
11412 | PAT Tester Verification Test (Seaward Supernova + PAT Checkbox) | Earth Bond Test TP2 (0.55Ω - 0.52Ω) | 3809
11412 | PAT Tester Verification Test (Seaward Supernova + PAT Checkbox) | Insulation Test (2.0MΩ - 1.9MΩ) | 3810
11412 | PAT Tester Verification Test (Seaward Supernova + PAT Checkbox) | Earth Leakage Test (0.9mA - 0.7mA) | 3811
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | X Confirm SN label does not have any CSA markings. | 3812
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | ✔ Confirm SN label does have the correct SEMI marking. | 3813
1234 | Print and Fit S/N Label (CSA UL certification) | ✔ Confirm SN label does have the correct CSA marking. | 3814
1234 | Print and Fit S/N Label (CSA UL certification) | X Confirm SN label does not have any SEMI marking. | 3815
1235 | Print and Fit S/N Label ( SEMI S2 S8 certification + CSA UL certification) | ✔Confirm SN label does have the correct CSA marking. | 3816
1235 | Print and Fit S/N Label ( SEMI S2 S8 certification + CSA UL certification) | ✔ Confirm SN label does have the correct SEMI marking. | 3817
1235 | Print and Fit S/N Label ( SEMI S2 S8 certification + CSA UL certification) | ✔ Confirm SN label does have the correct CSA marking. | 3818
1235 | Print and Fit S/N Label ( SEMI S2 S8 certification + CSA UL certification) | ✔ Confirm SN label does have the correct SEMI marking. | 3819
1236 | Print and Fit SN Label (NO CSA, NO SEMI) | X Confirm SN label does not have any CSA markings. | 3820
1236 | Print and Fit SN Label (NO CSA, NO SEMI) | X Confirm SN label does not have any SEMI markings. | 3821
1236 | Print and Fit SN Label (NO CSA, NO SEMI) | X Confirm NO CSA markings. | 3822
1236 | Print and Fit SN Label (NO CSA, NO SEMI) | X Confirm NO SEMI markings. | 3823
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | X Confirm NO CSA markings. | 3824
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | ✔ Confirm SN label does have the correct SEMI marking. | 3825
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | X Confirm NO CSA markings. | 3826
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | ✔ Confirm DOES HAVE SEMI marking. | 3827
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | X Confirm NO CSA markings. | 3828
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | ✔ Confirm SEMI marking on SN label. | 3829
316 | Initial Resolution Images | Actual for 1kV = | 3830
316 | Initial Resolution Images | Store 1KV Resolution Images and Pate image | 3831
316 | Initial Resolution Images | Actual for 15kV = | 3832
316 | Initial Resolution Images | Store 15KV Resolution Images and Pate image | 3833
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | Confirm SN label does not have any CSA markings. | 3834
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | ✔ SEMI marking is on SN label. | 3835
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | ✘ No CSA marking on SN label. | 3836
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | ✔ SEMI marking is on SN label. | 3837
1236 | Print and Fit SN Label (NO CSA, NO SEMI) | ✘ No CSA marking on SN label.. | 3838
1236 | Print and Fit SN Label (NO CSA, NO SEMI) | ✘ No SEMI marking on SN label. | 3839
10647 | Check Calibration Figures in SEM Status are acceptable | Confirm SEM Status that (C1 Double only), C2 & C3 Cal are reasonable | 3840
10647 | Check Calibration Figures in SEM Status are acceptable | System Align Offset Angle Approx: (-21 for X-Beam) (-51 for Sphinx) (-110 for Raith) (N/A for D/C) | 3841
1236 | Print and Fit SN Label (NO CSA, NO SEMI) | ✘ 𝗡𝗢 CSA marking on SN label.. | 3842
1236 | Print and Fit SN Label (NO CSA, NO SEMI) | ✘ 𝗡𝗢 SEMI marking on SN label. | 3843
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | ✘ 𝗡𝗢 CSA marking on SN label. | 3844
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | ✔ SEMI marking is on SN label. | 3845
1234 | Print and Fit S/N Label (CSA UL certification) | ✔ CSA marking is on SN label. | 3846
1234 | Print and Fit S/N Label (CSA UL certification) | Confirm SN label does not have any SEMI marking. | 3847
1234 | Print and Fit S/N Label (CSA UL certification) | ✔ CSA marking is on SN label. | 3848
1234 | Print and Fit S/N Label (CSA UL certification) | ✘ 𝗡𝗢 SEMI marking on SN label. | 3849
1235 | Print and Fit S/N Label ( SEMI S2 S8 certification + CSA UL certification) | ✔ CSA marking is on SN label. | 3850
1235 | Print and Fit S/N Label ( SEMI S2 S8 certification + CSA UL certification) | Confirm SN label does have the correct SEMI marking. | 3851
1235 | Print and Fit S/N Label ( SEMI S2 S8 certification + CSA UL certification) | ✔ CSA marking is on SN label. | 3852
1235 | Print and Fit S/N Label ( SEMI S2 S8 certification + CSA UL certification) | ✔ SEMI marking is on SN label. | 3853
3252 | EO P.C.B Versions | F/W Version NOTE MUST BE VERSION 2 IF HITECH STAGE IS FITTED | 3854
3252 | EO P.C.B Versions | Hardware Version | 3855
10831 | Indigo Filter Upgrade | Confirm Indigo Filter has been fitted to CL detector | 3856
11243 | Install BSD1 | Configure compressed airline for detector mechanics | 3857
11243 | Install BSD1 | Configure compressed airline for pneumatic detector mechanics | 3858
1271 | Install BSD1 VP | Configure compressed airline for pneumatic detector mechanics | 3859
10829 | Install BSD Push/Pull Assembly | Fit Adaptor Plate to Chamber | 3860
10829 | Install BSD Push/Pull Assembly | Fit Amplifier | 3861
10829 | Install BSD Push/Pull Assembly | Set up Rod Distance to 75mm | 3862
10829 | Install BSD Push/Pull Assembly | Fit Push Pull Assembly | 3863
10829 | Install BSD Push/Pull Assembly | Set up Rod Distance to 75mm | 3864
10829 | Install BSD Push/Pull Assembly | Fit Push Pull Assembly | 3865
1358 | Install BSD4 | Configure compressed airline for pneumatic detector mechanics | 3866
3252 | EO P.C.B Versions | F/W Version NOTE MUST BE VERSION 2 IF HITTEK STAGE IS FITTED | 3867
3252 | EO P.C.B Versions | Hardware Version | 3868
10647 | Check Calibration Figures in SEM Status are acceptable | Confirm in SEM Status that (C1 Double only), C2 & C3 Cal are reasonable | 3869
10647 | Check Calibration Figures in SEM Status are acceptable | System Align Offset Angle Approx: (-21 for X-Beam) (-51 for Sphinx) (-110 for Raith) (N/A for D/C) | 3870
10647 | Check Calibration Figures in SEM Status are acceptable | Confirm in SEM Status that Correction Angle and Correction Factors (8 values) are not zero | 3871
10832 | Install VPSE G4 Detector | Ensure Indigo Filter is fitted. | 3872
10834 | Install C2D Detector | Set C2D length to positioin 2. | 3873
10843 | Install STEM Detector | Configure compressed airline for pneumatic detector mechanics | 3874
1549 | Insight | Objective Cap - No particles visible | 3875
1549 | Insight | Lower Aperture - No large particles visible (See WI for details) | 3876
1549 | Insight | Upper Aperture - No cut-off on image and no particles visible. | 3877
329 | Stigmation / Aperture Checks | Stig: 30μm, All EHT’s are ≤ 20% | 3878
329 | Stigmation / Aperture Checks | Stig: Off Axis Aper, All EHT’s are ≤ 30% | 3879
329 | Stigmation / Aperture Checks | Aper/Gun: 30μm, All EHT’s (Excluding 99v) are ≤ 60% | 3880
10935 | Route Detector Cables | Add Detector Labels | 3881
10935 | Route Detector Cables | Add Detector Labels | 3882
10935 | Route Detector Cables | Order Detecors in H-Asap as per configurator | 3883
10935 | Route Detector Cables | Add Detector Labels | 3884
10935 | Route Detector Cables | Order Detecors in H-ASAP as per configurator | 3885
10838 | Install Specimen Current Detector | Install SCD Amplifier | 3886
10838 | Install Specimen Current Detector | Fit SCD Relay PCB - Sigma500 Only | 3887
10838 | Install Specimen Current Detector | Install SCD Amplifier | 3888
10842 | Install Airlock 80 | Install Airlock FW Version 2.0.0 | 3889
10841 | Install Airlock 200 | Install Airlock FW Version 2.0.0 | 3890
10827 | Fit Stage Cabling | Stage Cable | 3891
10827 | Fit Stage Cabling | Specimen Touch Cable | 3892
11213 | Widefield Mode Calibration | Calibration wizard completed. | 3893
11071 | Mag Cal Accuracy check (G500/560) | Check Mag cal errors are <2% in X & Y axes in all 3 Mag Ranges. Orthog. ± 0.1deg | 3894
11071 | Mag Cal Accuracy check (G500/560) | Store accuracy images (Mag 1, Mag 2, Mag 3 and Orthog) | 3895
11071 | Mag Cal Accuracy check (G560/XB750) | Check Mag cal errors are <2% in X & Y axes in all 3 Mag Ranges. Orthog. ± 0.1deg | 3896
11071 | Mag Cal Accuracy check (G560/XB750) | Store accuracy images (Mag 1, Mag 2, Mag 3 and Orthog) | 3897
11069 | Dynamic Focus Test | Dynamic focus function check - Whole image in focus | 3898
11069 | Dynamic Focus Test | Y - Axis measurement (400μm +/- 2%) | 3899
11069 | Dynamic Focus Test | Orthogonality | 3900
281 | ESB concentricity & hole diameter check (standard objective) | Check concentricity of ESB wrt scanning system. | 3901
281 | ESB concentricity & hole diameter check (standard objective) | Record the ESB serial number - This is the S/N that is shown on the ESB label (1026-XXX or 1125-XXX). | 3902
281 | ESB concentricity & hole diameter check (standard objective) | Measure the ESB hole diameter. | 3903
281 | ESB concentricity & hole diameter check (standard objective) | Store ESB Concentricity image. Add ESB serial number as a text annotation in SmartSEM. (filename must be exactly esbdiameter.tif) | 3904
396 | ESB concentricity & hole diameter check (low kV objective) | Check concentricity of ESB wrt scanning system. (Limit: <= 50um) | 3905
396 | ESB concentricity & hole diameter check (low kV objective) | Measure the ESB hole diameter. | 3906
396 | ESB concentricity & hole diameter check (low kV objective) | Record the ESB serial number - This is the S/N that is shown on the ESB label (1026-XXX or 1125-XXX). | 3907
396 | ESB concentricity & hole diameter check (low kV objective) | Store ESB Concentricity image. Add ESB serial number as a text annotation in SmartSEM. (filename must be exactly esbdiameter.tif) | 3908
1591 | Magnification Calibration | Take and save images at Mag 200, 2kX and 35kX. | 3909
1591 | Magnification Calibration | Mag is within +/- 1% for each. | 3910
1591 | Magnification Calibration | Orthognality is within +/- 0.1 deg. | 3911
1568 | System Shutdown Gemini | Remove docking brackets and fit transit brackets | 3912
1568 | System Shutdown Gemini | Power system down. | 3913
1568 | System Shutdown Gemini | Blow water out of the cooling system. | 3914
1568 | System Shutdown Gemini | Put thread lock on pendulum screws and transit block Screws. | 3915
1568 | System Shutdown Gemini | Chamber secured to vibration plate | 3916
1568 | System Shutdown Gemini | Chamber Sercuring Screws Tightened to 25NM | 3917
1568 | System Shutdown Gemini | Chamber Sercuring Screws Tightness Checked By Name: | 3918
1568 | System Shutdown Gemini | Secure Airfeet with foam blocks. | 3919
1568 | System Shutdown Gemini | Fit foam blocks around column cladding. | 3920
1568 | System Shutdown Gemini | Secure Turbo pump with foam blocks. | 3921
1568 | System Shutdown Gemini | Fit Covers to all unused connectors. | 3922
1568 | System Shutdown Gemini | Remove Cekon Plug From Mains Cable (Except Polo + Systems) | 3923
1568 | System Shutdown Gemini | Fit Live and Neutral Labels to Mains Cable (Except Polo + Systems) | 3924
1568 | System Shutdown Gemini | Mains and Earth cable stowed on Mains unit. | 3925
1568 | System Shutdown Gemini | Fit shipping Eyes to transit blocks. | 3926
1568 | System Shutdown Gemini | Confirm all column cladding is secure and screws Loctited in place. | 3927
1568 | System Shutdown Gemini | All doors closed and secured. | 3928
1568 | System Shutdown Gemini | All plinth panels removed and stored in transit trolley. | 3929
1568 | System Shutdown Gemini | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel). | 3930
1568 | System Shutdown Gemini | Fit M8 screw to secure stage door to chamber. | 3931
1568 | System Shutdown Gemini | Fit dust covers and secure / remove cables. | 3932
1568 | System Shutdown Gemini | Prepare column cladding transport jig . | 3933
1568 | System Shutdown Gemini | Print shipping stickers . | 3934
1592 | Magnification Calibration | Take and save images at Mag 200, 3kX and 35kX. | 3935
1592 | Magnification Calibration | Mag is within +/- 2% for each. | 3936
1592 | Magnification Calibration | Orthognality is within +/- 0.1 deg | 3937
11123 | System Shutdown Crossbeam | Remove docking brackets and fit transit brackets | 3938
11123 | System Shutdown Crossbeam | Power system down. | 3939
11123 | System Shutdown Crossbeam | Blow water out of the cooling system. | 3940
11123 | System Shutdown Crossbeam | Put thread lock on pendulum screws and transit block Screws. | 3941
11123 | System Shutdown Crossbeam | Chamber secured to vibration plate (N/A to XB 550L) | 3942
11123 | System Shutdown Crossbeam | Chamber securing screws tightened to 25nm (N/A to XB 550L) | 3943
11123 | System Shutdown Crossbeam | Chamber securing screws tightness checked. Type your initials (For XB 550L type N/A) | 3944
11123 | System Shutdown Crossbeam | Secure Airfeet with foam blocks. | 3945
11123 | System Shutdown Crossbeam | Secure Turbo pump with foam blocks. | 3946
11123 | System Shutdown Crossbeam | Fit Covers to all unused connectors. | 3947
11123 | System Shutdown Crossbeam | Mains and Earth cable stowed on Mains unit. | 3948
11123 | System Shutdown Crossbeam | Fit shipping Eyes to transit blocks. | 3949
11123 | System Shutdown Crossbeam | All doors closed and secured. | 3950
11123 | System Shutdown Crossbeam | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel). | 3951
11123 | System Shutdown Crossbeam | Print shipping stickers. | 3952
11123 | System Shutdown Crossbeam | Fit 6x M8x16mm cap head screws to stage door | 3953
1250 | Stage Repeatability Macro - Hittech Stage Only | Download Macro folder | 3954
1250 | Stage Repeatability Macro - Hittech Stage Only | Run Kompass 3.08.exe | 3955
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 1 ($[StageRepeatability]) | 3956
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 2 ($[StageRepeatability]) | 3957
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 3 ($[StageRepeatability]) | 3958
1250 | Stage Repeatability Macro - Hittech Stage Only | Image of Macro Results | 3959
10516 | Stage Acceptance Check X and Y Axis - Cartesian Stages Only | X Axis | 3960
10516 | Stage Acceptance Check X and Y Axis - Cartesian Stages Only | Y Axis | 3961
10516 | Stage Acceptance Check X and Y Axis - Cartesian Stages Only | X+Y Axis | 3962
10516 | Stage Acceptance Check X and Y Axis - Cartesian Stages Only | Y+X Axis | 3963
10516 | Stage Acceptance Check X and Y Axis - Cartesian Stages Only | 10mm Square | 3964
10517 | Stage Acceptance Check Z Axis - Cartesian Stages Only | Z Axis | 3965
10518 | Stage Backlash Measurement X & Y - Cartesian Stages Only | X Axis | 3966
10518 | Stage Backlash Measurement X & Y - Cartesian Stages Only | Y Axis | 3967
1250 | Stage Repeatability Macro - Hittech Stage Only | Download Macro folder | 3968
1250 | Stage Repeatability Macro - Hittech Stage Only | Run Kompass 3.08.exe | 3969
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 1 ($[StageRepeatability]) | 3970
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 2 ($[StageRepeatability]) | 3971
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 3 ($[StageRepeatability]) | 3972
1250 | Stage Repeatability Macro - Hittech Stage Only | Image of Macro Results | 3973
214 | Photograph Oxford Sales Number Label | Take Photo of Label | 3974
214 | Photograph Oxford Sales Number Label | Oxford Sales Order | 3975
214 | Photograph Oxford Sales Number Label | Store Oxford Sales Order Label Images | 3976
1250 | Stage Repeatability Macro - Hittech Stage Only | Download Macro folder | 3977
1250 | Stage Repeatability Macro - Hittech Stage Only | Run Kompass 3.08.exe | 3978
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 1 (MAX 5um) | 3979
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 2 (MAX 5um) | 3980
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 3 (MAX 5um) | 3981
1250 | Stage Repeatability Macro - Hittech Stage Only | Image of Macro Results | 3982
6928 | Stage Repeatability Macro | Select Stage Type | 3983
6928 | Stage Repeatability Macro | Macro Result 1 (Max 5) | 3984
6928 | Stage Repeatability Macro | Macro Result 2 (Max 5) | 3985
6928 | Stage Repeatability Macro | Macro Result 3 (Max 5) | 3986
6928 | Stage Repeatability Macro | Image of Macro Results | 3987
3242 | PC Installation | Select PC type | 3988
3242 | PC Installation | PC S/No. | 3989
3242 | PC Installation | Is SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, | 3990
3242 | PC Installation | Hard Drive 1 Size | 3991
3242 | PC Installation | Hard Drive 2 Size | 3992
3242 | PC Installation | RAM | 3993
3242 | PC Installation | Processor | 3994
3242 | PC Installation | Operating System | 3995
3242 | PC Installation | Change Desktop Settings | 3996
3242 | PC Installation | Install SmartSEM and record Version | 3997
3242 | PC Installation | Check for Service Packs and install them | 3998
3242 | PC Installation | Upgrade ASAP Firmware (check Service Pack for latest version) and API | 3999
3242 | PC Installation | Copy the SmartSEM sotware into a Folder in C:Support and Name it SmartSEM X.x (i.e 8.2) Software | 4000
3242 | PC Installation | Installation of Zen Core EM if Hasap is fitted | 4001
3242 | PC Installation | Set BIOS AC Power loss to ‘ON’ | 4002
3242 | PC Installation | Set the Time and Date on the PC to current UK | 4003
3242 | PC Installation | Attach Cable Markers to USB Cables | 4004
8699 | Cable Up | System Mains Voltage | 4005
8699 | Cable Up | IF LaB6 and 120V - Transformer Installed | 4006
8699 | Cable Up | Cable Up | 4007
8699 | Cable Up | Label Cables | 4008
8699 | Cable Up | Tape PC Configuration Image to PC | 4009
10862 | Chamber & Table Top | Chamber is Free From Defects | 4010
10862 | Chamber & Table Top | Chamber Pumping Pipe Is Fitted Centrally To The Table Top | 4011
10862 | Chamber & Table Top | Table Top Is Free From Defects | 4012
348 | Basic Interlock tests | Chamber Door Interlock | 4013
348 | Basic Interlock tests | Service Switch | 4014
348 | Basic Interlock tests | Lens De-activation | 4015
348 | Basic Interlock tests | Detector De-activation | 4016
348 | Basic Interlock tests | EHT + IPPSU De-activation | 4017
349 | Earth Continuity Tests. | Record SEAWARD PRIMETEST 100 SN: | 4018
349 | Earth Continuity Tests. | Calibration Due Date: | 4019
349 | Earth Continuity Tests. | Tester verification test; Confirm earth short circuit = (<0.03Ω) | 4020
349 | Earth Continuity Tests. | Test all earth points as shown in procedure. (Max = 0.1Ω) Record max resistance shown | 4021
349 | Earth Continuity Tests. | Confirm correct SN label applied to product: (CSA Compliant / Non-CSA compliant) | 4022
354 | Install Customer Licences | Delete the Production Test Licences | 4023
354 | Install Customer Licences | Obtain the customer Entitlement Certificate | 4024
354 | Install Customer Licences | Download into the Engineer folder | 4025
354 | Install Customer Licences | Install the Customer Licences | 4026
354 | Install Customer Licences | Copy Contents of Image Folder to the Product Folder | 4027
354 | Install Customer Licences | Copy EMServer.db to the Product Folder | 4028
354 | Install Customer Licences | Print out the Customer Entitlement Certicate - To be sent to the customer | 4029
356 | Hardware Dongle Check | Check System And Dongle S/No. Match In Service Panel | 4030
356 | Hardware Dongle Check | Un-tick Omit Dongle In Factory Panel | 4031
356 | Hardware Dongle Check | Check Dongle fits in USB HUB (“ICY BOX”) and replace if needed. | 4032
1596 | Customer licenses | Check that the licences installed in the Zeiss Licence Management Tool are correct to the Sales Order and that they functio | 4033
371 | Generate the following Documents (Gemini) | X-Ray Certificate | 4034
371 | Generate the following Documents (Gemini) | Quality Certificate | 4035
371 | Generate the following Documents (Gemini) | Customer Entitlement Certificate | 4036
371 | Generate the following Documents (Gemini) | Quality Feedback Form | 4037
371 | Generate the following Documents (Gemini) | DENKA Data Sheet | 4038
371 | Generate the following Documents (Gemini) | Options Removed packing List | 4039
371 | Generate the following Documents (Gemini) | Options Retained packing List | 4040
371 | Generate the following Documents (Gemini) | Standard Packing List | 4041
371 | Generate the following Documents (Gemini) | PC Inferface assignment | 4042
371 | Generate the following Documents (Gemini) | Declaration of conformity | 4043
371 | Generate the following Documents (Gemini) | Installation report | 4044
371 | Generate the following Documents (Gemini) | Acceptance test record | 4045
371 | Generate the following Documents (Gemini) | Licence Generator report | 4046
10467 | Install SmartSEM + ASAP Engine | Install SmartSEM + ASAP Engine | 4047
10467 | Install SmartSEM + ASAP Engine | Copy the SmartSEM sotware into the SmartSEM Software folder (C:Support/Data/SmartSEM Software) | 4048
10467 | Install SmartSEM + ASAP Engine | SmartSEM Version | 4049
10467 | Install SmartSEM + ASAP Engine | SmartSEM Service Pack | 4050
10467 | Install SmartSEM + ASAP Engine | SmartSEM Hotfix | 4051
10467 | Install SmartSEM + ASAP Engine | ASAP Driver Scan Control Service Version | 4052
10467 | Install SmartSEM + ASAP Engine | Installation of Zen Core EM | 4053
10671 | Fit Stage PCB (Klein / HiTech) | Stage PCB S/No | 4054
10671 | Fit Stage PCB (Klein / HiTech) | Fit Stage Light Barrier PCB | 4055
10671 | Fit Stage PCB (Klein / HiTech) | All Connections Are Secured with Screws | 4056
10671 | Fit Stage PCB (Klein / HiTech) | Cables Are Neatly Routed and Secured | 4057
3291 | Ripple (Target $[MaxRipple]nm) | X = | 4058
3291 | Ripple (Target $[MaxRipple]nm) | Store X Ripple Images and Ripplelyzer image for 5 x consecutive images | 4059
3291 | Ripple (Target $[MaxRipple]nm) | Y = | 4060
3291 | Ripple (Target $[MaxRipple]nm) | Store Y Ripple Images and Ripplelyzer image for 5 x consecutive images | 4061
396 | ESB concentricity & hole diameter check (low kV objective) | Check concentricity of ESB wrt scanning system. (Limit: <= 50um) | 4062
396 | ESB concentricity & hole diameter check (low kV objective) | Measure the ESB hole diameter. | 4063
396 | ESB concentricity & hole diameter check (low kV objective) | Record the ESB serial number - This is the S/N that is shown on the ESB label (1026-XXX or 1125-XXX). | 4064
396 | ESB concentricity & hole diameter check (low kV objective) | Store ESB Concentricity image. Add ESB serial number as a text annotation in SmartSEM. (filename must be exactly esbdiameter.tif) | 4065
396 | ESB concentricity & hole diameter check (low kV objective) | Check concentricity of ESB wrt scanning system. (Limit: <= 50um) | 4066
396 | ESB concentricity & hole diameter check (low kV objective) | Measure the ESB hole diameter. | 4067
396 | ESB concentricity & hole diameter check (low kV objective) | Record the ESB serial number - This is the S/N that is shown on the ESB label (1026-XXX or 1125-XXX). | 4068
396 | ESB concentricity & hole diameter check (low kV objective) | Store ESB Concentricity image. Add ESB serial number as a text annotation in SmartSEM. (filename must be exactly esbdiameter.tif) | 4069
759 | Run EO Test From Service Panel | Run EO Test | 4070
11077 | ZEN | ZEN Software Installed | 4071
10926 | ZEN | ZEN Software Installed | 4072
86 | VP-operation (No Beamsleeve) | Chamber vaccuum tracks demand | 4073
10671 | Fit Stage PCB (Klein / HiTech) | Stage PCB S/No | 4074
10671 | Fit Stage PCB (Klein / HiTech) | Fit Stage Light Barrier PCB | 4075
10671 | Fit Stage PCB (Klein / HiTech) | All Connections Are Secured with Screws | 4076
10671 | Fit Stage PCB (Klein / HiTech) | Cables Are Neatly Routed and Secured | 4077
1319 | Fit Stage PCB (Cartesian) | Stage PCB S/No | 4078
1319 | Fit Stage PCB (Cartesian) | All Connections Are Secured with Screws | 4079
1319 | Fit Stage PCB (Cartesian) | Cables Are Neatly Routed and Secured | 4080
10567 | Install Customer Licences | Delete the Production Test Licences | 4081
10567 | Install Customer Licences | Obtain the customer Entitlement Certificate | 4082
10567 | Install Customer Licences | Download into the C:Support/Data/Licence folder | 4083
10567 | Install Customer Licences | Install the Customer Licences | 4084
10567 | Install Customer Licences | Copy Contents of Image Folder to the Product Folder | 4085
10567 | Install Customer Licences | Copy EMServer.db to the Product Folder | 4086
10567 | Install Customer Licences | Copy EMServer.db to the C: Support/Data/Log Files folder | 4087
10567 | Install Customer Licences | Print out the Customer Entitlement Certicate - To be sent to the customer | 4088
10567 | Install Customer Licences | Store EM Server.db | 4089
10567 | Install Customer Licences | Copy EMServer.log to the Product Folder | 4090
10567 | Install Customer Licences | Store EM Server Log | 4091
10567 | Install Customer Licences | Copy EMServer Screening.txt to the Product Folder | 4092
10800 | Fit HiTech Stage | Confirm stage door is set up correctly, adjust if neccesary. | 4093
10567 | Install Customer Licences | Delete the Production Test Licences | 4094
10567 | Install Customer Licences | Obtain the customer Entitlement Certificate | 4095
10567 | Install Customer Licences | Download into the C:Support/Data/Licence folder | 4096
10567 | Install Customer Licences | Install the Customer Licences | 4097
10567 | Install Customer Licences | Copy Contents of Image Folder to the Product Folder | 4098
10567 | Install Customer Licences | Copy EMServer.db to the Product Folder | 4099
10567 | Install Customer Licences | Copy EMServer.db to the C: Support/Data/Log Files folder | 4100
10567 | Install Customer Licences | Print out the Customer Entitlement Certicate - To be sent to the customer | 4101
10567 | Install Customer Licences | Store EM Server.db | 4102
10567 | Install Customer Licences | Copy EMServer.log to the Product Folder | 4103
10567 | Install Customer Licences | Store EM Server Log | 4104
1392 | Stage Links | Stage Links Set Correctly | 4105
6927 | Stage Axis Cycle Test | Macro run time | 4106
916 | Volutome | 1. Unpack Volutome and check all parts on the packing list are present (packing list link can be found in attachments) | 4107
916 | Volutome | 1. Upload data from USB stick to the O drive product folder | 4108
916 | Volutome | 1. Remove transit clamps from Volutome stage | 4109
916 | Volutome | 1. Ensure both safety switches are fitted to the Volutome door | 4110
916 | Volutome | 2. Ensure hardware has been configured according to work instructions | 4111
916 | Volutome | 2. Fit Volutome stage | 4112
916 | Volutome | 2. Check protection circuit for Piezo is fitted (349567-9013-000) | 4113
916 | Volutome | 2. Connect Volutome stage and controller to the SEM | 4114
916 | Volutome | 2. Fit Stemi holder bracket to Volutome stage door (STEMI adjustment bracket needs to fit / disconnect without force from door) | 4115
916 | Volutome | 3. Ethernet Connection Setup | 4116
916 | Volutome | 4. Install the latest version of Zen - Ensure to configure it for Volutome (Refer to Zen installation WI) | 4117
916 | Volutome | 4. Ensure Athena inspector manager has all green lights, Physik microtome and Physik stage are present - save a screenshot in the Volutome folder | 4118
916 | Volutome | 4. Install PI Micromove | 4119
916 | Volutome | 4. Run getversions.exe and save text file in the Volutome folder | 4120
916 | Volutome | 4. Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator | 4121
916 | Volutome | 4. Set up the Volutome workbench in Zen | 4122
916 | Volutome | 5. PID parameters configured | 4123
916 | Volutome | 6. Record VolumeBSD serial number | 4124
916 | Volutome | 6. Record Volutome stage serial number | 4125
916 | Volutome | 6. Record VolumeBSD amplifier serial number | 4126
916 | Volutome | 6. Record VolumeBSD mechanics serial number | 4127
916 | Volutome | 6. Record C-885 PI controller serial number | 4128
916 | Volutome | 6. Record E-727 PI controller serial number | 4129
916 | Volutome | 7. Volutome 'internal LED light' function check | 4130
916 | Volutome | 8. Volutome 'External cold light' function check | 4131
916 | Volutome | 9. Adjust Volutome stage centre (Crosshairs should be within 4 central squares) | 4132
916 | Volutome | 10. Install Nav cam in-line connections (If applicable) | 4133
916 | Volutome | 11. Perform the microtome and stage functionality checks | 4134
916 | Volutome | 12. Fit and align the VolumeBSD diode | 4135
916 | Volutome | 12. Check VolumeBSD alignment repeatability | 4136
916 | Volutome | 12. Check all gain ranges of the VolumeBSD (normal/inverted) | 4137
916 | Volutome | 12. Perform Signal/noise ratio check | 4138
916 | Volutome | 12. Ensure BSD diode protector is fitted and adjusted in chamber | 4139
916 | Volutome | 12. Install cable tie to loop BSD signal cable | 4140
916 | Volutome | 12. Test BSD functionality in Zen - Insert, retract and gain ranges all functional | 4141
916 | Volutome | 13. Check SCM with Volutome stage | 4142
916 | Volutome | 14. Noise check - Check there is no visible oscillation when imaging with a pixel size of 5nm | 4143
916 | Volutome | 15. Oscillator continuity check | 4144
916 | Volutome | 16. Perform X-ray check | 4145
916 | Volutome | 17. Pump Down Time | 4146
916 | Volutome | 17. Ultimate Vacuum | 4147
916 | Volutome | 18. Focal CC fitted to Volutome stage | 4148
916 | Volutome | 18. Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar | 4149
916 | Volutome | 18. Remove carbon needle from focal CC after testing and place in the engineers toolbox container | 4150
916 | Volutome | 19. Generate live image from Axiocam | 4151
916 | Volutome | 20. Dummy experiment completed with no erros reported in Zen - Results saved in Volutome folder | 4152
916 | Volutome | 21. Re-package Volutome and check it's completeness | 4153
916 | Volutome | 22. Check functionality of all additional software installed e.g Atlas | 4154
916 | Volutome | 23. Ensure Arivis Software Licence is available (1 year or permanent) | 4155
916 | Volutome | 24. ManEng/ Senior Eng/ Production Manger Sign off | 4156
10467 | Install SmartSEM + ASAP Engine | Install SmartSEM + ASAP Engine | 4157
10467 | Install SmartSEM + ASAP Engine | Copy the SmartSEM sotware into the SmartSEM Software folder (C:Support/Data/SmartSEM Software) | 4158
10467 | Install SmartSEM + ASAP Engine | SmartSEM Version | 4159
10467 | Install SmartSEM + ASAP Engine | SmartSEM Service Pack | 4160
10467 | Install SmartSEM + ASAP Engine | SmartSEM Hotfix | 4161
10467 | Install SmartSEM + ASAP Engine | ASAP Driver Scan Control Service Version | 4162
10467 | Install SmartSEM + ASAP Engine | Installation of Zen Core EM if Hasap is fitted | 4163
10671 | Fit Stage PCB (Klein / HiTech) | Stage PCB S/No | 4164
10671 | Fit Stage PCB (Klein / HiTech) | Fit Stage Light Barrier PCB | 4165
10671 | Fit Stage PCB (Klein / HiTech) | Set up the stage links as outlined in the procedure | 4166
10671 | Fit Stage PCB (Klein / HiTech) | All Connections Are Secured with Screws | 4167
10671 | Fit Stage PCB (Klein / HiTech) | Cables Are Neatly Routed and Secured | 4168
1392 | Stage Links | Set up stage links as outlined in the procedure. | 4169
10671 | Fit Stage PCB (Klein / HiTech) | Stage PCB S/No | 4170
10671 | Fit Stage PCB (Klein / HiTech) | Fit Stage Light Barrier PCB | 4171
10671 | Fit Stage PCB (Klein / HiTech) | Set up the stage links as outlined in the procedure | 4172
10671 | Fit Stage PCB (Klein / HiTech) | All Connections Are Secured with Screws | 4173
10671 | Fit Stage PCB (Klein / HiTech) | Cables Are Neatly Routed and Secured | 4174
1319 | Fit Stage PCB (Cartesian) | Stage PCB S/No | 4175
1319 | Fit Stage PCB (Cartesian) | Set up the stage links as outlined in the procedure | 4176
1319 | Fit Stage PCB (Cartesian) | All Connections Are Secured with Screws | 4177
1319 | Fit Stage PCB (Cartesian) | Cables Are Neatly Routed and Secured | 4178
1319 | Fit Stage PCB (Cartesian) | Stage PCB S/No | 4179
1319 | Fit Stage PCB (Cartesian) | Set up the stage links as outlined in the procedure | 4180
1319 | Fit Stage PCB (Cartesian) | All Connections Are Secured with Screws | 4181
1319 | Fit Stage PCB (Cartesian) | Cables Are Neatly Routed and Secured | 4182
1392 | EO Links | Remove links 103 + 104 from EO PCB. | 4183
10861 | Air & Water System | Air Pipes Are Correctly Routed and Kink Free | 4184
10861 | Air & Water System | Pressure Regulators Are Fitted In The Correct Orientation | 4185
10861 | Air & Water System | WaterPipes Are Correctly Routed And Free From Kinks | 4186
10849 | Power Up To Standby (Yellow Button) | Rotary Pump Pumping | 4187
10849 | Power Up To Standby (Yellow Button) | Penning Gauge LED is On | 4188
10849 | Power Up To Standby (Yellow Button) | Light Is Lit On EHT Set | 4189
10848 | Plinth Connections | Confirm there is no pressure drop over the 2 mins | 4190
10848 | Plinth Connections | Backing Pump | 4191
10848 | Plinth Connections | Cooling Water | 4192
10848 | Plinth Connections | Protective Earth | 4193
10848 | Plinth Connections | 16 Amp Mains | 4194
10851 | Run Up SmartSEM | ASAP FW/HW Present | 4195
10851 | Run Up SmartSEM | Record SmartSEM Version: | 4196
10851 | Run Up SmartSEM | ASAP FW/HW Present | 4197
10851 | Run Up SmartSEM | Record SmartSEM Version: | 4198
10851 | Run Up SmartSEM | Conirm no errors flagged in EM Server upon start up | 4199
10852 | Run UP SmartSEM Service Centre | PSU Tests Passed | 4200
10859 | System Cabling | All Cables Are Neatly Routed | 4201
10859 | System Cabling | All Plugs Are Secured Correctly | 4202
10859 | System Cabling | All Earth Cables Are Tightly Secured | 4203
10859 | System Cabling | Check Opto Devices Are Not Damaged On All PCBs & Mains Box | 4204
10859 | System Cabling | Check Stage Board Link Configuration | 4205
10798 | Clean Chamber | Clean Internal Chamber Surfaces with Carbaflow and lint free cloth. | 4206
10798 | Clean Chamber | Clean Internal Chamber Surfaces with IPA and lint free cloth. | 4207
10798 | Clean Chamber | Fit chamber meshes | 4208
10798 | Clean Chamber | Clean Internal Chamber Surfaces with Carbaflow and lint free cloth. | 4209
10798 | Clean Chamber | Clean Internal Chamber Surfaces with IPA and lint free cloth. | 4210
10798 | Clean Chamber | Fit Chamber Meshes | 4211
1232 | SmartSEM conditions | System Vacuum Reading - Confirm system vacuum reads < 1x10-5mbar | 4212
1232 | SmartSEM conditions | Gun Vacuum Reading OK - Confirm gun vacuum <3x10-6 mbar (If vacuum showing no improvment please contact ME) | 4213
1232 | SmartSEM conditions | Initialise Stage - Confirm stage initialises without errors | 4214
1232 | SmartSEM conditions | Water flow test - Confirm SmartSEM reports 'Water loss error' when water switched off at services. | 4215
1232 | SmartSEM conditions | System Vacuum Reading - Confirm system vacuum reads < 1x10-5mbar | 4216
1232 | SmartSEM conditions | Gun Vacuum Reading OK - Confirm gun vacuum <3x10-6 mbar (If vacuum showing no improvment please contact ME) | 4217
1232 | SmartSEM conditions | Set the stage type in Admin: | 4218
1232 | SmartSEM conditions | Initialise Stage - Confirm stage initialises without errors | 4219
1232 | SmartSEM conditions | Water flow test - Confirm SmartSEM reports 'Water loss error' when water switched off at services. | 4220
1232 | SmartSEM conditions | System Vacuum Reading - Confirm system vacuum reads < 1x10-5mbar | 4221
1232 | SmartSEM conditions | Gun Vacuum Reading OK - Confirm gun vacuum <3x10-6 mbar (If vacuum showing no improvment please contact ME) | 4222
1232 | SmartSEM conditions | Set Stage Type in Admin | 4223
1232 | SmartSEM conditions | Initialise Stage - Confirm stage initialises without errors | 4224
1232 | SmartSEM conditions | Water flow test - Confirm SmartSEM reports 'Water loss error' when water switched off at services. | 4225
1232 | SmartSEM conditions | System Vacuum Reading - Confirm system vacuum reads < 1x10-5mbar | 4226
1232 | SmartSEM conditions | Gun Vacuum Reading OK - Confirm gun vacuum <3x10-6 mbar (If vacuum showing no improvment please contact ME) | 4227
1232 | SmartSEM conditions | Set Stage Type in Admin | 4228
1232 | SmartSEM conditions | Initialise Stage - Confirm stage initialises without errors | 4229
1232 | SmartSEM conditions | Water flow test - Confirm SmartSEM reports 'Water loss error' when water switched off at services. | 4230
1232 | SmartSEM conditions | System Vacuum Reading - Confirm system vacuum reads < 1x10-5mbar | 4231
1232 | SmartSEM conditions | Gun Vacuum Reading OK - Confirm gun vacuum <3x10-6 mbar (If vacuum showing no improvment please contact ME) | 4232
1232 | SmartSEM conditions | Set Stage Type in Admin | 4233
1232 | SmartSEM conditions | Initialise Stage - Confirm stage initialises without errors | 4234
1232 | SmartSEM conditions | Water flow test - Confirm SmartSEM reports 'Water loss error' when water switched off at services. | 4235
1319 | Fit Stage PCB (Cartesian) | Stage PCB S/No | 4236
1319 | Fit Stage PCB (Cartesian) | Set up the stage links as outlined in the procedure | 4237
1319 | Fit Stage PCB (Cartesian) | All Connections Are Secured with Screws | 4238
1319 | Fit Stage PCB (Cartesian) | Cables Are Neatly Routed and Secured | 4239
1167 | Install Production Test Licence | Check that Time & Date on PC are set to current UK | 4240
1167 | Install Production Test Licence | Download Licence from SharePoint | 4241
1167 | Install Production Test Licence | Delete the Licence in SharePoint | 4242
1167 | Install Production Test Licence | Install the licence using the Zeiss License Management Tool. | 4243
1167 | Install Production Test Licence | hASAP Only - Installation of ZEN Core EM licence using the Zeiss activation tool | 4244
1399 | Initial Resolution Images | Actual for 1kV = | 4245
1399 | Initial Resolution Images | Store 1KV Resolution Images and Pate image | 4246
1399 | Initial Resolution Images | Actual for 15kV = | 4247
1399 | Initial Resolution Images | Store 15KV Resolution Images and Pate image | 4248
1399 | Initial Resolution Images (XB 750) | Actual for 1kV = | 4249
1399 | Initial Resolution Images (XB 750) | Store 1KV Resolution Images and Pate image | 4250
1399 | Initial Resolution Images (XB 750) | Actual for 15kV = | 4251
1399 | Initial Resolution Images (XB 750) | Store 15KV Resolution Images and Pate image | 4252
1399 | Initial Resolution Images (XB 750) | Actual for 1kV = | 4253
1399 | Initial Resolution Images (XB 750) | Store 1KV Resolution Images and Pate image | 4254
1399 | Initial Resolution Images (XB 750) | Actual for 15kV = | 4255
1399 | Initial Resolution Images (XB 750) | Store 15KV Resolution Images and Pate image | 4256
1399 | Initial Resolution Images (XB 750) | Actual for 1kV (0.5 - 2mm WD) = | 4257
1399 | Initial Resolution Images (XB 750) | Store 1KV (0.5 - 2mm WD) Resolution Images and Pate image | 4258
1399 | Initial Resolution Images (XB 750) | Actual for 15kV = | 4259
1399 | Initial Resolution Images (XB 750) | Store 15KV Resolution Images and Pate image | 4260
11255 | SCS Check | Check SCS Folder/Sales Order | 4261
11255 | SCS Check | SCS Number (If applicable) | 4262
1393 | Sense BSD | Record Diode Serial Number: | 4263
1393 | Sense BSD | Mechanical Alignment | 4264
1393 | Sense BSD | Configure BSD in Administrator | 4265
1393 | Sense BSD | Input Offset Cal (Calibration Wizard) | 4266
1393 | Sense BSD | Input Offset Cal (ASAP Test Panel) | 4267
1393 | Sense BSD | Check BIAS operation | 4268
1393 | Sense BSD | Signal To Noise Test (Result >12%) | 4269
1393 | Sense BSD | Store Signal To Noise Result (Screen-shot) | 4270
1393 | Sense BSD | Set Up Diode Protection Plate | 4271
1393 | Sense BSD | Vac Degradation < 3x10-5 mbar | 4272
1393 | Sense BSD | Fit Detector Mechanics Cover | 4273
1393 | Sense BSD | Mechanical Alignment | 4274
1393 | Sense BSD | Check BIAS operation | 4275
1393 | Sense BSD | Signal To Noise Test (Result >12%) | 4276
1393 | Sense BSD | Set Up Diode Protection Plate | 4277
1393 | Sense BSD | Vac Degradation < 3x10-5 mbar | 4278
1393 | Sense BSD | Mechanical Alignment | 4279
1393 | Sense BSD | Check BIAS operation | 4280
1393 | Sense BSD | Check Images | 4281
1393 | Sense BSD | Vac Degradation < 3x10-5 mbar | 4282
739 | Sense BSD | Record Diode Serial Number: | 4283
739 | Sense BSD | Mechanical Alignment | 4284
739 | Sense BSD | Configure BSD in Administrator | 4285
739 | Sense BSD | Input Offset Cal (Calibration Wizard) | 4286
739 | Sense BSD | Input Offset Cal (ASAP Test Panel) | 4287
739 | Sense BSD | Check BIAS operation | 4288
739 | Sense BSD | Signal To Noise Test (Result >12%) | 4289
739 | Sense BSD | Store Signal To Noise Result (Screen-shot) | 4290
739 | Sense BSD | Set Up Diode Protection Plate | 4291
739 | Sense BSD | Vac Degradation < 3x10-5 mbar | 4292
739 | Sense BSD | Fit Detector Mechanics Cover | 4293
916 | Volutome | 1. Unpack Volutome and check all parts on the packing list are present (packing list link can be found in attachments) | 4294
916 | Volutome | 1. Upload data from USB stick to the O drive product folder | 4295
916 | Volutome | 1. Remove transit clamps from Volutome stage | 4296
916 | Volutome | 1. Ensure both safety switches are fitted to the Volutome door | 4297
916 | Volutome | 2. Ensure hardware has been configured according to work instructions | 4298
916 | Volutome | 2. Fit Volutome stage | 4299
916 | Volutome | 2. Check protection circuit for Piezo is fitted (349567-9013-000) | 4300
916 | Volutome | 2. Connect Volutome stage and controller to the SEM | 4301
916 | Volutome | 2. Fit Stemi holder bracket to Volutome stage door (STEMI adjustment bracket needs to fit / disconnect without force from door) | 4302
916 | Volutome | 3. Ethernet Connection Setup | 4303
916 | Volutome | 4. Install the latest version of Zen - Ensure to configure it for Volutome (Refer to Zen installation WI) | 4304
916 | Volutome | 4. Ensure Athena inspector manager has all green lights, Physik microtome and Physik stage are present - save a screenshot in the Volutome folder | 4305
916 | Volutome | 4. Install PI Micromove | 4306
916 | Volutome | 4. Run getversions.exe and save text file in the Volutome folder | 4307
916 | Volutome | 4. Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator | 4308
916 | Volutome | 4. Set up the Volutome workbench in Zen | 4309
916 | Volutome | 5. PID parameters configured | 4310
916 | Volutome | 6. Record VolumeBSD serial number | 4311
916 | Volutome | 6. Record Volutome stage serial number | 4312
916 | Volutome | 6. Record VolumeBSD amplifier serial number | 4313
916 | Volutome | 6. Record VolumeBSD mechanics serial number | 4314
916 | Volutome | 6. Record C-885 PI controller serial number | 4315
916 | Volutome | 6. Record E-727 PI controller serial number | 4316
916 | Volutome | 7. Volutome 'internal LED light' function check | 4317
916 | Volutome | 8. Volutome 'External cold light' function check | 4318
916 | Volutome | 9. Adjust Volutome stage centre (Crosshairs should be within 4 central squares) | 4319
916 | Volutome | 10. Install Nav cam in-line connections (If applicable) | 4320
916 | Volutome | 11. Perform the microtome and stage functionality checks | 4321
916 | Volutome | 12. Fit and align the VolumeBSD diode | 4322
916 | Volutome | 12. Check VolumeBSD alignment repeatability | 4323
916 | Volutome | 12. Check all gain ranges of the VolumeBSD (normal/inverted) | 4324
916 | Volutome | 12. Perform Signal/noise ratio check | 4325
916 | Volutome | 12. Ensure BSD diode protector is fitted and adjusted in chamber | 4326
916 | Volutome | 12. Install cable tie to loop BSD signal cable | 4327
916 | Volutome | 12. Test BSD functionality in Zen - Insert, retract and gain ranges all functional | 4328
916 | Volutome | 13. Check SCM with Volutome stage | 4329
916 | Volutome | 14. Noise check - Check there is no visible oscillation when imaging with a pixel size of 5nm | 4330
916 | Volutome | 15. Oscillator continuity check | 4331
916 | Volutome | 16. Perform X-ray check | 4332
916 | Volutome | 17. Pump Down Time | 4333
916 | Volutome | 17. Ultimate Vacuum | 4334
916 | Volutome | 18. Focal CC fitted to Volutome stage | 4335
916 | Volutome | 18. Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar | 4336
916 | Volutome | 18. Remove carbon needle from focal CC after testing and place in the engineers toolbox container | 4337
916 | Volutome | 19. Generate live image from Axiocam | 4338
916 | Volutome | 20. Dummy experiment completed with no erros reported in Zen - Results saved in Volutome folder | 4339
916 | Volutome | 21. Re-package Volutome and check it's completeness | 4340
916 | Volutome | 22. Check functionality of all additional software installed e.g Atlas | 4341
916 | Volutome | 23. Ensure Arivis Software Licence is available (1 year or permanent) | 4342
10670 | Fit EO PCB (Standard) | Confirm Standard EO PCB (356133-9004-000) | 4343
10670 | Fit EO PCB (Standard) | EO PCB S/No | 4344
10670 | Fit EO PCB (Standard) | Fit EO PCB | 4345
10670 | Fit EO PCB (Standard) | All Connections Are Secured with Screws | 4346
10670 | Fit EO PCB (Standard) | check firmware on proms is v2 and ends with 020 | 4347
10670 | Fit EO PCB (Standard) | Cables Are Neatly Routed and Secured | 4348
1399 | Initial Resolution Images (XB 750) | Actual for 1kV (0.5 - 2mm WD) = | 4349
1399 | Initial Resolution Images (XB 750) | Store 1KV (0.5 - 2mm WD) Resolution Images and Pate image | 4350
1399 | Initial Resolution Images (XB 750) | Actual for 1kV (4mm WD) = | 4351
1399 | Initial Resolution Images (XB 750) | Store 1KV (4mm WD) Resolution Images and Pate image | 4352
1399 | Initial Resolution Images (XB 750) | Actual for 15kV = | 4353
1399 | Initial Resolution Images (XB 750) | Store 15KV Resolution Images and Pate image | 4354
1399 | Initial Resolution Images (XB 750) | Actual for 1kV (0.5 - 2mm WD) = | 4355
1399 | Initial Resolution Images (XB 750) | Store 1KV (0.5 - 2mm WD) Resolution Images and Pate image | 4356
1399 | Initial Resolution Images (XB 750) | Actual for 1kV (4mm WD) = | 4357
1399 | Initial Resolution Images (XB 750) | Store 1KV (4mm WD) Resolution Images and Pate image | 4358
1399 | Initial Resolution Images (XB 750) | Actual for 15kV = | 4359
1399 | Initial Resolution Images (XB 750) | Store 15KV Resolution Images and Pate image | 4360
425 | PC MAC Address | Record MAC address (Internet not MIC-NET) | 4361
10930 | FIB set up and Disabled | Run the ‘Disable FIB–all GIS.reg’ file | 4362
252 | Configure PC ASAP | PC S/N | 4363
252 | Configure PC ASAP | PC/Model (SAP No.) | 4364
252 | Configure PC ASAP | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448 | 4365
252 | Configure PC ASAP | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448 , then the PC needs reimaging – follow the attached WI. The correct image version is V2. | 4366
252 | Configure PC ASAP | Hard disk Image (Version) | 4367
252 | Configure PC ASAP | Record Operating System | 4368
252 | Configure PC ASAP | Check the Entitlement certificate to determine which version of Smartsem to install | 4369
252 | Configure PC ASAP | Install Smart SEM (See WI 9693) Record Version | 4370
252 | Configure PC ASAP | Check for available SmartSEMService Packs and install them. (N/A if none) | 4371
252 | Configure PC ASAP | Check for available SmartSEM hotfixes and install them. (N/A if none) | 4372
252 | Configure PC ASAP | Install CZ Support Tools. Ver | 4373
252 | Configure PC ASAP | Disable Automatic Updates | 4374
252 | Configure PC ASAP | Run ASAP Screening Program | 4375
252 | Configure PC ASAP | System Specific Datazone Loaded | 4376
252 | Configure PC ASAP | In SEM status set 'Mains Sync' to off | 4377
254 | Install Customer Language Option | Language | 4378
972 | Install Service centre Software | Service centre S/W installed on Customer PC. | 4379
972 | Install Service centre Software | Confirm Data Logging is running and External Logging is enabled. | 4380
11270 | Install Predictive Service Software | Predictive Service S/W installed on Customer PC. | 4381
11270 | Install Predictive Service Software | Predictive Service has been activated | 4382
1399 | Initial Resolution Images (XB 750) | Actual for 1kV (4mm WD) = | 4383
1399 | Initial Resolution Images (XB 750) | Store 1KV (4mm WD) Resolution Images and Pate image | 4384
1399 | Initial Resolution Images (XB 750) | Actual for 15kV = | 4385
1399 | Initial Resolution Images (XB 750) | Store 15KV Resolution Images and Pate image | 4386
303 | Stage Drift Program | Actual (Spec: 40nm /6 min) | 4387
303 | Stage Drift Program | Actual (Spec: 40nm /6 min) | 4388
303 | Stage Drift Program | Store stage drift macro results image | 4389
6931 | Stage Drift Program | Drift | 4390
6931 | Stage Drift Program | Drift result 1 | 4391
6931 | Stage Drift Program | Drift result 2 | 4392
427 | Column Alignment | Confirm that the contents of Tablefiller are deleted before starting the procedure. Stored values can affect the alignment routine. | 4393
427 | Column Alignment | Mechanically align the column using the Merlin Alignment Wizard | 4394
427 | Column Alignment | Confirm that the Align 1 figures obtained during Mechanical Alignment are <15%. If they are not, consult a Supervisor or Man Eng before continuing. | 4395
427 | Column Alignment | Tighten Gun Head clamping screws and confirm the 3 out of 4 accessible are a minimum torque of 1Nm with torque driver. | 4396
427 | Column Alignment | Once adjusters are backed off, tighten C1 + C2 clamping screws and confirm they are a minimum torque of 1Nm with torque driver. | 4397
184 | Gun Conditioning Test when at final vacuum | Start Gun Vacuum Tracing | 4398
184 | Gun Conditioning Test when at final vacuum | Save Gun Vacuum Trace | 4399
184 | Gun Conditioning Test when at final vacuum | Store Gun Conditioning Trace as .PDF | 4400
184 | Gun Conditioning Test when at final vacuum | Start Gun Vacuum Tracing | 4401
184 | Gun Conditioning Test when at final vacuum | Save Gun Vacuum Trace | 4402
184 | Gun Conditioning Test when at final vacuum | Store Gun Conditioning Trace as .PDF | 4403
184 | Gun Conditioning Test when at final vacuum | Start Gun Vacuum Tracing | 4404
184 | Gun Conditioning Test when at final vacuum | Store Gun Conditioning Trace as .PDF | 4405
10487 | Condition Gun - Tungsten | Gun Monitor Setup | 4406
10487 | Condition Gun - Tungsten | Test No. 96 Selected | 4407
10487 | Condition Gun - Tungsten | EHT Stable at 32KV - ≥15 minutes | 4408
10487 | Condition Gun - Tungsten | 32KV Gun Monitor Traced Saved | 4409
10487 | Condition Gun - Tungsten | Gun Monitor Setup | 4410
10487 | Condition Gun - Tungsten | Test No. 96 Selected | 4411
10487 | Condition Gun - Tungsten | EHT Stable at 32KV - ≥15 minutes | 4412
10487 | Condition Gun - Tungsten | 32KV Gun Monitor Traced Saved | 4413
10488 | Obtain Image | Ensure Gun Monitor is Running and Test 96 Selected. | 4414
10488 | Obtain Image | Carry Out Initial Gun Run Up | 4415
10488 | Obtain Image | Carry Out Gun Run Up | 4416
10488 | Obtain Image | X Gun Tilt ≤50% | 4417
10488 | Obtain Image | Y Gun Tilt ≤50% | 4418
10488 | Obtain Image | X Gun Shift ≤50% | 4419
10488 | Obtain Image | Y Gun Shift ≤50% | 4420
10498 | Condition Gun - LAB6 | Connect The Heinzinger PSU | 4421
10498 | Condition Gun - LAB6 | Check that all the connections are correct | 4422
10498 | Condition Gun - LAB6 | Run up to 35KV and leave for 15 minute | 4423
10498 | Condition Gun - LAB6 | Connect The Heinzinger PSU | 4424
10498 | Condition Gun - LAB6 | Check that all the connections are correct | 4425
10498 | Condition Gun - LAB6 | Run up to 35KV and leave for 15 minute | 4426
10498 | Condition Gun - LAB6 | No Strom Current is Displayed | 4427
10498 | Condition Gun - LAB6 | Switch Off the Heinzinger and disconnect | 4428
10498 | Condition Gun - LAB6 | Reconnect the Gun Cable to the Spellman EHT Set | 4429
10498 | Condition Gun - LAB6 | Connect The Heinzinger PSU | 4430
10498 | Condition Gun - LAB6 | Check that all the connections are correct | 4431
10498 | Condition Gun - LAB6 | Run up to 35KV and leave for 15 minute | 4432
10498 | Condition Gun - LAB6 | No Strom Current is Displayed | 4433
10498 | Condition Gun - LAB6 | Switch Off the Heinzinger and disconnect | 4434
10498 | Condition Gun - LAB6 | Reconnect the Gun Cable to the Spellman EHT Set | 4435
10498 | Condition Gun - LAB6 | Connect The Heinzinger PSU | 4436
10498 | Condition Gun - LAB6 | Check that all the connections are correct | 4437
10498 | Condition Gun - LAB6 | Run up to 35KV and leave for 30 minute | 4438
10498 | Condition Gun - LAB6 | No Strom Current is Displayed | 4439
10498 | Condition Gun - LAB6 | Check No Spikes in Gun Vacuum Trace | 4440
10498 | Condition Gun - LAB6 | Switch Off the Heinzinger and disconnect | 4441
10498 | Condition Gun - LAB6 | Reconnect the Gun Cable to the Spellman EHT Set | 4442
10498 | Condition Gun - LAB6 | Connect The Heinzinger PSU | 4443
10498 | Condition Gun - LAB6 | Check that all the connections are correct | 4444
10498 | Condition Gun - LAB6 | Run up to 35KV and leave for 30 minute | 4445
10498 | Condition Gun - LAB6 | No Strom Current is Displayed | 4446
10498 | Condition Gun - LAB6 | Check No Spikes in Gun Vacuum Trace | 4447
10498 | Condition Gun - LAB6 | Switch Off the Heinzinger and disconnect | 4448
10498 | Condition Gun - LAB6 | Remove interlock cable and plug the gun interlock cable back into the Vice PCB. | 4449
10498 | Condition Gun - LAB6 | Reconnect the Gun Cable to the Spellman EHT Set | 4450
1403 | Install Quiet Mode | Install customer quiet mode | 4451
1403 | Install Quiet Mode | Set Test tool Quiet Mode to 'OFF' (In service cupboard) | 4452
10492 | Column Alignment | LaB6 Only - Loosen Ion Pump Support Bracket | 4453
10492 | Column Alignment | C3 Reversal-≤20µm | 4454
10492 | Column Alignment | 30KV-5KV-≤20µm | 4455
10492 | Column Alignment | Image Shift-≤3000nm | 4456
10492 | Column Alignment | LaB6 Only - Loosen Ion Pump Support Bracket | 4457
10492 | Column Alignment | Align Column To Within Spec. | 4458
10492 | Column Alignment | LaB6 Only - Tighten the Ion Pump Support Bracket. | 4459
10492 | Column Alignment | C3 Reversal-≤20µm | 4460
10492 | Column Alignment | 30KV-5KV-≤20µm | 4461
10492 | Column Alignment | Image Shift-≤3000nm | 4462
10492 | Column Alignment | LaB6 Only - Loosen the Ion Pump Support Bracket | 4463
10492 | Column Alignment | Align Column To Within Spec. | 4464
10492 | Column Alignment | LaB6 Only - Tighten the Ion Pump Support Bracket. | 4465
10492 | Column Alignment | C3 Reversal-≤20µm | 4466
10492 | Column Alignment | 30KV-5KV-≤20µm | 4467
10492 | Column Alignment | Image Shift-≤3000nm | 4468
10492 | Column Alignment | LaB6 Only - Loosen the Ion Pump Support Bracket | 4469
10492 | Column Alignment | Align Column To Within Spec. | 4470
10492 | Column Alignment | LaB6 Only - Tighten the Ion Pump Support Bracket. | 4471
10492 | Column Alignment | Check C3 Reversal-≤20µm | 4472
10492 | Column Alignment | Check Image Movement - 30KV-5KV-≤20µm | 4473
10492 | Column Alignment | Check Image Shift-≤3000nm | 4474
10492 | Column Alignment | LaB6 Only - Loosen the Ion Pump Support Bracket | 4475
10492 | Column Alignment | Align Column To Within Spec. | 4476
10492 | Column Alignment | LaB6 Only - Tighten the Ion Pump Support Bracket. | 4477
10492 | Column Alignment | Tighten the Column Clamping Screws | 4478
10492 | Column Alignment | Check C3 Reversal-≤20µm | 4479
10492 | Column Alignment | Check Image Movement - 30KV-5KV-≤20µm | 4480
10492 | Column Alignment | Check Image Shift-≤3000nm | 4481
252 | Configure PC ASAP | PC S/N | 4482
252 | Configure PC ASAP | PC/Model (SAP No.) | 4483
252 | Configure PC ASAP | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448 | 4484
252 | Configure PC ASAP | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448 , then the PC needs reimaging – follow the attached WI. The correct image version is V2. | 4485
252 | Configure PC ASAP | Hard disk Image (Version) | 4486
252 | Configure PC ASAP | Record Operating System | 4487
252 | Configure PC ASAP | Check the Entitlement certificate to determine which version of Smartsem to install | 4488
252 | Configure PC ASAP | Install Smart SEM (See WI 9693) Record Version | 4489
252 | Configure PC ASAP | Check for available SmartSEMService Packs and install them. (N/A if none) | 4490
252 | Configure PC ASAP | Check for available SmartSEM hotfixes and install them. (N/A if none) | 4491
252 | Configure PC ASAP | Installation of Zen Core EM | 4492
252 | Configure PC ASAP | Install CZ Support Tools. Ver | 4493
252 | Configure PC ASAP | Disable Automatic Updates | 4494
252 | Configure PC ASAP | Run ASAP Screening Program | 4495
252 | Configure PC ASAP | System Specific Datazone Loaded | 4496
252 | Configure PC ASAP | In SEM status set 'Mains Sync' to off | 4497
11274 | Fit Column Cables + Pipes | Fit and route all column cables | 4498
11274 | Fit Column Cables + Pipes | Confirm IGP cable locking screws are secured and grub screw tightened | 4499
11274 | Fit Column Cables + Pipes | Fit and route objective water pipes | 4500
11274 | Fit Column Cables + Pipes | Safety – Torque the earth stud to 15 N.m | 4501
11274 | Fit Column Cables + Pipes | Torque both earth zinc plated M8 nuts to 15 N.m | 4502
11274 | Fit Column Cables + Pipes | Confirm HV cable is routed though plinth, and secured on top of plinth | 4503
3246 | Connect Plinth Services | Air Pressure O/P = 6bar | 4504
3246 | Connect Plinth Services | Water | 4505
3246 | Connect Plinth Services | Mains | 4506
3246 | Connect Plinth Services | Ensure Test tool Quiet Mode is set to 'ON' in service cupboard (Unless customer Quiet mode fitted) | 4507
1278 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 4508
1278 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 4509
1278 | ZEN Core EM | Install Customer's ZEN Licence | 4510
1278 | ZEN Core EM | Setup SmartSEM Toggle Button | 4511
1278 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 4512
1278 | ZEN Core EM | Mains Sync Set To OFF | 4513
1278 | ZEN Core EM | Select Your ZEN core EM Runtime Profile | 4514
1278 | ZEN Core EM | Create User Supervisor EM/LM - Only required if ZEN EM and LM are installed | 4515
1278 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 4516
1278 | ZEN Core EM | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems. | 4517
1278 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 4518
1278 | ZEN Core EM | Configue Airlock Cam and Setup Exchange Position $Exchange (use Upper case E in ZEN) | 4519
1278 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 4520
1278 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 4521
1278 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 4522
1278 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 4523
1278 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 4524
1278 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 4525
1279 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 4526
1279 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 4527
1279 | ZEN Core EM | Install Customer's ZEN Licence | 4528
1279 | ZEN Core EM | Setup SmartSEM Toggle Button | 4529
1279 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 4530
1279 | ZEN Core EM | Mains Sync Set To OFF | 4531
1279 | ZEN Core EM | Choose your ZEN core EM runtime profile | 4532
1279 | ZEN Core EM | Create User Supervisor EM/LM - Only required if ZEN EM and LM are installed | 4533
1279 | ZEN Core EM | Create User Supervisor Volutome (only applicable for Volutome systems) | 4534
1279 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 4535
1279 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 4536
1279 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 4537
1279 | ZEN Core EM | Configue $NavCam (Capital N and C) | 4538
1279 | ZEN Core EM | Configue $Exchange (capital E) | 4539
1279 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 4540
1279 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 4541
1279 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 4542
1279 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 4543
1279 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 4544
1279 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 4545
1280 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 4546
1280 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 4547
1280 | ZEN Core EM | Install Customer's ZEN Licence | 4548
1280 | ZEN Core EM | Setup SmartSEM Toggle Button | 4549
1280 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 4550
1280 | ZEN Core EM | Mains Sync Set To OFF | 4551
1280 | ZEN Core EM | Choose your ZEN core EM runtime profile | 4552
1280 | ZEN Core EM | Create User Supervisor EM/LM - Only required if ZEN EM and LM are installed | 4553
1280 | ZEN Core EM | Create User Supervisor Volutome (only applicable for Volutome systems) | 4554
1280 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 4555
1280 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 4556
1280 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 4557
1280 | ZEN Core EM | Configue $NavCam (Capital N and C) | 4558
1280 | ZEN Core EM | Configue $Exchange (capital E) | 4559
1280 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 4560
1280 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 4561
1280 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 4562
1280 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 4563
1280 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 4564
1280 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 4565
916 | Volutome | 1. Unpack Volutome and check all parts on the packing list are present (packing list link can be found in attachments) | 4566
916 | Volutome | 1. Upload data from USB stick to the O drive product folder | 4567
916 | Volutome | 1. Remove transit clamps from Volutome stage | 4568
916 | Volutome | 1. Ensure both safety switches are fitted to the Volutome door | 4569
916 | Volutome | 2. Ensure hardware has been configured according to work instructions | 4570
916 | Volutome | 2. Fit Volutome stage | 4571
916 | Volutome | 2. Check protection circuit for Piezo is fitted (349567-9013-000) | 4572
916 | Volutome | 2. Connect Volutome stage and controller to the SEM | 4573
916 | Volutome | 2. Fit Stemi holder bracket to Volutome stage door (STEMI adjustment bracket needs to fit / disconnect without force from door) | 4574
916 | Volutome | 3. Ethernet Connection Setup | 4575
916 | Volutome | 4. Install the latest version of Zen - Ensure to configure it for Volutome (Refer to Zen installation WI) | 4576
916 | Volutome | 4. Ensure Athena inspector manager has all green lights, Physik microtome and Physik stage are present - save a screenshot in the Volutome folder | 4577
916 | Volutome | 4. Install PI Micromove | 4578
916 | Volutome | 4. Run getversions.exe and save text file in the Volutome folder | 4579
916 | Volutome | 4. Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator | 4580
916 | Volutome | 4. Set up the Volutome workbench in Zen | 4581
916 | Volutome | 5. PID parameters configured | 4582
916 | Volutome | 6. Record VolumeBSD serial number | 4583
916 | Volutome | 6. Record Volutome stage serial number | 4584
916 | Volutome | 6. Record VolumeBSD amplifier serial number | 4585
916 | Volutome | 6. Record VolumeBSD mechanics serial number | 4586
916 | Volutome | 6. Record C-885 PI controller serial number | 4587
916 | Volutome | 6. Record E-727 PI controller serial number | 4588
916 | Volutome | 7. Volutome 'internal LED light' function check | 4589
916 | Volutome | 8. Volutome 'External cold light' function check | 4590
916 | Volutome | 9. Adjust Volutome stage centre (Crosshairs should be within 4 central squares) | 4591
916 | Volutome | 10. Install Nav cam in-line connections (If applicable) | 4592
916 | Volutome | 11. Perform the microtome and stage functionality checks | 4593
916 | Volutome | 12. Fit and align the VolumeBSD diode | 4594
916 | Volutome | 12. Check VolumeBSD alignment repeatability | 4595
916 | Volutome | 12. Check all gain ranges of the VolumeBSD (normal/inverted) | 4596
916 | Volutome | 12. Perform Signal/noise ratio check | 4597
916 | Volutome | 12. Ensure BSD diode protector is fitted and adjusted in chamber | 4598
916 | Volutome | 12. Install cable tie to loop BSD signal cable | 4599
916 | Volutome | 12. Test BSD functionality in Zen - Insert, retract and gain ranges all functional | 4600
916 | Volutome | 13. Check SCM with Volutome stage | 4601
916 | Volutome | 14. Noise check - Check there is no visible oscillation when imaging with a pixel size of 5nm | 4602
916 | Volutome | 15. Oscillator continuity check | 4603
916 | Volutome | 16. Perform X-ray check | 4604
916 | Volutome | 17. Pump Down Time | 4605
916 | Volutome | 17. Ultimate Vacuum | 4606
916 | Volutome | 18. Focal CC fitted to Volutome stage | 4607
916 | Volutome | 18. Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar | 4608
916 | Volutome | 18. Remove carbon needle from focal CC after testing and place in the engineers toolbox container | 4609
916 | Volutome | 19. Generate live image from Axiocam | 4610
916 | Volutome | 20. Dummy experiment completed with no erros reported in Zen - Results saved in Volutome folder | 4611
916 | Volutome | 21. Re-package Volutome and check it's completeness | 4612
916 | Volutome | 22. Check functionality of all additional software installed e.g Atlas | 4613
916 | Volutome | 22. Check four standard Volutome ZEN licenses are active (ZEN core, GeminiSEM, Volutome, Automated Imaging) | 4614
916 | Volutome | 23. Ensure Arivis Software Licence is available (1 year or permanent) | 4615
371 | Generate the following Documents (Gemini) | X-Ray Certificate | 4616
371 | Generate the following Documents (Gemini) | Quality Certificate | 4617
371 | Generate the following Documents (Gemini) | Customer Entitlement Certificate | 4618
371 | Generate the following Documents (Gemini) | DENKA Data Sheet | 4619
371 | Generate the following Documents (Gemini) | Options Removed packing List | 4620
371 | Generate the following Documents (Gemini) | Options Retained packing List | 4621
371 | Generate the following Documents (Gemini) | Standard Packing List | 4622
371 | Generate the following Documents (Gemini) | PC Inferface assignment | 4623
371 | Generate the following Documents (Gemini) | Declaration of conformity | 4624
371 | Generate the following Documents (Gemini) | Acceptance test record | 4625
371 | Generate the following Documents (Gemini) | Licence Generator report | 4626
427 | Column Alignment | Confirm that the contents of Tablefiller are deleted before starting the procedure. Stored values can affect the alignment routine. | 4627
427 | Column Alignment | Mechanically align the column using the Merlin Alignment Wizard. Note:- Gun head and C1/C2 clamping screws should be tightened to a minimum torque of 1nM using a torque driver. | 4628
427 | Column Alignment | Confirm that the Align 1 figures obtained during Mechanical Alignment are <15%. If they are not, consult a Supervisor or Man Eng before continuing. | 4629
427 | Column Alignment | Once the column is mechanically aligned, go to 10kv Analytic Column mode and set maximum Probe current. In the Column calibration window adjust Gun Alignment to maximise Probe current into a FC. Gun Align values should be <60%. If they are >60% the column will need mechanically aligning again. | 4630
303 | Stage Drift Program | Actual (Spec: 40nm /6 min) | 4631
303 | Stage Drift Program | Actual (Spec: 40nm /6 min) | 4632
303 | Stage Drift Program | Store stage drift macro results image | 4633
10728 | Fit Column Cabling | Scan Cables | 4634
10728 | Fit Column Cabling | Lens Cables | 4635
10728 | Fit Column Cabling | EHT Cables | 4636
10728 | Fit Column Cabling | Penning Cable | 4637
10730 | Fit LaB6 Cabling | Valves Cables | 4638
10730 | Fit LaB6 Cabling | Heater Cables | 4639
10730 | Fit LaB6 Cabling | Ion Pump Cable | 4640
10730 | Fit LaB6 Cabling | Ion Pump PSU Serial Number | 4641
10730 | Fit LaB6 Cabling | Fit Ion Pump PSU Bracket | 4642
10730 | Fit LaB6 Cabling | Fit Ion Pump PSU | 4643
10730 | Fit LaB6 Cabling | PSU Qualification Marks Checked | 4644
10731 | Fit EHT Set | EHT Set Type Fitted | 4645
10731 | Fit EHT Set | EHT Serial Number | 4646
10731 | Fit EHT Set | Revision/Mod Level | 4647
10731 | Fit EHT Set | Qualification Marks Checked | 4648
10731 | Fit EHT Set | Torque Earth Connection to 3Nm | 4649
10710 | Fit Chamber Assembly | Check Stage Rail Grub Screws are Tight | 4650
10710 | Fit Chamber Assembly | Torque Chamber Bolts to 35Nm | 4651
10710 | Fit Chamber Assembly | Record Chamber SN | 4652
10710 | Fit Chamber Assembly | Check Bellows is concentric to the hole in the desktop. | 4653
1134 | Print Shipment, Packing Labels and Identify System Material | Print Shipment Labels | 4654
1134 | Print Shipment, Packing Labels and Identify System Material | Sort out customer material, label if possible | 4655
1134 | Print Shipment, Packing Labels and Identify System Material | Place left over stickes in Document Holder | 4656
1134 | Print Shipment, Packing Labels and Identify System Material | Print Shipment Labels | 4657
1134 | Print Shipment, Packing Labels and Identify System Material | Sort out customer material in box, label if possible | 4658
1134 | Print Shipment, Packing Labels and Identify System Material | Place left over stickes in Document Holder | 4659
10610 | Print Shipping Documents | Print a copy of sales order (See order inbox link) | 4660
10465 | Install and SetUp Windows 10 | Windows Version | 4661
10465 | Install and SetUp Windows 10 | Disable Drivers and Windows Updates | 4662
10467 | Install SmartSEM + ASAP Engine | Install SmartSEM + ASAP Engine | 4663
10467 | Install SmartSEM + ASAP Engine | Copy the SmartSEM sotware into the SmartSEM Software folder (C:Support/Data/SmartSEM Software) | 4664
10467 | Install SmartSEM + ASAP Engine | SmartSEM Version | 4665
10467 | Install SmartSEM + ASAP Engine | SmartSEM Service Pack | 4666
10467 | Install SmartSEM + ASAP Engine | ASAP Driver Scan Control Service Version | 4667
10467 | Install SmartSEM + ASAP Engine | Installation of Zen Core EM | 4668
1167 | Install Production Test Licence | Check that Time & Date on PC are set to current UK | 4669
1167 | Install Production Test Licence | Download Licence from SharePoint | 4670
1167 | Install Production Test Licence | Delete the Licence in SharePoint | 4671
1167 | Install Production Test Licence | Install the licence using the Zeiss License Management Tool. | 4672
1167 | Install Production Test Licence | Installation of ZEN Core EM licence using the Zeiss activation tool | 4673
8699 | Cable Up | System Mains Voltage | 4674
8699 | Cable Up | IF LaB6 and 120V - Transformer Installed | 4675
8699 | Cable Up | Cable Up | 4676
8699 | Cable Up | Label Cables | 4677
929 | Print and Fit S/N Label | Print and fit the relevant S/N label to the system. See WI for details. | 4678
11260 | Install C2D Detector | Select Detector Type | 4679
11194 | Fit SE Detector | Check that the Assembly is the Correct Part Number | 4680
11194 | Fit SE Detector | Serial Number (or NA) | 4681
11194 | Fit SE Detector | Correct Collector Beak Fitted | 4682
11194 | Fit SE Detector | Fit Detector Assembly | 4683
11194 | Fit SE Detector | Apply Emersion Oil to PM Tube Face | 4684
11194 | Fit SE Detector | Fit PM Head Amp Assembly confirm thst wires are still connected after fitment | 4685
1167 | Install Production Test Licence | Check that Time & Date on PC are set to current UK | 4686
1167 | Install Production Test Licence | Download SmartSEM Licence from SharePoint | 4687
1167 | Install Production Test Licence | Delete the SmartSEM Licence in SharePoint | 4688
1167 | Install Production Test Licence | Download ZEN Licence from SharePoint | 4689
1167 | Install Production Test Licence | Delete the ZEN Licence in SharePoint | 4690
1167 | Install Production Test Licence | Install the SmartSEM licence using the Zeiss License Management Tool. | 4691
1167 | Install Production Test Licence | Installation of ZEN Core EM licence using the Zeiss Licence Management Tool | 4692
1134 | Print Shipment, Packing Labels and Identify System Material | Print Shipment Labels | 4693
1134 | Print Shipment, Packing Labels and Identify System Material | Sort out customer material in box, label if possible | 4694
1134 | Print Shipment, Packing Labels and Identify System Material | Place left over stickes in Document Holder | 4695
1405 | Cable Up | System Mains Voltage | 4696
1405 | Cable Up | IF LaB6 and 120V - Transformer Installed | 4697
1405 | Cable Up | Cable Up | 4698
1405 | Cable Up | Label Cables | 4699
8699 | Cable Up (240V Mains) | System Mains Voltage | 4700
8699 | Cable Up (240V Mains) | IF LaB6 and 120V - Transformer Installed | 4701
8699 | Cable Up (240V Mains) | Cable Up | 4702
8699 | Cable Up (240V Mains) | Label Cables | 4703
1405 | Cable Up (120V Mains) | System Mains Voltage | 4704
1405 | Cable Up (120V Mains) | IF LaB6 and 120V - Transformer Installed | 4705
1405 | Cable Up (120V Mains) | Cable Up | 4706
1405 | Cable Up (120V Mains) | Label Cables | 4707
1655 | UHV Chamber S/N | S/N = | 4708
1655 | UHV Chamber S/N | Please confirm UHV supplier | 4709
1655 | UHV Chamber S/N |  | 4710
10652 | SmartEDX | Type of Detector | 4711
10652 | SmartEDX | Detector Serial number | 4712
10652 | SmartEDX | DDP Module Serial Number | 4713
10652 | SmartEDX | Apply Krytox Grease to the Adaptor Plate Inner O Ring. | 4714
10652 | SmartEDX | Fit Flange & Mount Detector | 4715
10652 | SmartEDX | Install EDS PCB's Into PC | 4716
10652 | SmartEDX | Installation of DPP Module & Cables | 4717
10652 | SmartEDX | Configure PC Network & Test | 4718
10652 | SmartEDX | Disable Windows10 Fast Startup | 4719
10652 | SmartEDX | Install ApexKey.edx Licences | 4720
10652 | SmartEDX | Install Genesis.edx Licences (if Applicable) | 4721
10652 | SmartEDX | Fit Dongle | 4722
10652 | SmartEDX | APEX Software Version | 4723
10652 | SmartEDX | Install APEX Software | 4724
10652 | SmartEDX | Install Remoting | 4725
10652 | SmartEDX | Setup RConfigure | 4726
10652 | SmartEDX | Run API Test | 4727
10652 | SmartEDX | Configure Imaging-Run SGShell | 4728
10652 | SmartEDX | Calibrate Detector-Run eDPP Shell | 4729
10652 | SmartEDX | Save Calibration Report | 4730
10652 | SmartEDX | Detector Resolution (measured using Cu) ≈ 129 eV | 4731
10652 | SmartEDX | Store Calibration Report | 4732
10652 | SmartEDX | Save Calibration on Exit | 4733
10652 | SmartEDX | Run APEX | 4734
10652 | SmartEDX | Check Spectrum | 4735
10652 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected | 4736
10652 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected | 4737
10652 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected | 4738
10652 | SmartEDX | Check Mapping | 4739
10652 | SmartEDX | Check Line Scanning | 4740
10652 | SmartEDX | Save Aquired Spectrum, Mappimg & Line Scanning to the Image Folder | 4741
10652 | SmartEDX | Store Aquired Spectrum (STD, Quick, HD), Mappimg & Line Scanning Images | 4742
10652 | SmartEDX | Save Calibration Data to the folder C: Support/Data/ Log Files | 4743
10652 | SmartEDX | Store Calibration Data File | 4744
10652 | SmartEDX | Save Calibration Data to the SmartSEM Image Folder | 4745
449 | BSD-1 | Configure BSD in Administrator | 4746
449 | BSD-1 | Mechanical Alignment | 4747
449 | BSD-1 | Input Offset Cal | 4748
449 | BSD-1 | Black Level Cal | 4749
449 | BSD-1 | Diode Segment Test | 4750
449 | BSD-1 | Diode Mirror Test | 4751
449 | BSD-1 | Vac Degradation < 3x10-5 mbar | 4752
449 | BSD-1 | All Gain Ranges working | 4753
449 | BSD-1 | Check all channels in Invert mode | 4754
1278 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 4755
1278 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 4756
1278 | ZEN Core EM | Install Customer's ZEN Licence | 4757
1278 | ZEN Core EM | Setup SmartSEM Toggle Button | 4758
1278 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 4759
1278 | ZEN Core EM | Mains Sync Set To OFF | 4760
1278 | ZEN Core EM | Select Your ZEN core EM Runtime Profile | 4761
1278 | ZEN Core EM | Create User Supervisor EM/LM - Only required if ZEN EM and LM are installed (Note: This is rare) | 4762
1278 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 4763
1278 | ZEN Core EM | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems. | 4764
1278 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 4765
1278 | ZEN Core EM | Configue Airlock Cam and Setup Exchange Position $Exchange (use Upper case E in ZEN) | 4766
1278 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 4767
1278 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 4768
1278 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 4769
1278 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 4770
1278 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 4771
1278 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 4772
10526 | Chamberscope | Number of Chamberscope's on Order | 4773
10526 | Chamberscope | Chamberscope 1 - Select Port | 4774
10526 | Chamberscope | Check Camera for focus | 4775
10526 | Chamberscope | Functional Check-Black/White + Colour Mode | 4776
10526 | Chamberscope | Chamberscope 2 - Select Port or Not Fitted | 4777
10526 | Chamberscope | Chamberscope 2 - Check Camera 2 for Focus | 4778
10526 | Chamberscope | Chamberscope 2 - Functional Check-Black/White + Colour Mode | 4779
10526 | Chamberscope | Chamberscope 2 - Add Icon to Toolbar | 4780
3314 | SmartEDX | Type of Detector | 4781
3314 | SmartEDX | Detector Serial number | 4782
3314 | SmartEDX | DDP Module Serial Number | 4783
3314 | SmartEDX | Apply Krytox Grease to the Adaptor Plate Inner O Ring. | 4784
3314 | SmartEDX | Fit Flange & Mount Detector | 4785
3314 | SmartEDX | Install EDS PCBs into PC | 4786
3314 | SmartEDX | Installation of DPP Module & Cables | 4787
3314 | SmartEDX | Configure PC Network & Test | 4788
3314 | SmartEDX | Disable Windows10 Fast Startup | 4789
3314 | SmartEDX | Install ApexKey.edx Licences | 4790
3314 | SmartEDX | Install Genesis.edx Licences (if Applicable) | 4791
3314 | SmartEDX | Fit Dongle | 4792
3314 | SmartEDX | APEX Software Version | 4793
3314 | SmartEDX | Install APEX Software | 4794
3314 | SmartEDX | Install Remoting | 4795
3314 | SmartEDX | Setup RConfigure | 4796
3314 | SmartEDX | Run API Test | 4797
3314 | SmartEDX | Configure imaging – Run SGShell | 4798
3314 | SmartEDX | Calibrate Detector – Run eDPPShell | 4799
3314 | SmartEDX | Save Calibration Report | 4800
3314 | SmartEDX | Detector Resolution - ≥ 129 eV | 4801
3314 | SmartEDX | Save calibration data on exit | 4802
3314 | SmartEDX | Run APEX | 4803
3314 | SmartEDX | Check Spectrum | 4804
3314 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected | 4805
3314 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected | 4806
3314 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected | 4807
3314 | SmartEDX | Check Mapping | 4808
3314 | SmartEDX | Check Line Scanning | 4809
3314 | SmartEDX | Save Aquired Spectrum, Mapping and Line Scanning Images to Image Folder | 4810
3314 | SmartEDX | Save Calibration Data to the Licence Folderon the SmartSem USB stick. | 4811
3314 | SmartEDX | Save Calibration Data to the Image Folder | 4812
449 | BSD-1 | Configure BSD in Administrator | 4813
449 | BSD-1 | Mechanical Alignment | 4814
449 | BSD-1 | Input Offset Cal | 4815
449 | BSD-1 | Black Level Cal | 4816
449 | BSD-1 | Diode Segment Test | 4817
449 | BSD-1 | Diode Mirror Test | 4818
449 | BSD-1 | Vac Degradation < 3x10-5 mbar | 4819
449 | BSD-1 | All Gain Ranges working | 4820
449 | BSD-1 | Check all channels in Invert mode | 4821
451 | BSD-4 | Configure BSD in Administrator | 4822
451 | BSD-4 | Input Offset Cal | 4823
451 | BSD-4 | Black Level Cal | 4824
451 | BSD-4 | Mechanical Alignment | 4825
451 | BSD-4 | Diode Segment Test | 4826
451 | BSD-4 | Diode Mirror Test | 4827
451 | BSD-4 | Vac Degradation < 3x10-5 mbar | 4828
451 | BSD-4 | All Gain Ranges working | 4829
451 | BSD-4 | Check all channels in Invert mode | 4830
10553 | Standard W Resoltion Checks | 30KV Res- ≤2.2nm | 4831
10553 | Standard W Resoltion Checks | 3KV Res- ≤8nm | 4832
10553 | Standard W Resoltion Checks | 3KV Res- ≤7.5nm - If Unable to Achieve Contact Man Eng. | 4833
10553 | Standard W Resoltion Checks | 1KV Res- ≤15nm | 4834
10553 | Standard W Resoltion Checks | 1KV Res- ≤14nm - If Unable To Achieve Contact Man Eng. | 4835
10553 | Standard W Resoltion Checks | 200V Res- Not measured | 4836
10553 | Standard W Resoltion Checks | Store Standard Resolution Images | 4837
916 | Volutome | 1. Unpack Volutome and check all parts on the packing list are present (packing list link can be found in attachments) | 4838
916 | Volutome | 1. Upload data from USB stick to the O drive product folder | 4839
916 | Volutome | 1. Remove transit clamps from Volutome stage | 4840
916 | Volutome | 1. Ensure both safety switches are fitted to the Volutome door | 4841
916 | Volutome | 1. Check Engineers toolbox for completeness. | 4842
916 | Volutome | 1. Check Excentric pin fitment within holder. | 4843
916 | Volutome | 2. Ensure hardware has been configured according to work instructions | 4844
916 | Volutome | 2. Fit Volutome stage | 4845
916 | Volutome | 2. Check protection circuit for Piezo is fitted (349567-9013-000) | 4846
916 | Volutome | 2. Connect Volutome stage and controller to the SEM | 4847
916 | Volutome | 2. Fit Stemi holder bracket to Volutome stage door (STEMI adjustment bracket needs to fit / disconnect without force from door) | 4848
916 | Volutome | 2. Check all lights and fan come on when the electonics rack is powered on. | 4849
916 | Volutome | 3. Ethernet Connection Setup | 4850
916 | Volutome | 4. Install the latest version of Zen - Ensure to configure it for Volutome (Refer to Zen installation WI) | 4851
916 | Volutome | 4. Ensure Athena inspector manager has all green lights, Physik microtome and Physik stage are present - save a screenshot in the Volutome folder | 4852
916 | Volutome | 4. Install PI Micromove | 4853
916 | Volutome | 4. Run getversions.exe and save text file in the Volutome folder | 4854
916 | Volutome | 4. Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator | 4855
916 | Volutome | 4. Set up the Volutome workbench in Zen | 4856
916 | Volutome | 4. Athena database scan paramaters have been checked and adjusted if neccessary. | 4857
916 | Volutome | 5. PID parameters configured | 4858
916 | Volutome | 5. Retractor piezo auto zero completed (after at least 24 hours at vacuum and electronics rack powered). | 4859
916 | Volutome | 6. Record VolumeBSD serial number | 4860
916 | Volutome | 6. Record Volutome stage serial number | 4861
916 | Volutome | 6. Record VolumeBSD amplifier serial number | 4862
916 | Volutome | 6. Record VolumeBSD mechanics serial number | 4863
916 | Volutome | 6. Record C-885 PI controller serial number | 4864
916 | Volutome | 6. Record E-727 PI controller serial number | 4865
916 | Volutome | 7. Volutome 'internal LED light' function check | 4866
916 | Volutome | 8. Volutome 'External cold light' function check | 4867
916 | Volutome | 9. Adjust Volutome stage centre (Crosshairs should be within 4 central squares) | 4868
916 | Volutome | 10. Install Nav cam in-line connections (If applicable) | 4869
916 | Volutome | 11. Perform the microtome and stage functionality checks | 4870
916 | Volutome | 12. Fit and align the VolumeBSD diode | 4871
916 | Volutome | 12. Check VolumeBSD alignment repeatability | 4872
916 | Volutome | 12. Check all gain ranges of the VolumeBSD (normal/inverted) | 4873
916 | Volutome | 12. Perform Signal/noise ratio check | 4874
916 | Volutome | 12. Ensure BSD diode protector is fitted and adjusted in chamber | 4875
916 | Volutome | 12. Install cable tie to loop BSD signal cable | 4876
916 | Volutome | 12. Test BSD functionality in Zen - Insert, retract and gain ranges all functional | 4877
916 | Volutome | 13. Check SCM with Volutome stage | 4878
916 | Volutome | 14. Noise check - Check there is no visible oscillation when imaging with a pixel size of 5nm | 4879
916 | Volutome | 15. Oscillator continuity check | 4880
916 | Volutome | 16. Perform X-ray check | 4881
916 | Volutome | 17. Pump Down Time | 4882
916 | Volutome | 17. Ultimate Vacuum | 4883
916 | Volutome | 18. Focal CC fitted to Volutome stage | 4884
916 | Volutome | 18. Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar | 4885
916 | Volutome | 18. Remove carbon needle from focal CC after testing and place in the engineers toolbox container | 4886
916 | Volutome | 19. Generate live image from Axiocam | 4887
916 | Volutome | 20. Dummy experiment completed with no erros reported in Zen - Results saved in Volutome folder | 4888
916 | Volutome | 21. Re-package Volutome and check it's completeness | 4889
916 | Volutome | 22. Check functionality of all additional software installed e.g Atlas | 4890
916 | Volutome | 22. Check four standard Volutome ZEN licenses are active (ZEN core, GeminiSEM, Volutome, Automated Imaging) | 4891
916 | Volutome | 23. Ensure Arivis Software Licence is available (1 year or permanent) | 4892
1134 | Print Shipment, Packing Labels and Identify System Material | Print Shipment Labels | 4893
1134 | Print Shipment, Packing Labels and Identify System Material | Sort out customer material in box, label if possible | 4894
1134 | Print Shipment, Packing Labels and Identify System Material | Place left over stickes in Document Holder | 4895
10610 | Print Shipping Documents | Print a copy of sales order (See order inbox link) | 4896
10610 | Print Shipping Documents | Print a copy of sales order (See order inbox link) | 4897
916 | Volutome | 1. Unpack Volutome and check all parts on the packing list are present (packing list link can be found in attachments) | 4898
916 | Volutome | 1. Upload data from USB stick to the O drive product folder | 4899
916 | Volutome | 1. Remove transit clamps from Volutome stage | 4900
916 | Volutome | 1. Ensure both safety switches are fitted to the Volutome door | 4901
916 | Volutome | 1. Check Engineers toolbox for completeness. | 4902
916 | Volutome | 1. Check Excentric pin fitment within holder. | 4903
916 | Volutome | 2. Ensure hardware has been configured according to work instructions | 4904
916 | Volutome | 2. Fit Volutome stage | 4905
916 | Volutome | 2. Check protection circuit for Piezo is fitted (349567-9013-000) | 4906
916 | Volutome | 2. Connect Volutome stage and controller to the SEM | 4907
916 | Volutome | 2. Fit Stemi holder bracket to Volutome stage door (STEMI adjustment bracket needs to fit / disconnect without force from door) | 4908
916 | Volutome | 2. Check all lights and fan come on when the electonics rack is powered on. | 4909
916 | Volutome | 3. Ethernet Connection Setup | 4910
916 | Volutome | 4. Install the latest version of Zen - Ensure to configure it for Volutome (Refer to Zen installation WI) | 4911
916 | Volutome | 4. Ensure Athena inspector manager has all green lights, Physik microtome and Physik stage are present - save a screenshot in the Volutome folder | 4912
916 | Volutome | 4. Install PI Micromove | 4913
916 | Volutome | 4. Run getversions.exe and save text file in the Volutome folder | 4914
916 | Volutome | 4. Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator | 4915
916 | Volutome | 4. Set up the Volutome workbench in Zen | 4916
916 | Volutome | 5. PID parameters configured | 4917
916 | Volutome | 5. Retractor piezo auto zero completed (after at least 24 hours at vacuum and electronics rack powered). | 4918
916 | Volutome | 6. Record VolumeBSD serial number | 4919
916 | Volutome | 6. Record Volutome stage serial number | 4920
916 | Volutome | 6. Record VolumeBSD amplifier serial number | 4921
916 | Volutome | 6. Record VolumeBSD mechanics serial number | 4922
916 | Volutome | 6. Record C-885 PI controller serial number | 4923
916 | Volutome | 6. Record E-727 PI controller serial number | 4924
916 | Volutome | 7. Volutome 'internal LED light' function check | 4925
916 | Volutome | 8. Volutome 'External cold light' function check | 4926
916 | Volutome | 9. Adjust Volutome stage centre (Crosshairs should be within 4 central squares) | 4927
916 | Volutome | 10. Install Nav cam in-line connections (If applicable) | 4928
916 | Volutome | 11. Perform the microtome and stage functionality checks | 4929
916 | Volutome | 12. Fit and align the VolumeBSD diode | 4930
916 | Volutome | 12. Check VolumeBSD alignment repeatability | 4931
916 | Volutome | 12. Check all gain ranges of the VolumeBSD (normal/inverted) | 4932
916 | Volutome | 12. Perform Signal/noise ratio check | 4933
916 | Volutome | 12. Ensure BSD diode protector is fitted and adjusted in chamber | 4934
916 | Volutome | 12. Install cable tie to loop BSD signal cable | 4935
916 | Volutome | 12. Test BSD functionality in Zen - Insert, retract and gain ranges all functional | 4936
916 | Volutome | 13. Check SCM with Volutome stage | 4937
916 | Volutome | 14. Noise check - Check there is no visible oscillation when imaging with a pixel size of 5nm | 4938
916 | Volutome | 15. Oscillator continuity check | 4939
916 | Volutome | 16. Perform X-ray check | 4940
916 | Volutome | 17. Pump Down Time | 4941
916 | Volutome | 17. Ultimate Vacuum | 4942
916 | Volutome | 18. Focal CC fitted to Volutome stage | 4943
916 | Volutome | 18. Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar | 4944
916 | Volutome | 18. Remove carbon needle from focal CC after testing and place in the engineers toolbox container | 4945
916 | Volutome | 19. Generate live image from Axiocam | 4946
916 | Volutome | 20. Dummy experiment completed with no erros reported in Zen - Results saved in Volutome folder | 4947
916 | Volutome | 21. Re-package Volutome and check it's completeness | 4948
916 | Volutome | 22. Check functionality of all additional software installed e.g Atlas | 4949
916 | Volutome | 22. Check four standard Volutome ZEN licenses are active (ZEN core, GeminiSEM, Volutome, Automated Imaging) | 4950
916 | Volutome | 23. Ensure Arivis Software Licence is available (1 year or permanent) | 4951
3251 | Vac P.C.B Versions | F/W Version | 4952
3251 | Vac P.C.B Versions | F/W Version | 4953
19 | Fit DC Unit and Level control module | Fit DC unit. - 2 Person lift or lifting equipment must be used! | 4954
19 | Fit DC Unit and Level control module | Fit level control PCB | 4955
389 | SmartSEM Administrator (Column / Stage Tab) | Tick option 'Low KV Objective Cap' | 4956
3252 | EO P.C.B Versions | F/W Version NOTE MUST BE VERSION 2 IF HITTEK STAGE IS FITTED | 4957
3252 | EO P.C.B Versions | Hardware Version | 4958
1416 | Oxford EDS Detector Installation and Test | Setup Ethernet Connection | 4959
1416 | Oxford EDS Detector Installation and Test | Install Aztec Software | 4960
1416 | Oxford EDS Detector Installation and Test | Aztec Software Version | 4961
1416 | Oxford EDS Detector Installation and Test | Licence Installation | 4962
1416 | Oxford EDS Detector Installation and Test | Run Aztec Assist+ | 4963
1416 | Oxford EDS Detector Installation and Test | Install Detector | 4964
1416 | Oxford EDS Detector Installation and Test | Imaging Setup | 4965
1416 | Oxford EDS Detector Installation and Test | Stage Calibration Setup | 4966
1416 | Oxford EDS Detector Installation and Test | Calibrate Detector | 4967
1416 | Oxford EDS Detector Installation and Test | Save Aztec Sign Off PDF and Calibration Folder | 4968
3253 | EHT Set Versions | F/W Version | 4969
3253 | EHT Set Versions | Hardware Version | 4970
277 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | Ext V = | 4971
277 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | 20um Aperture Measured Speciment Current (98pA +/- 2pA)= | 4972
277 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | 30um Aperture Measured Speciment Current (220pA +/- 15pA) = | 4973
277 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | Actual Ext V = | 4974
277 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | 20um Aperture Measured Speciment Current (98pA +/- 2pA)= | 4975
277 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | 30um Aperture Measured Speciment Current (220pA +/- 15pA) = | 4976
277 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | Datasheet Ext V (200 = | 4977
277 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | Actual Ext V = | 4978
277 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | 20um Aperture Measured Speciment Current (98pA +/- 2pA)= | 4979
277 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | 30um Aperture Measured Speciment Current (220pA +/- 15pA) = | 4980
277 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | Datasheet Ext V (@200uA/sr) = | 4981
277 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | Actual Ext V = | 4982
277 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | 20um Aperture Measured Speciment Current (98pA +/- 2pA)= | 4983
277 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | 30um Aperture Measured Speciment Current (220pA +/- 15pA) = | 4984
1408 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | Datasheet Ext V (@200uA/sr) = | 4985
1408 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | Actual Ext V = | 4986
1408 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | 20um Aperture Measured Speciment Current (98pA +/- 2pA)= | 4987
1408 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | 30um Aperture Measured Speciment Current (220pA +/- 15pA) = | 4988
1408 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | 20um Aperture Measured Speciment Current (98pA +/- 2pA)= | 4989
1408 | Check Specimen Current @ 10kV WD 5mm (low kV objective) | 30um Aperture Measured Speciment Current (220pA +/- 15pA) = | 4990
1408 | Low kV Aperture Contamination Check | 20um Aperture Measured Speciment Current (98pA +/- 2pA)= | 4991
1408 | Low kV Aperture Contamination Check | 30um Aperture Measured Speciment Current (220pA +/- 15pA) = | 4992
1408 | Low kV Aperture Contamination Check | 20um Aperture Measured Speciment Current (98pA +/- 2pA)= | 4993
1408 | Low kV Aperture Contamination Check | 30um Aperture Measured Speciment Current (220pA +/- 15pA) = | 4994
1408 | Low kV Aperture Contamination Check | Confirm the specimen current ratio between apertures has not increased significantly | 4995
1408 | Low kV Aperture Contamination Check | 20um Aperture Measured Speciment Current (98pA +/- 2pA)= | 4996
1408 | Low kV Aperture Contamination Check | 30um Aperture Measured Speciment Current (220pA +/- 15pA) = | 4997
1408 | Low kV Aperture Contamination Check | Confirm the specimen current ratio between apertures has not increased significantly compared to previous measurment. | 4998
1408 | Low kV Aperture Contamination Check | 20um Aperture Measured Speciment Current (98pA +/- 2pA)= | 4999
1408 | Low kV Aperture Contamination Check | 30um Aperture Measured Speciment Current (220pA +/- 15pA) = | 5000
1408 | Low kV Aperture Contamination Check | Confirm the specimen current ratio between apertures has not increased compared to previous measurment. | 5001
1408 | Low kV Aperture Contamination Check | 20um Aperture Measured Speciment Current (98pA +/- 2pA)= | 5002
1408 | Low kV Aperture Contamination Check | 30um Aperture Measured Speciment Current = | 5003
1408 | Low kV Aperture Contamination Check | Confirm the specimen current ratio between apertures has not increased compared to previous measurment. | 5004
1409 | Low kV Aperture Contamination Check | 20um Aperture Measured Speciment Current (98pA +/- 2pA)= | 5005
1409 | Low kV Aperture Contamination Check | 30um Aperture Measured Speciment Current = | 5006
1409 | Low kV Aperture Contamination Check | Confirm the specimen current ratio between apertures has not increased compared to previous measurment. | 5007
1567 | Pre Shutdown Checks | Remove Docking Brackets and fit Transit clamps. | 5008
1567 | Pre Shutdown Checks | Airfeet deflated and free to move. | 5009
1567 | Pre Shutdown Checks | Confirm that the 3 'o' rings are fitted to the flexi hose backing line | 5010
1567 | Pre Shutdown Checks | Remove Docking Brackets and fit Transit clamps. | 5011
1567 | Pre Shutdown Checks | Airfeet deflated and free to move. | 5012
11408 | Pre Shutdown Checks (Raith) | Remove Test Tool SE2 and fit to warehouse stand. Fit SE2 blank plate to chamber and SHV connector to Spellman (00000-0302-367). | 5013
11408 | Pre Shutdown Checks (Raith) | Confirm that Raith top mounted SE2 detector is mounted to a warehouse stand and that a protective cap is fitted over the light guide. | 5014
11408 | Pre Shutdown Checks (Raith) | Remove SCM and pack for shipment (include screws) | 5015
11408 | Pre Shutdown Checks (Raith) | Remove Docking Brackets and fit Transit clamps. | 5016
11408 | Pre Shutdown Checks (Raith) | Airfeet deflated and free to move. | 5017
1567 | Pre Shutdown Checks | Remove Docking Brackets and fit Transit clamps. | 5018
1567 | Pre Shutdown Checks | Airfeet deflated and free to move. | 5019
11408 | Pre Shutdown Checks (Raith) | Remove Test Tool SE2 and fit to warehouse stand. Fit SE2 blank plate to chamber and SHV connector to Spellman (00000-0302-367). | 5020
11408 | Pre Shutdown Checks (Raith) | Confirm that Raith top mounted SE2 detector is mounted to a warehouse stand and that a protective cap is fitted over the light guide. | 5021
11408 | Pre Shutdown Checks (Raith) | Remove SCM and pack for shipment (include screws) | 5022
11408 | Pre Shutdown Checks (Raith) | Remove Docking Brackets and fit Transit clamps. | 5023
11408 | Pre Shutdown Checks (Raith) | Airfeet deflated and free to move. | 5024
10649 | Complete X-Ray Certificate | Complete X-Ray certificate | 5025
10649 | Complete X-Ray Certificate | Confirm X-Ray certificate has been signed | 5026
1424 | China RoHS Compliance Label | Fit China RoHS Compliance Label | 5027
11237 | Ensure Open Anode (Non EBBU) | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed | 5028
10531 | C2D/C2D F2 | Select Type | 5029
10531 | C2D/C2D F2 | Black Level Cal | 5030
10531 | C2D/C2D F2 | C2D Bias Target Check | 5031
10531 | C2D/C2D F2 | Gain Range Check | 5032
10531 | C2D/C2D F2 | Auto Level Check | 5033
10531 | C2D/C2D F2 | Auto Range Check | 5034
10531 | C2D/C2D F2 | Oxford EDX Detector Insulator Removed | 5035
10531 | C2D/C2D F2 | Function Image | 5036
10531 | C2D/C2D F2 | Check Beam Is On at Max VP Pressure | 5037
10531 | C2D/C2D F2 | Store C2D Functional Image | 5038
114 | C2D/C2D F2 | Black Level Cal | 5039
114 | C2D/C2D F2 | C2D Bias Target Check | 5040
114 | C2D/C2D F2 | Gain Range Check | 5041
114 | C2D/C2D F2 | Auto Level Check | 5042
114 | C2D/C2D F2 | Auto Range Check | 5043
114 | C2D/C2D F2 | Oxford EDX Detector Insulator Removed | 5044
114 | C2D/C2D F2 | Function Image | 5045
114 | C2D/C2D F2 | Store C2D Functional Image | 5046
11084 | C2D F2 | VP Mode Macro selects C2D when going to VP | 5047
11084 | C2D F2 | Black Level calibration | 5048
11084 | C2D F2 | Functional Checks ok | 5049
11084 | C2D F2 | Oxford EDX Detector Insulator Removed | 5050
11084 | C2D F2 | Take Functional Image | 5051
1153 | Fit Dynamic EO board - Double Scan (346001-9110-305) | Confirm EO PCB PN is '346001-9110-305' | 5052
1153 | Fit Dynamic EO board - Double Scan (346001-9110-305) | Fit EO Dynamic PCB | 5053
13 | Fit Dynamic EO board - Single Scan (346001-9110-315) | Confirm EO Dynamic PCB PN is '346001-9110-315' | 5054
13 | Fit Dynamic EO board - Single Scan (346001-9110-315) | Fit EO Dynamic PCB | 5055
3251 | Vac P.C.B Versions | F/W Version | 5056
13 | Fit Dynamic EO board - Single Scan (346001-9110-315) | Confirm EO Dynamic PCB PN is '346001-9110-315' | 5057
13 | Fit Dynamic EO board - Single Scan (346001-9110-315) | Fit EO Dynamic PCB | 5058
1154 | Fit Static EO board DC (346001-9100-400) | Confirm EO Static PCB PN is '346001-9100-400' | 5059
1154 | Fit Static EO board DC (346001-9100-400) | Fit EO Static PCB | 5060
13 | Fit Dynamic EO board - Single Scan (346001-9110-315) | Confirm EO Dynamic PCB PN is '346001-9110-315' | 5061
13 | Fit Dynamic EO board - Single Scan (346001-9110-315) | Fit EO Dynamic PCB | 5062
10463 | Record PC Details | Select PC | 5063
10463 | Record PC Details | PC Serial No | 5064
3242 | PC Installation | Select PC | 5065
3242 | PC Installation | PC S/No. | 5066
3242 | PC Installation | Hard Drive 1 Size | 5067
3242 | PC Installation | Hard Drive 2 Size | 5068
3242 | PC Installation | RAM | 5069
3242 | PC Installation | Processor | 5070
3242 | PC Installation | Operating System | 5071
3242 | PC Installation | Change Desktop Settings | 5072
3242 | PC Installation | Install SmartSEM and record Version | 5073
3242 | PC Installation | Check for Service Packs and install them | 5074
3242 | PC Installation | Upgrade ASAP Firmware (check Service Pack for latest version) and API | 5075
3242 | PC Installation | Copy the SmartSEM sotware into a Folder in C:Support and Name it SmartSEM X.x (i.e 8.2) Software | 5076
3242 | PC Installation | Installation of Zen Core EM if Hasap is fitted | 5077
3242 | PC Installation | Set BIOS AC Power loss to ‘ON’ | 5078
3242 | PC Installation | Set the Time and Date on the PC to current UK | 5079
3242 | PC Installation | Attach Cable Markers to USB Cables | 5080
11414 | Print and Fit S/N Label | Confirm label has CSA mark and SN and PN are correct. | 5081
11272 | Install Predictive Service Software | Predictive Service S/W installed on Customer PC. | 5082
11272 | Install Predictive Service Software | Do Not Activate | 5083
11272 | Install Predictive Service Software | Copy the Predictive Service.json file into C:Support/Data/Licence folder. DO NOT DO ON POLO SYSTEMS | 5084
11272 | Install Predictive Service Software | Predictive Service S/W installed on Customer PC. | 5085
11272 | Install Predictive Service Software | Do Not Activate | 5086
11272 | Install Predictive Service Software | Copy the Predictive Service.json file into C:Support/Data/Licence folder. | 5087
11029 | Prepare Open X Configuration | Clear EPROM on VAC Extension | 5088
10567 | Install Customer Licences | Delete the Production Test Licences | 5089
10567 | Install Customer Licences | Obtain the customer Entitlement Certificate | 5090
10567 | Install Customer Licences | Download into the C:Support/Data/Licence folder | 5091
10567 | Install Customer Licences | Install the Customer Licences | 5092
10567 | Install Customer Licences | Copy Contents of Image Folder to the Product Folder | 5093
10567 | Install Customer Licences | Copy EMServer.db to the Product Folder | 5094
10567 | Install Customer Licences | Copy EMServer.db to the C: Support/Data/Log Files folder | 5095
10567 | Install Customer Licences | Print out the Customer Entitlement Certicate - To be sent to the customer | 5096
10567 | Install Customer Licences | Store EM Server.db | 5097
10567 | Install Customer Licences | Copy EMServer.log to the Product Folder | 5098
10567 | Install Customer Licences | Store EM Server Log | 5099
10465 | Install and SetUp Windows 11 | Windows Version | 5100
10465 | Install and SetUp Windows 11 | Disable Drivers and Windows Updates | 5101
10465 | Install and SetUp Windows | Windows Version | 5102
10465 | Install and SetUp Windows | Disable Drivers and Windows Updates | 5103
10465 | Install and SetUp Windows | Windows Version | 5104
10465 | Install and SetUp Windows | Disable Drivers and Windows Updates | 5105
10465 | Install and SetUp Windows | Windows Version | 5106
10465 | Install and SetUp Windows | Disable Drivers and Windows Updates | 5107
3242 | PC Installation | Select PC | 5108
3242 | PC Installation | PC S/No. | 5109
3242 | PC Installation | Hard Drive 1 Size | 5110
3242 | PC Installation | Hard Drive 2 Size | 5111
3242 | PC Installation | RAM | 5112
3242 | PC Installation | Processor | 5113
3242 | PC Installation | Operating System | 5114
3242 | PC Installation | Change Desktop Settings | 5115
3242 | PC Installation | Install SmartSEM and record Version | 5116
3242 | PC Installation | Check for Service Packs and install them | 5117
3242 | PC Installation | Upgrade ASAP Firmware (check Service Pack for latest version) and API | 5118
3242 | PC Installation | Copy the SmartSEM sotware into a Folder in C:Support and Name it SmartSEM X.x (i.e 8.2) Software | 5119
3242 | PC Installation | Installation of Zen Core EM if Hasap is fitted | 5120
3242 | PC Installation | Set BIOS AC Power loss to ‘ON’ | 5121
3242 | PC Installation | Set the Time and Date on the PC to current UK | 5122
3242 | PC Installation | Attach Cable Markers to USB Cables | 5123
11271 | Install Predictive Service Software | Predictive Service S/W installed on Customer PC. | 5124
11271 | Install Predictive Service Software | Do Not Activate | 5125
11271 | Install Predictive Service Software | Copy the Predictive Service.json file into C:Support/Data/Licence folder. | 5126
64 | SmartSEM | Start SmartSEM and confirm there are no errors in EM server. | 5127
64 | SmartSEM | Initialise Stage can be initialised with no errors. | 5128
1422 | SmartSEM | Start SmartSEM and confirm there are no errors in EM server - With the exception of Stage related errors as no Stage / Stage PCB fitted. | 5129
1422 | SmartSEM | Start SmartSEM and confirm there are no errors in EM server - With the exception of Stage related errors as no Stage / Stage PCB fitted. | 5130
1422 | SmartSEM (Raith) | Start SmartSEM and confirm there are no errors in EM server - With the exception of Stage related errors as no Stage / Stage PCB fitted. | 5131
3321 | Install Customer Licences | Delete the Production Test Licences | 5132
3321 | Install Customer Licences | Copy EMServer.db to the C: Support/Data/Log Files folder | 5133
3321 | Install Customer Licences | Obtain the customer Entitlement Certificate | 5134
3321 | Install Customer Licences | Download into the C:Support/Data/Licence folder | 5135
3321 | Install Customer Licences | Print out the Customer Entitlement Certificate - To be sent to the customer | 5136
3321 | Install Customer Licences | Load customer licences and check against order | 5137
3321 | Install Customer Licences | Test customer licenses | 5138
3327 | Copy files to server | Em Server.db | 5139
3327 | Copy files to server | Em Server.txt Log | 5140
3327 | Copy files to server | Optiprobe.txt Log | 5141
3327 | Copy files to server | All Images | 5142
3252 | EO P.C.B Versions | F/W Version NOTE MUST BE VERSION 2 IF HITTEK STAGE IS FITTED | 5143
3252 | EO P.C.B Versions | Hardware Version | 5144
3253 | EHT Set Versions | F/W Version | 5145
3253 | EHT Set Versions | Hardware Version | 5146
3255 | Stage Versions | Controller Version | 5147
3255 | Stage Versions | Axes Version | 5148
3255 | Stage Versions | Hardware Version | 5149
3254 | ASAP / SVSEM Versions | F/W Version | 5150
3254 | ASAP / SVSEM Versions | Hardware Version | 5151
275 | Align Centre Aperture - 20nA system | Are Gun Head clamping screws secure? | 5152
275 | Align Centre Aperture - 20nA system | Tighten Gun Head clamping screws and confirm the 3 out of 4 accessible are a minimum torque of 1Nm with torque driver. | 5153
275 | Align Centre Aperture - 20nA system | Are Condensor clamping screws secure? | 5154
275 | Align Centre Aperture - 20nA system | Tighten condensor clamping screws and confirm they are a minimum torque of 1Nm with torque driver. | 5155
275 | Align Centre Aperture - 20nA system | Store 1kV emission image (Condensor off) | 5156
275 | Align Centre Aperture - 20nA system | Check emission image against 'emissions distortion criteria' | 5157
275 | Align Centre Aperture - 20nA system | Store 10kV emission image (Condensor off) | 5158
11035 | Raith On Axis Alignment Procedure | With Cond. Off Max Aper. X & Y < 30% | 5159
11035 | Raith On Axis Alignment Procedure | Gun Head Clamping Screws Secure | 5160
11035 | Raith On Axis Alignment Procedure | Tighten Gun Head clamping screws and confirm the 3 out of 4 accessible are a minimum torque of 1Nm with torque driver. | 5161
11035 | Raith On Axis Alignment Procedure | Condenser Clamping Screws Secure | 5162
11035 | Raith On Axis Alignment Procedure | Tighten condensor clamping screws and confirm they are a minimum torque of 1Nm with torque driver. | 5163
1278 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 5164
1278 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 5165
1278 | ZEN Core EM | Install Customer's ZEN Licence | 5166
1278 | ZEN Core EM | Setup SmartSEM Toggle Button | 5167
1278 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 5168
1278 | ZEN Core EM | Mains Sync Set To OFF | 5169
1278 | ZEN Core EM | Select Your ZEN core EM Runtime Profile | 5170
1278 | ZEN Core EM | Create User Supervisor EM/LM - Only required if ZEN EM and LM are installed (Note: This is rare) | 5171
1278 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 5172
1278 | ZEN Core EM | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems. | 5173
1278 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 5174
1278 | ZEN Core EM | Configue Airlock Cam and Setup Exchange Position $Exchange (use Upper case E in ZEN) | 5175
1278 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 5176
1278 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 5177
1278 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 5178
1278 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 5179
1278 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 5180
1278 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 5181
1279 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 5182
1279 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 5183
1279 | ZEN Core EM | Install Customer's ZEN Licence | 5184
1279 | ZEN Core EM | Setup SmartSEM Toggle Button | 5185
1279 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 5186
1279 | ZEN Core EM | Mains Sync Set To OFF | 5187
1279 | ZEN Core EM | Choose your ZEN core EM runtime profile | 5188
1279 | ZEN Core EM | Create User Supervisor EM/LM - Only required if ZEN EM and LM are installed | 5189
1279 | ZEN Core EM | Create User Supervisor Volutome (only applicable for Volutome systems) | 5190
1279 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 5191
1279 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 5192
1279 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 5193
1279 | ZEN Core EM | Configue $NavCam (Capital N and C) | 5194
1279 | ZEN Core EM | Configue $Exchange (capital E) | 5195
1279 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 5196
1279 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 5197
1279 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 5198
1279 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 5199
1279 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 5200
1279 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 5201
10508 | Stigmator & Alignment Check -EDS BeamSleeve (500um) | EDS BeamSleeve 500um 30KV @ 8mm X- ≤35% | 5202
10508 | Stigmator & Alignment Check -EDS BeamSleeve (500um) | EDS BeamSleeve 500um 30KV @ 8mm Y- ≤35% | 5203
10508 | Stigmator & Alignment Check -EDS BeamSleeve (500um) | EDS BeamSleeve 500um Alignment- ≤70% | 5204
1280 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 5205
1280 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 5206
1280 | ZEN Core EM | Install Customer's ZEN Licence | 5207
1280 | ZEN Core EM | Setup SmartSEM Toggle Button | 5208
1280 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 5209
1280 | ZEN Core EM | Mains Sync Set To OFF | 5210
1280 | ZEN Core EM | Choose your ZEN core EM runtime profile | 5211
1280 | ZEN Core EM | Create User Supervisor EM/LM - Only required if ZEN EM and LM are installed | 5212
1280 | ZEN Core EM | Create User Supervisor Volutome (only applicable for Volutome systems) | 5213
1280 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 5214
1280 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 5215
1280 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 5216
1280 | ZEN Core EM | Configue $NavCam (Capital N and C) | 5217
1280 | ZEN Core EM | Configue $Exchange (capital E) | 5218
1280 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 5219
1280 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 5220
1280 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 5221
1280 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 5222
1280 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 5223
1280 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 5224
1280 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 5225
1280 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 5226
1280 | ZEN Core EM | Install Customer's ZEN Licence | 5227
1280 | ZEN Core EM | Setup SmartSEM Toggle Button | 5228
1280 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 5229
1280 | ZEN Core EM | Mains Sync Set To OFF | 5230
1280 | ZEN Core EM | Choose your ZEN core EM runtime profile | 5231
1280 | ZEN Core EM | Create User Supervisor EM/LM - Only required if ZEN EM and LM are installed. This is rare! | 5232
1280 | ZEN Core EM | Create User Supervisor Volutome (only applicable for Volutome systems) | 5233
1280 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 5234
1280 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 5235
1280 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 5236
1280 | ZEN Core EM | Configue $NavCam (Capital N and C) | 5237
1280 | ZEN Core EM | Configue $Exchange (capital E) | 5238
1280 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 5239
1280 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 5240
1280 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 5241
1280 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 5242
1280 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 5243
1280 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 5244
10509 | Stigmator & Alignment Check - EDS BeamSleeve (1000um) | EDS BeamSleeve 1000um 30KV @ 8mm X- ≤35% | 5245
10509 | Stigmator & Alignment Check - EDS BeamSleeve (1000um) | EDS BeamSleeve 1000um 30KV @ 8mm Y- ≤35% | 5246
10509 | Stigmator & Alignment Check - EDS BeamSleeve (1000um) | EDS BeamSleeve 1000um Alignment- ≤70% | 5247
252 | Configure PC ASAP | PC S/N | 5248
252 | Configure PC ASAP | PC/Model (SAP No.) | 5249
252 | Configure PC ASAP | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448 | 5250
252 | Configure PC ASAP | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448 , then the PC needs reimaging – follow the attached WI. The correct image version is V2. | 5251
252 | Configure PC ASAP | Hard disk Image (Version) | 5252
252 | Configure PC ASAP | Record Operating System | 5253
252 | Configure PC ASAP | Check the Entitlement certificate to determine which version of Smartsem to install | 5254
252 | Configure PC ASAP | Install Smart SEM (See WI 9693) Record Version | 5255
252 | Configure PC ASAP | Check for available SmartSEMService Packs and install them. (N/A if none) | 5256
252 | Configure PC ASAP | Check for available SmartSEM hotfixes and install them. (N/A if none) | 5257
252 | Configure PC ASAP | Installation of Zen Core EM | 5258
252 | Configure PC ASAP | Install CZ Support Tools. Ver | 5259
252 | Configure PC ASAP | Disable Automatic Updates | 5260
252 | Configure PC ASAP | Run ASAP Screening Program | 5261
252 | Configure PC ASAP | System Specific Datazone Loaded | 5262
252 | Configure PC ASAP | In SEM status set 'Mains Sync' to off | 5263
3242 | PC Installation | Select PC | 5264
3242 | PC Installation | PC S/No. | 5265
3242 | PC Installation | Hard Drive 1 Size | 5266
3242 | PC Installation | Hard Drive 2 Size | 5267
3242 | PC Installation | RAM | 5268
3242 | PC Installation | Processor | 5269
3242 | PC Installation | Operating System | 5270
3242 | PC Installation | Change Desktop Settings | 5271
3242 | PC Installation | Install SmartSEM and record Version | 5272
3242 | PC Installation | Check for Service Packs and install them | 5273
3242 | PC Installation | Upgrade ASAP Firmware (check Service Pack for latest version) and API | 5274
3242 | PC Installation | Copy the SmartSEM sotware into a Folder in C:Support and Name it SmartSEM X.x (i.e 8.2) Software | 5275
3242 | PC Installation | Installation of Zen Core EM if Hasap is fitted | 5276
3242 | PC Installation | Set BIOS AC Power loss to ‘ON’ | 5277
3242 | PC Installation | Set the Time and Date on the PC to current UK | 5278
3242 | PC Installation | Attach Cable Markers to USB Cables | 5279
11028 | On/Off Switch Panel | Curved Switch Panel Fitted. NOTE: Only applicable to MC OEM2 and not GS360 OEM7. | 5280
10498 | Condition Gun - LAB6 | Connect The Earth Cable To the Column | 5281
10498 | Condition Gun - LAB6 | Connect the Gun Cable to the Test Cable HV Socket | 5282
10498 | Condition Gun - LAB6 | Check that all the connections are correct | 5283
10498 | Condition Gun - LAB6 | Run up to 35KV and leave for 30 minute | 5284
10498 | Condition Gun - LAB6 | No Strom Current is Displayed | 5285
10498 | Condition Gun - LAB6 | Check No Spikes in Gun Vacuum Trace | 5286
10498 | Condition Gun - LAB6 | Switch Off the Heinzinger and disconnect | 5287
10498 | Condition Gun - LAB6 | Remove interlock cable and plug the gun interlock cable back into the Vice PCB. | 5288
10498 | Condition Gun - LAB6 | Reconnect the Gun Cable to the Spellman EHT Set | 5289
10498 | Condition Gun - LAB6 | Connect The Earth Cable To the Column | 5290
10498 | Condition Gun - LAB6 | Connect the Gun Cable to the Test Cable HV Socket | 5291
10498 | Condition Gun - LAB6 | Connect the Interlock Cable | 5292
10498 | Condition Gun - LAB6 | Check that all the connections are correct | 5293
10498 | Condition Gun - LAB6 | Run up to 35KV and leave for 30 minute | 5294
10498 | Condition Gun - LAB6 | No Strom Current is Displayed | 5295
10498 | Condition Gun - LAB6 | Check No Spikes in Gun Vacuum Trace | 5296
10498 | Condition Gun - LAB6 | Switch Off the Heinzinger and disconnect | 5297
10498 | Condition Gun - LAB6 | Remove interlock cable and plug the gun interlock cable back into the Vice PCB. | 5298
10498 | Condition Gun - LAB6 | Reconnect the Gun Cable to the Spellman EHT Set | 5299
10498 | Condition Gun - LAB6 | Connect The Earth Cable To the Column | 5300
10498 | Condition Gun - LAB6 | Connect the Gun Cable to the Test Cable HV Socket | 5301
10498 | Condition Gun - LAB6 | Connect the Interlock Cable | 5302
10498 | Condition Gun - LAB6 | Check that all the connections are correct | 5303
10498 | Condition Gun - LAB6 | Conect the Mains to the Heinzinger PSU | 5304
10498 | Condition Gun - LAB6 | Run up to 35KV and leave for 30 minute | 5305
10498 | Condition Gun - LAB6 | No Strom Current is Displayed | 5306
10498 | Condition Gun - LAB6 | Check No Spikes in Gun Vacuum Trace | 5307
10498 | Condition Gun - LAB6 | Switch Off the Heinzinger and disconnect | 5308
10498 | Condition Gun - LAB6 | Remove interlock cable and plug the gun interlock cable back into the Vice PCB. | 5309
10498 | Condition Gun - LAB6 | Reconnect the Gun Cable to the Spellman EHT Set | 5310
10498 | Condition Gun - LAB6 | Connect The Earth Cable To the Column | 5311
10498 | Condition Gun - LAB6 | Connect the Gun Cable to the Test Cable HV Socket Using the Modified Clamp Ring | 5312
10498 | Condition Gun - LAB6 | Connect the Interlock Cable | 5313
10498 | Condition Gun - LAB6 | Check that all the connections are correct | 5314
10498 | Condition Gun - LAB6 | Conect the Mains to the Heinzinger PSU | 5315
10498 | Condition Gun - LAB6 | Run up to 35KV and leave for 30 minute | 5316
10498 | Condition Gun - LAB6 | No Strom Current is Displayed | 5317
10498 | Condition Gun - LAB6 | Check No Spikes in Gun Vacuum Trace | 5318
10498 | Condition Gun - LAB6 | Switch Off the Heinzinger and disconnect the Mains Cable | 5319
10498 | Condition Gun - LAB6 | Refit the Gun Cable with the Original Clamp Ring. | 5320
10498 | Condition Gun - LAB6 | Remove interlock cable and plug the gun interlock cable back into the Vice PCB. | 5321
10498 | Condition Gun - LAB6 | Reconnect the Gun Cable to the Spellman EHT Set | 5322
10498 | Condition Gun - LAB6 | Connect The Earth Cable To the Column | 5323
10498 | Condition Gun - LAB6 | Connect the Gun Cable to the Test Cable HV Socket Using the Modified Clamp Ring | 5324
10498 | Condition Gun - LAB6 | Connect the Interlock Cable | 5325
10498 | Condition Gun - LAB6 | Check that all the connections are correct | 5326
10498 | Condition Gun - LAB6 | Conect the Mains to the Heinzinger PSU | 5327
10498 | Condition Gun - LAB6 | Run up to 35KV and leave for 30 minute | 5328
10498 | Condition Gun - LAB6 | No Strom Current is Displayed | 5329
10498 | Condition Gun - LAB6 | Check No Spikes in Gun Vacuum Trace | 5330
10498 | Condition Gun - LAB6 | Switch Off the Heinzinger and disconnect the Mains Cable | 5331
10498 | Condition Gun - LAB6 | Refit the Gun Cable to the Spellman EHT Set with the Original Clamp Ring. | 5332
10498 | Condition Gun - LAB6 | Remove interlock cable and plug the gun interlock cable back into the Vice PCB. | 5333
10498 | Condition Gun - LAB6 | Reconnect the Gun Cable to the Spellman EHT Set | 5334
11028 | On/Off Switch Panel | Curved Switch Panel Fitted. | 5335
10498 | Condition Gun - LAB6 | Connect The Earth Cable To the Column | 5336
10498 | Condition Gun - LAB6 | Connect the Gun Cable to the Test Cable HV Socket Using the Modified Clamp Ring | 5337
10498 | Condition Gun - LAB6 | Connect the Interlock Cable | 5338
10498 | Condition Gun - LAB6 | Check that all the connections are correct | 5339
10498 | Condition Gun - LAB6 | Conect the Mains to the Heinzinger PSU | 5340
10498 | Condition Gun - LAB6 | Run up to 35KV and leave for 30 minute | 5341
10498 | Condition Gun - LAB6 | No Strom Current is Displayed | 5342
10498 | Condition Gun - LAB6 | Check No Spikes in Gun Vacuum Trace | 5343
10498 | Condition Gun - LAB6 | Switch Off the Heinzinger and disconnect the Mains Cable | 5344
10498 | Condition Gun - LAB6 | Remove interlock cable and plug the gun interlock cable back into the Vice PCB. | 5345
10498 | Condition Gun - LAB6 | Reconnect the Gun Cable to the Spellman EHT Set using the Original Clamp Ring. | 5346
10498 | Condition Gun - LAB6 | Remove the Earth Cable From the Column | 5347
10498 | Condition Gun - LAB6 | Connect The Earth Cable To the Column | 5348
10498 | Condition Gun - LAB6 | Connect the Gun Cable to the Test Cable HV Socket Using the Modified Clamp Ring | 5349
10498 | Condition Gun - LAB6 | Connect the Interlock Cable | 5350
10498 | Condition Gun - LAB6 | Check that all the connections are correct | 5351
10498 | Condition Gun - LAB6 | Please sign the logbook to certify that you have inspected the equipment and confirm it is in good condition to use | 5352
10498 | Condition Gun - LAB6 | Conect the Mains to the Heinzinger PSU | 5353
10498 | Condition Gun - LAB6 | Run up to 35KV and leave for 30 minute | 5354
10498 | Condition Gun - LAB6 | No Strom Current is Displayed | 5355
10498 | Condition Gun - LAB6 | Check No Spikes in Gun Vacuum Trace | 5356
10498 | Condition Gun - LAB6 | Switch Off the Heinzinger and disconnect the Mains Cable | 5357
10498 | Condition Gun - LAB6 | Remove interlock cable and plug the gun interlock cable back into the Vice PCB. | 5358
10498 | Condition Gun - LAB6 | Reconnect the Gun Cable to the Spellman EHT Set using the Original Clamp Ring. | 5359
10498 | Condition Gun - LAB6 | Remove the Earth Cable From the Column | 5360
1279 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 5361
1279 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 5362
1279 | ZEN Core EM | Install Customer's ZEN Licence | 5363
1279 | ZEN Core EM | Setup SmartSEM Toggle Button | 5364
1279 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 5365
1279 | ZEN Core EM | Mains Sync Set To OFF | 5366
1279 | ZEN Core EM | Choose your ZEN core EM runtime profile | 5367
1279 | ZEN Core EM | Create User Supervisor EM/LM - Only required if ZEN EM and LM are installed. This is rare! | 5368
1279 | ZEN Core EM | Create User Supervisor Volutome (only applicable for Volutome systems) | 5369
1279 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 5370
1279 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 5371
1279 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 5372
1279 | ZEN Core EM | Configue $NavCam (Capital N and C) | 5373
1279 | ZEN Core EM | Configue $Exchange (capital E) | 5374
1279 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 5375
1279 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 5376
1279 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 5377
1279 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 5378
1279 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 5379
1279 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 5380
10508 | Stigmator & Alignment Check -EDS BeamSleeve (500um) | BeamSleeve 500um 30KV @ 8mm X- ≤35% | 5381
10508 | Stigmator & Alignment Check -EDS BeamSleeve (500um) | BeamSleeve 500um 30KV @ 8mm Y- ≤35% | 5382
10508 | Stigmator & Alignment Check -EDS BeamSleeve (500um) | BeamSleeve 500um Alignment- ≤70% | 5383
11030 | Configure Open X Interfaces X21 & X121 | Configue Raith Open X Interface | 5384
11410 | Set default Aligning System Offset Angle | Align offset angle set to -110 degrees. | 5385
11410 | Set default Aligning System Offset Angle (Raith MC OEM2) | Align offset angle set to -110 degrees. | 5386
1428 | Set default Aligning System Offset Angle (Gemini 360 OEM7) | Align offset angle set to -140 degrees. | 5387
11037 | Calibrate Stage Angle | Stage Angle is close to 133Degs | 5388
11037 | Calibrate Stage Angle (Raith MC OEM2) | Stage Angle is close to 133Degs | 5389
1429 | Calibrate Stage Angle (Gemini 360 OEM7) | Stage Angle is close to 163Degs | 5390
10509 | Stigmator & Alignment Check - EDS BeamSleeve (1000um) | BeamSleeve 1000um 30KV @ 8mm X- ≤35% | 5391
10509 | Stigmator & Alignment Check - EDS BeamSleeve (1000um) | BeamSleeve 1000um 30KV @ 8mm Y- ≤35% | 5392
10509 | Stigmator & Alignment Check - EDS BeamSleeve (1000um) | BeamSleeve 1000um Alignment- ≤70% | 5393
10464 | Power Up and Enter Bios | Check That AC Loss Policy Depend On, is Set To Power On | 5394
10464 | Power Up and Enter Bios | Set the Time and Date on the PC to current UK | 5395
10464 | Power Up and Enter Bios | Disable Windows Fast Start UP | 5396
10464 | Power Up and Enter Bios | Disable USB Power Save | 5397
10464 | Power Up and Enter Bios | Check That AC Loss Policy Depend On, is Set To Power On | 5398
10464 | Power Up and Enter Bios | Set the Time and Date on the PC to current UK | 5399
10464 | Power Up and Enter Bios | Disable Windows Fast Start UP | 5400
10464 | Power Up and Enter Bios | Disable USB Power Save | 5401
10464 | Power Up and Enter Bios | Check PC Part Number | 5402
10464 | Power Up and Enter Bios | Check That AC Loss Policy Depend On, is Set To Power On | 5403
10464 | Power Up and Enter Bios | Set the Time and Date on the PC to current UK | 5404
10464 | Power Up and Enter Bios | Disable Windows Fast Start UP | 5405
10464 | Power Up and Enter Bios | Disable USB Power Save | 5406
10464 | Power Up and Enter Bios | Check PC Part Number | 5407
10464 | Power Up and Enter Bios | Check Windows 11 Certification label is fitted | 5408
3242 | PC Installation | Select PC | 5409
3242 | PC Installation | PC S/No. | 5410
3242 | PC Installation | Hard Drive 1 Size | 5411
3242 | PC Installation | Hard Drive 2 Size | 5412
3242 | PC Installation | RAM | 5413
3242 | PC Installation | Processor | 5414
3242 | PC Installation | Operating System | 5415
3242 | PC Installation | Change Desktop Settings | 5416
3242 | PC Installation | Install SmartSEM and record Version | 5417
3242 | PC Installation | Check for Service Packs and install them | 5418
3242 | PC Installation | Upgrade ASAP Firmware (check Service Pack for latest version) and API | 5419
3242 | PC Installation | Copy the SmartSEM sotware into a Folder in C:Support and Name it SmartSEM X.x (i.e 8.2) Software | 5420
3242 | PC Installation | Installation of Zen Core EM if Hasap is fitted | 5421
3242 | PC Installation | Check That AC Loss Policy Depend On, is Set To Power On | 5422
3242 | PC Installation | Set the Time and Date on the PC to current UK | 5423
3242 | PC Installation | Disable Windows Fast Start UP | 5424
3242 | PC Installation | Disable USB Power Save | 5425
3242 | PC Installation | Check PC Part Number | 5426
3242 | PC Installation | Check Windows 11 Certification label is fitted | 5427
3242 | PC Installation | Attach Cable Markers to USB Cables | 5428
10509 | Stigmator & Alignment Check - EDS BeamSleeve (1000um) | EDS BeamSleeve 1000um 30KV @ 8mm X- ≤35% | 5429
10509 | Stigmator & Alignment Check - EDS BeamSleeve (1000um) | EDS BeamSleeve 1000um 30KV @ 8mm Y- ≤35% | 5430
10509 | Stigmator & Alignment Check - EDS BeamSleeve (1000um) | BeamSleeve 1000um Alignment- ≤70% | 5431
3242 | PC Installation | Select PC | 5432
3242 | PC Installation | PC S/No. | 5433
3242 | PC Installation | Operating System | 5434
3242 | PC Installation | Change Desktop Settings | 5435
3242 | PC Installation | Install SmartSEM and record Version | 5436
3242 | PC Installation | Check for Service Packs and install them | 5437
3242 | PC Installation | Upgrade ASAP Firmware (check Service Pack for latest version) and API | 5438
3242 | PC Installation | Copy the SmartSEM sotware into a Folder in C:Support and Name it SmartSEM X.x (i.e 8.2) Software | 5439
3242 | PC Installation | Installation of Zen Core EM if Hasap is fitted | 5440
3242 | PC Installation | Check That AC Loss Policy Depend On, is Set To Power On | 5441
3242 | PC Installation | Set the Time and Date on the PC to current UK | 5442
3242 | PC Installation | Disable Windows Fast Start UP | 5443
3242 | PC Installation | Disable USB Power Save | 5444
3242 | PC Installation | Check PC Part Number | 5445
3242 | PC Installation | Check Windows 11 Certification label is fitted | 5446
3242 | PC Installation | Attach Cable Markers to USB Cables | 5447
10602 | Magnification Calibration | Check Calibration in each mag range. DO NOT USE THE CAL WIZARD. JUST CHECK USING THE ANNOTATION CURSORS | 5448
10602 | Magnification Calibration | Check Stage Angle has been setup correctly | 5449
13 | Fit Dynamic EO board - Single Scan (346001-9110-315) | Confirm EO Dynamic PCB PN is '346001-9110-315' | 5450
13 | Fit Dynamic EO board - Single Scan (346001-9110-315) | Fit EO Dynamic PCB (Record SN) | 5451
1153 | Fit Dynamic EO board - Double Scan (346001-9110-305) | Confirm EO PCB PN is '346001-9110-305' | 5452
1153 | Fit Dynamic EO board - Double Scan (346001-9110-305) | Fit EO Dynamic PCB | 5453
1153 | Fit Dynamic EO board - Double Scan (346001-9110-305) | Confirm EO PCB PN is '346001-9110-305' | 5454
1153 | Fit Dynamic EO board - Double Scan (346001-9110-305) | Fit EO Dynamic PCB (Record SN) | 5455
15 | Fit Vac board | Fit Vac PCB (Record SN) | 5456
14 | Fit Static EO board SC (346001-9100-410) | Confirm EO Static PCB PN is '346001-9100-410' | 5457
14 | Fit Static EO board SC (346001-9100-410) | Fit EO Static PCB (Record SN) | 5458
329 | Stigmation / Aperture Checks | Stig: 30μm, All EHT’s are ≤ 20% | 5459
329 | Stigmation / Aperture Checks | Stig: Off Axis Aper, All EHT’s are ≤ 30% | 5460
329 | Stigmation / Aperture Checks | Aper/Gun: 30μm, All EHT’s (Excluding 99v) are ≤ 60% | 5461
705 | Raith Stigmation /Aperture Checks | Stigs: 30um Aperture, All EHT ≤20% | 5462
705 | Raith Stigmation /Aperture Checks | Stigs: Off Axis Apertures, All EHT ≤30% | 5463
705 | Raith Stigmation /Aperture Checks | Aper & Gun: 30um Aperture, All EHT ≤50% | 5464
461 | SC Manual Tablefiller Entries (Raith) | 22kv - All Apertures at Normal and High current WD5 | 5465
461 | SC Manual Tablefiller Entries (Raith) | 30kv - Aperture 1 and 7 at Normal and High current at WD5 | 5466
1440 | Fit KC Label | KC Label fitted | 5467
1236 | Print and Fit SN Label (NO CSA, NO SEMI) | ✘ 𝗡𝗢 CSA marking on SN label.. | 5468
1236 | Print and Fit SN Label (NO CSA, NO SEMI) | ✘ 𝗡𝗢 SEMI marking on SN label. | 5469
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | ✘ 𝗡𝗢 CSA marking on SN label. | 5470
11345 | Print and Fit S/N Label (SEMI S2 S8 certification) | ✔ SEMI marking is on SN label. | 5471
1154 | Fit Static EO board DC (346001-9100-400) | Confirm EO Static PCB PN is '346001-9100-400' | 5472
1154 | Fit Static EO board DC (346001-9100-400) | Fit EO Static PCB | 5473
1234 | Print and Fit S/N Label (CSA UL certification) | ✔ CSA marking is on SN label. | 5474
1234 | Print and Fit S/N Label (CSA UL certification) | ✘ 𝗡𝗢 SEMI marking on SN label. | 5475
1235 | Print and Fit S/N Label ( SEMI S2 S8 certification + CSA UL certification) | ✔ CSA marking is on SN label. | 5476
1235 | Print and Fit S/N Label ( SEMI S2 S8 certification + CSA UL certification) | ✔ SEMI marking is on SN label. | 5477
11426 | 560 Ressoution Measurements | 15kv 0.7nm | 5478
11426 | 560 Ressoution Measurements | Store 15KV Resolution Images and Pate image for 2 x consecutive images | 5479
11426 | 560 Ressoution Measurements | 1kv 1.2nm | 5480
11426 | 560 Ressoution Measurements | Store 1KV Resolution Images and Pate image for 2 x consecutive images | 5481
11426 | 560 Ressoution Measurements | 500V 1.5nm | 5482
11426 | 560 Ressoution Measurements | Store 500V Resolution Images and Pate image for 2 x consecutive images | 5483
11425 | 360 Ressoution Measurements | 15kv 0.9nm | 5484
11425 | 360 Ressoution Measurements | Store 15KV Resolution Images and Pate image for 2 x consecutive images | 5485
11425 | 360 Ressoution Measurements | 1kv 1.3nm | 5486
11425 | 360 Ressoution Measurements | Store 1KV Resolution Images and Pate image for 2 x consecutive images | 5487
11425 | 360 Ressoution Measurements | 500V 1.9nm | 5488
11425 | 360 Ressoution Measurements | Store 500V Resolution Images and Pate image for 2 x consecutive images | 5489
3292 | Standard Resolution (2 x consecutive images) | 1kV (Spec: $[Resolution1kV]nm) | 5490
3292 | Standard Resolution (2 x consecutive images) | Store 1KV Resolution Images and Pate image for 2 x consecutive images | 5491
3292 | Standard Resolution (2 x consecutive images) | 15kV (Spec: $[Resolution15kV]nm) | 5492
3292 | Standard Resolution (2 x consecutive images) | Store 15KV Resolution Images and Pate image for 2 x consecutive images | 5493
3292 | Standard Resolution (2 x consecutive images) | 1kV (Spec: $[Resolution1kV]nm) | 5494
3292 | Standard Resolution (2 x consecutive images) | Store 1KV Resolution Images and Pate image for 2 x consecutive images | 5495
3292 | Standard Resolution (2 x consecutive images) | 15kV (Spec: $[Resolution15kV]nm) | 5496
3292 | Standard Resolution (2 x consecutive images) | Store 15KV Resolution Images and Pate image for 2 x consecutive images | 5497
3292 | Standard Resolution (2 x consecutive images) | 1kV (Spec: $[Resolution1kV]nm) | 5498
3292 | Standard Resolution (2 x consecutive images) | Store 1KV Resolution Images and Pate image for 2 x consecutive images | 5499
3292 | Standard Resolution (2 x consecutive images) | 15kV (Spec: $[Resolution15kV]nm) | 5500
3292 | Standard Resolution (2 x consecutive images) | Store 15KV Resolution Images and Pate image for 2 x consecutive images | 5501
3292 | Standard Resolution (2 x consecutive images) | 1kV (Spec: $[Resolution1kV]nm) | 5502
3292 | Standard Resolution (2 x consecutive images) | Store 1KV Resolution Images and Pate image for 2 x consecutive images | 5503
3292 | Standard Resolution (2 x consecutive images) | 15kV (Spec: $[Resolution15kV]nm) | 5504
3292 | Standard Resolution (2 x consecutive images) | Store 15KV Resolution Images and Pate image for 2 x consecutive images | 5505
1154 | Fit Static EO board DC (346001-9100-400) | Confirm EO Static PCB PN is '346001-9100-400' | 5506
1154 | Fit Static EO board DC (346001-9100-400) | Fit EO Static PCB (Record SN) | 5507
3246 | Connect Plinth Services | Air Pressure O/P = 6bar | 5508
3246 | Connect Plinth Services | Water | 5509
3246 | Connect Plinth Services | Connect Mains, record SN from power supply box (344732-9010-000) | 5510
3246 | Connect Plinth Services | Ensure Test tool Quiet Mode is set to 'ON' in service cupboard (Unless customer Quiet mode fitted) | 5511
8699 | Cable Up (240V Mains) | System Mains Voltage | 5512
8699 | Cable Up (240V Mains) | IF LaB6 and 120V - Transformer Installed | 5513
8699 | Cable Up (240V Mains) | Cable Up | 5514
8699 | Cable Up (240V Mains) | Label Cables | 5515
1405 | Cable Up (120V Mains) | System Mains Voltage | 5516
1405 | Cable Up (120V Mains) | IF LaB6 and 120V - Transformer Installed | 5517
1405 | Cable Up (120V Mains) | Cable Up | 5518
1405 | Cable Up (120V Mains) | Label Cables | 5519
1442 | Cable Up (240V Mains) | System Mains Voltage | 5520
1442 | Cable Up (240V Mains) | IF LaB6 and 120V - Transformer Installed | 5521
1442 | Cable Up (240V Mains) | Cable Up | 5522
1442 | Cable Up (240V Mains) | Label Cables | 5523
10506 | VP Functional Checks | Vp Functional Check - Switch Off the TestTool Quite Mode Before Proceeding. | 5524
10506 | VP Functional Checks | Vac Status = Ready and EHT Vac Ready = Yes, are always present at all VP values. (i.e. The beam remains on/can be switched on). | 5525
10506 | VP Functional Checks | Fit Label to the Black Plastic Aperture Box | 5526
3275 | VP Functional Checks | VP Functional Check - Switch Off the TestTool Quite Mode Before Proceeding. | 5527
1442 | Cable Up (240V Mains) | System Mains Voltage | 5528
1442 | Cable Up (240V Mains) | Cable Up | 5529
1442 | Cable Up (240V Mains) | Label Cables | 5530
8699 | Cable Up ( Lab6 240V Mains) | System Mains Voltage | 5531
8699 | Cable Up ( Lab6 240V Mains) | IF LaB6 and 120V - Transformer Installed | 5532
8699 | Cable Up ( Lab6 240V Mains) | Cable Up | 5533
8699 | Cable Up ( Lab6 240V Mains) | Fit LAB6 Heater Control, Record SN: | 5534
8699 | Cable Up ( Lab6 240V Mains) | Label Cables | 5535
1405 | Cable Up (Lab6 120V Mains) | System Mains Voltage | 5536
1405 | Cable Up (Lab6 120V Mains) | IF LaB6 and 120V - Transformer Installed | 5537
1405 | Cable Up (Lab6 120V Mains) | Fit LAB6 Heater Control, Record SN: | 5538
1405 | Cable Up (Lab6 120V Mains) | Cable Up | 5539
1405 | Cable Up (Lab6 120V Mains) | Label Cables | 5540
1442 | Cable Up (240V Mains) | System Mains Voltage | 5541
1442 | Cable Up (240V Mains) | Cable Up | 5542
1442 | Cable Up (240V Mains) | Label Cables | 5543
1442 | Cable Up (240V Mains) | System Mains Voltage | 5544
1442 | Cable Up (240V Mains) | Cable Up | 5545
1442 | Cable Up (240V Mains) | Label Cables | 5546
1442 | Cable Up (240V Mains) | System Mains Voltage | 5547
1442 | Cable Up (240V Mains) | Cable Up | 5548
1442 | Cable Up (240V Mains) | Label Cables | 5549
1405 | Cable Up (Lab6 120V Mains) | System Mains Voltage | 5550
1405 | Cable Up (Lab6 120V Mains) | IF LaB6 and 120V - Transformer Installed | 5551
1405 | Cable Up (Lab6 120V Mains) | Cable Up | 5552
1405 | Cable Up (Lab6 120V Mains) | Fit LAB6 Heater Control, Record SN: | 5553
1405 | Cable Up (Lab6 120V Mains) | Label Cables | 5554
8699 | Cable Up ( Lab6 240V Mains) | System Mains Voltage | 5555
8699 | Cable Up ( Lab6 240V Mains) | IF LaB6 and 120V - Transformer Installed | 5556
8699 | Cable Up ( Lab6 240V Mains) | Cable Up | 5557
8699 | Cable Up ( Lab6 240V Mains) | Fit LAB6 Heater Control Box, Record SN: | 5558
8699 | Cable Up ( Lab6 240V Mains) | Label Cables | 5559
1405 | Cable Up (Lab6 120V Mains) | System Mains Voltage | 5560
1405 | Cable Up (Lab6 120V Mains) | IF LaB6 and 120V - Transformer Installed | 5561
1405 | Cable Up (Lab6 120V Mains) | Cable Up | 5562
1405 | Cable Up (Lab6 120V Mains) | Fit LAB6 Heater Control Box, Record SN: | 5563
1405 | Cable Up (Lab6 120V Mains) | Label Cables | 5564
252 | Configure PC ASAP | PC S/N | 5565
252 | Configure PC ASAP | PC/Model (SAP No.) | 5566
252 | Configure PC ASAP | Hard disk Image (Version) | 5567
252 | Configure PC ASAP | Record Operating System | 5568
252 | Configure PC ASAP | Check the Entitlement certificate to determine which version of Smartsem to install | 5569
252 | Configure PC ASAP | Install Smart SEM (See WI 9693) Record Version | 5570
252 | Configure PC ASAP | Check for available SmartSEMService Packs and install them. (N/A if none) | 5571
252 | Configure PC ASAP | Check for available SmartSEM hotfixes and install them. (N/A if none) | 5572
252 | Configure PC ASAP | Installation of Zen Core EM | 5573
252 | Configure PC ASAP | Install CZ Support Tools. Ver | 5574
252 | Configure PC ASAP | Disable Automatic Updates | 5575
252 | Configure PC ASAP | Run ASAP Screening Program | 5576
252 | Configure PC ASAP | System Specific Datazone Loaded | 5577
252 | Configure PC ASAP | In SEM status set 'Mains Sync' to off | 5578
1443 | PC Power Up and Enter Bios | Check That AC Loss Policy Depend On, is Set To Power On | 5579
1443 | PC Power Up and Enter Bios | Set the Time and Date on the PC to current UK | 5580
1443 | PC Power Up and Enter Bios | Disable Windows Fast Start UP | 5581
1443 | PC Power Up and Enter Bios | Disable USB Power Save | 5582
1443 | PC Power Up and Enter Bios | Check PC Part Number | 5583
1443 | PC Power Up and Enter Bios | Check Windows 11 Certification label is fitted | 5584
1443 | PC Power Up and Enter Bios | Check That AC Loss Policy Depend On, is Set To Power On | 5585
1443 | PC Power Up and Enter Bios | Set the Time and Date on the PC to current UK | 5586
1443 | PC Power Up and Enter Bios | Disable Windows Fast Start UP | 5587
1443 | PC Power Up and Enter Bios | Disable USB Power Save | 5588
1443 | PC Power Up and Enter Bios | Check PC Part Number | 5589
1443 | PC Power Up and Enter Bios | Check Windows 11 Certification label is fitted | 5590
482 | Run Tablefiller | Mu metal shield is fitted | 5591
482 | Run Tablefiller | DC Tablefiller successfully completed . | 5592
252 | Configure PC ASAP | PC S/N | 5593
252 | Configure PC ASAP | Hard disk Image (Version) | 5594
252 | Configure PC ASAP | Record Operating System | 5595
252 | Configure PC ASAP | Check the Entitlement certificate to determine which version of Smartsem to install | 5596
252 | Configure PC ASAP | Install Smart SEM (See WI 9693) Record Version | 5597
252 | Configure PC ASAP | Check for available SmartSEMService Packs and install them. (N/A if none) | 5598
252 | Configure PC ASAP | Check for available SmartSEM hotfixes and install them. (N/A if none) | 5599
252 | Configure PC ASAP | Installation of Zen Core EM | 5600
252 | Configure PC ASAP | Install CZ Support Tools. Ver | 5601
252 | Configure PC ASAP | Disable Automatic Updates | 5602
252 | Configure PC ASAP | Run ASAP Screening Program | 5603
252 | Configure PC ASAP | System Specific Datazone Loaded | 5604
252 | Configure PC ASAP | In SEM status set 'Mains Sync' to off | 5605
1445 | DC Tablefiller Checks | Mu metal shield is fitted | 5606
1445 | DC Tablefiller Checks | Table filler Finalimgs directory images are acceptable | 5607
1445 | DC Tablefiller Checks | All entries - Stigmation <20% | 5608
1445 | DC Tablefiller Checks | Store table editor stigmation Image | 5609
1445 | DC Tablefiller Checks | All entries - Align1 <60% | 5610
1445 | DC Tablefiller Checks | Store table editor Align1 Image | 5611
1445 | DC Tablefiller Checks | All entries - Gun <60% | 5612
1445 | DC Tablefiller Checks | Store table editor Gun Image | 5613
1445 | DC Tablefiller Checks | Beamshift check - All EHT enties at 2mm WD 0% DOF - POI doesn't deviate more than 10um | 5614
1445 | DC Tablefiller Checks | Beamshift check - All EHT enties at 2mm WD 100% DOF - POI doesn't deviate more than 10um | 5615
1445 | DC Tablefiller Checks | Beamshift check - All EHT enties at 2mm WD overview mode - POI doesn't deviate more than 10um | 5616
1445 | DC Tablefiller Checks | Beamshift check - All EHT enties at 5mm WD 0% DOF - POI doesn't deviate more than 10um | 5617
1445 | DC Tablefiller Checks | Beamshift check - All EHT enties at 5mm WD 100% DOF - POI doesn't deviate more than 10um | 5618
1445 | DC Tablefiller Checks | Beamshift check - All EHT enties at 5mm WD overview mode - POI doesn't deviate more than 10um | 5619
1445 | DC Tablefiller Checks | User align wobble <10% on all EHT entries - 2mm WD - 0% DOF | 5620
1445 | DC Tablefiller Checks | User align wobble <10% on all EHT entries - 2mm WD - 100% DOF | 5621
1445 | DC Tablefiller Checks | User align wobble <10% on all EHT entries - 2mm WD - Optimum mode | 5622
1445 | DC Tablefiller Checks | User align wobble <10% on all EHT entries - 5mm WD - 0% DOF | 5623
1445 | DC Tablefiller Checks | User align wobble <10% on all EHT entries - 5mm WD - 100% DOF | 5624
1445 | DC Tablefiller Checks | User align wobble <10% on all EHT entries - 5mm WD - Optium mode | 5625
482 | Run Tablefiller | Mu metal shield is fitted | 5626
482 | Run Tablefiller | DC Tablefiller successfully completed | 5627
952 | Declaration of conformity (Gem 360) | Please print 'Declaration of Conformity' from URL link | 5628
3248 | Heater Cables And Box Configuration | Plate | 5629
3248 | Heater Cables And Box Configuration | Screws | 5630
3248 | Heater Cables And Box Configuration | Spacers | 5631
3248 | Heater Cables And Box Configuration | Record heater box SN | 5632
10463 | Record PC Details | Check PC Part Number | 5633
10463 | Record PC Details | Select PC | 5634
10463 | Record PC Details | PC Serial No | 5635
10463 | Record PC Details | Check Windows 11 Certification label is fitted | 5636
1399 | Initial Resolution Images (XB 750) | Actual for 1kV (4mm WD) = | 5637
1399 | Initial Resolution Images (XB 750) | Store 1KV (4mm WD) Resolution Images and Pate image | 5638
1399 | Initial Resolution Images (XB 750) | Actual for 1kV (2mm WD) = | 5639
1399 | Initial Resolution Images (XB 750) | Store 1KV (2mm WD) Resolution Images and Pate image | 5640
1399 | Initial Resolution Images (XB 750) | Actual for 15kV = | 5641
1399 | Initial Resolution Images (XB 750) | Store 15KV Resolution Images and Pate image | 5642
10929 | Calibrate Compucentric Stage Centre | Calibrate Stage Centre (Right Click on Stgae tab) | 5643
10929 | Calibrate Compucentric Stage Centre | Copy the Centre Pos X and Y values into the Adminitrator, Stage Post Init. Position. | 5644
10929 | Calibrate Compucentric Stage Centre | Tilt Centre Calibration (Located in Docking panel) | 5645
8699 | Cable Up ( Lab6 240V Mains) | System Mains Voltage | 5646
8699 | Cable Up ( Lab6 240V Mains) | IF LaB6 and 120V - Transformer Installed | 5647
8699 | Cable Up ( Lab6 240V Mains) | Cable Up | 5648
8699 | Cable Up ( Lab6 240V Mains) | Label Cables | 5649
1405 | Cable Up (Lab6 120V Mains) | System Mains Voltage | 5650
1405 | Cable Up (Lab6 120V Mains) | IF LaB6 and 120V - Transformer Installed | 5651
1405 | Cable Up (Lab6 120V Mains) | Cable Up | 5652
1405 | Cable Up (Lab6 120V Mains) | Label Cables | 5653
329 | Stigmation / Aperture Checks | Stig: 30μm, All EHT’s are ≤ 20% | 5654
329 | Stigmation / Aperture Checks | Stig: Off Axis Aper, All EHT’s are ≤ 30% | 5655
329 | Stigmation / Aperture Checks | Aper/Gun: 30μm, All EHT’s (Excluding 99v) are ≤ 60% | 5656
329 | Stigmation / Aperture Checks | Store table editor Stigmation Image | 5657
329 | Stigmation / Aperture Checks | Store table editor Gun Alignment Image | 5658
329 | Stigmation / Aperture Checks | Store table editor Aperture Alignment Image | 5659
330 | Stigmation / Aperture Checks | Stig: 20μm Aper, 10kV is ≤ 35% | 5660
330 | Stigmation / Aperture Checks | Stig: Off Axis Aper, 10kV is ≤ 45% | 5661
330 | Stigmation / Aperture Checks | Stig: 20μm Aper, 1kV is ≤ 20% | 5662
330 | Stigmation / Aperture Checks | Stig: Off Axis Aper, 1kV is ≤ 30% | 5663
330 | Stigmation / Aperture Checks | Aper/Gun: 20μm, All EHT’s is ≤ 60% | 5664
330 | Stigmation / Aperture Checks | Store table editor Stigmation Image | 5665
330 | Stigmation / Aperture Checks | Store table editor Gun Alignment Image | 5666
330 | Stigmation / Aperture Checks | Store table editor Aperture Alignment Image | 5667
631 | VP/EP Function | VP/EP Function Checked | 5668
631 | VP/EP Function | Check Beam Is On at Max VP/EP Pressure | 5669
705 | Raith Stigmation /Aperture Checks | Stigs: 30um Aperture, All EHT ≤20% | 5670
705 | Raith Stigmation /Aperture Checks | Stigs: Off Axis Apertures, All EHT ≤30% | 5671
705 | Raith Stigmation /Aperture Checks | Aper & Gun: 30um Aperture, All EHT ≤50% | 5672
705 | Raith Stigmation /Aperture Checks | User align wobble <10% on all entries at 2mm WD | 5673
705 | Raith Stigmation /Aperture Checks | User align wobble <10% on all entries at 5mm WD | 5674
705 | Raith Stigmation /Aperture Checks | Store table editor Stigmation Image | 5675
705 | Raith Stigmation /Aperture Checks | Store table editor Gun Alignment Image | 5676
705 | Raith Stigmation /Aperture Checks | Store table editor Aperture Alignment Image | 5677
330 | Stigmation / Aperture Checks | Stig: 20μm Aper, 10kV is ≤ 35% | 5678
330 | Stigmation / Aperture Checks | Stig: Off Axis Aper, 10kV is ≤ 45% | 5679
330 | Stigmation / Aperture Checks | Stig: 20μm Aper, 1kV is ≤ 20% | 5680
330 | Stigmation / Aperture Checks | Stig: Off Axis Aper, 1kV is ≤ 30% | 5681
330 | Stigmation / Aperture Checks | Aper/Gun: 20μm, All EHT’s is ≤ 60% | 5682
330 | Stigmation / Aperture Checks | User align wobble <10% on all entries at 2mm WD | 5683
330 | Stigmation / Aperture Checks | User align wobble <10% on all entries at 5mm WD | 5684
330 | Stigmation / Aperture Checks | Store table editor Stigmation Image | 5685
330 | Stigmation / Aperture Checks | Store table editor Gun Alignment Image | 5686
330 | Stigmation / Aperture Checks | Store table editor Aperture Alignment Image | 5687
329 | Stigmation / Aperture Checks | Stig: 30μm, All EHT’s are ≤ 20% | 5688
329 | Stigmation / Aperture Checks | Stig: Off Axis Aper, All EHT’s are ≤ 30% | 5689
329 | Stigmation / Aperture Checks | Aper/Gun: 30μm, All EHT’s (Excluding 99v) are ≤ 60% | 5690
329 | Stigmation / Aperture Checks | User align wobble <10% on all entries at 2mm WD | 5691
329 | Stigmation / Aperture Checks | User align wobble <10% on all entries at 5mm WD | 5692
329 | Stigmation / Aperture Checks | Store table editor Stigmation Image | 5693
329 | Stigmation / Aperture Checks | Store table editor Gun Alignment Image | 5694
329 | Stigmation / Aperture Checks | Store table editor Aperture Alignment Image | 5695
1442 | Cable Up (Not Lab6) | System Mains Voltage | 5696
1442 | Cable Up (Not Lab6) | Cable Up | 5697
1442 | Cable Up (Not Lab6) | Label Cables | 5698
8699 | Cable Up ( Lab6 ) | System Mains Voltage | 5699
8699 | Cable Up ( Lab6 ) | IF LaB6 and 120V - Transformer Installed | 5700
8699 | Cable Up ( Lab6 ) | Cable Up | 5701
8699 | Cable Up ( Lab6 ) | Label Cables | 5702
1475 | ZEN core EM | ZEN Software Installed | 5703
1475 | ZEN core EM | Obtain a SE Image in ZEN | 5704
1475 | ZEN core EM | Obtain a Nav Cam Image and Check that a POI can be Navigated to. | 5705
3248 | Heater Cables And Box Configuration | Plate | 5706
3248 | Heater Cables And Box Configuration | Screws | 5707
3248 | Heater Cables And Box Configuration | Spacers | 5708
3248 | Heater Cables And Box Configuration | Record heater box SN: | 5709
8699 | Cable Up ( Lab6 ) | System Mains Voltage | 5710
8699 | Cable Up ( Lab6 ) | IF LaB6 and 120V - Transformer Installed | 5711
8699 | Cable Up ( Lab6 ) | Fit LAB6 Heater Control Box, Record SN: | 5712
8699 | Cable Up ( Lab6 ) | Cable Up | 5713
8699 | Cable Up ( Lab6 ) | Label Cables | 5714
6931 | Stage Drift Program | Drift result 1 | 5715
6931 | Stage Drift Program | Drift result 2 | 5716
6931 | Stage Drift Program | Drift result 3 | 5717
10464 | Setup Bios & Windows Setup | Check That AC Loss Policy Depend On, is Set To Power On | 5718
10464 | Setup Bios & Windows Setup | Set the Time and Date on the PC to current UK | 5719
10464 | Setup Bios & Windows Setup | Disable Windows Fast Start UP | 5720
10464 | Setup Bios & Windows Setup | Disable USB Power Save | 5721
10464 | Setup Bios & Windows Setup | Check PC Part Number | 5722
10464 | Setup Bios & Windows Setup | Check Windows 11 Certification label is fitted | 5723
10463 | Record PC Details | Select PC | 5724
10463 | Record PC Details | PC Serial No | 5725
10463 | Record PC Details | Check Windows 11 Certification label is fitted | 5726
10464 | Setup Bios & Windows Setup | Check That AC Loss Policy Depend On, is Set To Power On | 5727
10464 | Setup Bios & Windows Setup | Set the Time and Date on the PC to current UK | 5728
10464 | Setup Bios & Windows Setup | Disable Windows Fast Start UP | 5729
10464 | Setup Bios & Windows Setup | Disable USB Power Save | 5730
10464 | Setup Bios & More Windows Setup | Check That AC Loss Policy Depend On, is Set To Power On | 5731
10464 | Setup Bios & More Windows Setup | Set the Time and Date on the PC to current UK | 5732
10464 | Setup Bios & More Windows Setup | Disable Windows Fast Start UP | 5733
10464 | Setup Bios & More Windows Setup | Disable USB Power Save | 5734
10656 | TTL Functional Check | Vp Functional Check | 5735
10656 | TTL Functional Check | Vac Status = Ready and EHT Vac Ready = Yes, are always present at all VP values. (i.e. The beam remains on/can be switched on). | 5736
10652 | SmartEDX | Type of Detector | 5737
10652 | SmartEDX | Detector Serial number | 5738
10652 | SmartEDX | DDP Module Serial Number | 5739
10652 | SmartEDX | Apply Krytox Grease to the Adaptor Plate Inner O Ring. | 5740
10652 | SmartEDX | Fit Flange & Mount Detector | 5741
10652 | SmartEDX | Install EDS PCB's Into PC | 5742
10652 | SmartEDX | Installation of DPP Module & Cables | 5743
10652 | SmartEDX | Configure PC Network & Test | 5744
10652 | SmartEDX | Disable Windows10 Fast Startup | 5745
10652 | SmartEDX | Install ApexKey.edx Licences | 5746
10652 | SmartEDX | Install Genesis.edx Licences (if Applicable) | 5747
10652 | SmartEDX | Fit Dongle | 5748
10652 | SmartEDX | APEX Software Version | 5749
10652 | SmartEDX | Install APEX Software | 5750
10652 | SmartEDX | Install Remoting | 5751
10652 | SmartEDX | Setup RConfigure | 5752
10652 | SmartEDX | Run API Test | 5753
10652 | SmartEDX | Configure Imaging-Run SGShell | 5754
10652 | SmartEDX | Calibrate Detector-Run eDPP Shell | 5755
10652 | SmartEDX | Save Calibration Report | 5756
10652 | SmartEDX | Detector Resolution (measured using Cu) ≈ 129 eV | 5757
10652 | SmartEDX | Store Calibration Report | 5758
10652 | SmartEDX | Save Calibration on Exit | 5759
10652 | SmartEDX | Run APEX | 5760
10652 | SmartEDX | Check Spectrum | 5761
10652 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected | 5762
10652 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected | 5763
10652 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected | 5764
10652 | SmartEDX | Check Mapping | 5765
10652 | SmartEDX | Check Line Scanning | 5766
10652 | SmartEDX | Save Aquired Spectrum, Mappimg & Line Scanning to the Image Folder | 5767
10652 | SmartEDX | Store Aquired Spectrum (STD, Quick, HD), Mappimg & Line Scanning Images | 5768
10652 | SmartEDX | Save Calibration Data to the folder C: Support/Data/ Log Files | 5769
10652 | SmartEDX | Store Calibration Data File | 5770
10652 | SmartEDX | Save Calibration Data to the SmartSEM Image Folder | 5771
3314 | SmartEDX | Type of Detector | 5772
3314 | SmartEDX | Detector Serial number | 5773
3314 | SmartEDX | DDP Module Serial Number | 5774
3314 | SmartEDX | Apply Krytox Grease to the Adaptor Plate Inner O Ring. | 5775
3314 | SmartEDX | Fit Flange & Mount Detector | 5776
3314 | SmartEDX | Install EDS PCBs into PC | 5777
3314 | SmartEDX | Installation of DPP Module & Cables | 5778
3314 | SmartEDX | Configure PC Network & Test | 5779
3314 | SmartEDX | Disable Windows10 Fast Startup | 5780
3314 | SmartEDX | Install ApexKey.edx Licences | 5781
3314 | SmartEDX | Install Genesis.edx Licences (if Applicable) | 5782
3314 | SmartEDX | Fit Dongle | 5783
3314 | SmartEDX | APEX Software Version | 5784
3314 | SmartEDX | Install APEX Software | 5785
3314 | SmartEDX | Install Remoting | 5786
3314 | SmartEDX | Setup RConfigure | 5787
3314 | SmartEDX | Run API Test | 5788
3314 | SmartEDX | Configure imaging – Run SGShell | 5789
3314 | SmartEDX | Calibrate Detector – Run eDPPShell | 5790
3314 | SmartEDX | Save Calibration Report | 5791
3314 | SmartEDX | Detector Resolution - ≥ 129 eV | 5792
3314 | SmartEDX | Save calibration data on exit | 5793
3314 | SmartEDX | Run APEX | 5794
3314 | SmartEDX | Check Spectrum | 5795
3314 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected | 5796
3314 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected | 5797
3314 | SmartEDX | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected | 5798
3314 | SmartEDX | Check Mapping | 5799
3314 | SmartEDX | Check Line Scanning | 5800
3314 | SmartEDX | Save Aquired Spectrum, Mapping and Line Scanning Images to Image Folder | 5801
3314 | SmartEDX | Save Calibration Data to the Licence Folderon the SmartSem USB stick. | 5802
3314 | SmartEDX | Save Calibration Data to the Image Folder | 5803
10827 | Fit Stage Cabling | Stage Cable | 5804
10827 | Fit Stage Cabling | Specimen Touch Cable | 5805
1250 | Stage Repeatability Macro - Hittech Stage Only | Download Macro folder | 5806
1250 | Stage Repeatability Macro - Hittech Stage Only | Run Kompass 3.08.exe | 5807
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 1 (MAX 5um) | 5808
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 2 (MAX 5um) | 5809
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 3 (MAX 5um) | 5810
1250 | Stage Repeatability Macro - Hittech Stage Only | Image of Macro Results | 5811
1250 | Stage Repeatability Macro - Hittech Stage Only | Download Macro folder | 5812
1250 | Stage Repeatability Macro - Hittech Stage Only | Change the Directory to C:/Program Data/Carl Zeiss/SmartSEM/Images/Kompass. The Kompass sub-directory will need to be created | 5813
1250 | Stage Repeatability Macro - Hittech Stage Only | Run Kompass 3.08.exe | 5814
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 1 (MAX 5um) | 5815
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 2 (MAX 5um) | 5816
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 3 (MAX 5um) | 5817
1250 | Stage Repeatability Macro - Hittech Stage Only | Image of Macro Results | 5818
1250 | Stage Repeatability Macro - Hittech Stage Only | Download Macro folder | 5819
1250 | Stage Repeatability Macro - Hittech Stage Only | Obtain Image of Tin on Carbon Sample in a Outer Position @ 4K Mag, 20KV, 7mmWD | 5820
1250 | Stage Repeatability Macro - Hittech Stage Only | Change the Directory to C:/Program Data/Carl Zeiss/SmartSEM/Images/Kompass. The Kompass sub-directory will need to be created | 5821
1250 | Stage Repeatability Macro - Hittech Stage Only | Enter Stage SN | 5822
1250 | Stage Repeatability Macro - Hittech Stage Only | Run Kompass 3.08.exe | 5823
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 1 (MAX 5um) | 5824
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 2 (MAX 5um) | 5825
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 3 (MAX 5um) | 5826
1250 | Stage Repeatability Macro - Hittech Stage Only | Image of Macro Results | 5827
1250 | Stage Repeatability Macro - Hittech Stage Only | Download Macro folder | 5828
1250 | Stage Repeatability Macro - Hittech Stage Only | Obtain Image of Tin on Carbon Sample in a Outer Position @ 4K Mag, 20KV, 7mmWD | 5829
1250 | Stage Repeatability Macro - Hittech Stage Only | Change the Directory to C:/Program Data/Carl Zeiss/SmartSEM/Images/Kompass. The Kompass sub-directory will need to be created | 5830
1250 | Stage Repeatability Macro - Hittech Stage Only | Enter Stage SN | 5831
1250 | Stage Repeatability Macro - Hittech Stage Only | Execute Kompass 3.08.exe | 5832
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 1 (MAX 5um) | 5833
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 2 (MAX 5um) | 5834
1250 | Stage Repeatability Macro - Hittech Stage Only | Macro Result 3 (MAX 5um) | 5835
1250 | Stage Repeatability Macro - Hittech Stage Only | Image of Macro Results | 5836
10506 | VP Functional Checks | Vp Functional Check - Switch Off the TestTool Quite Mode Before Proceeding. | 5837
10506 | VP Functional Checks | Vac Status = Ready and EHT Vac Ready = Yes, are always present at all VP values. (i.e. The beam remains on/can be switched on). | 5838
10506 | VP Functional Checks | Fit Label to the Black Plastic Aperture Box | 5839
1454 | Raman | Test Valves Using Raman Test Interlock Box | 5840
1454 | Raman | Install SemRamanStage Software | 5841
1454 | Raman | Run SEM Raman Correlation | 5842
1454 | Raman | Save Calibration File as Raman | 5843
1454 | Raman | second chamberscope fitted | 5844
1454 | Raman | Run Smart Stage Mapping | 5845
252 | Configure PC ASAP | PC S/N | 5846
252 | Configure PC ASAP | Hard disk Image (Version) | 5847
252 | Configure PC ASAP | Record Operating System | 5848
252 | Configure PC ASAP | Check the Entitlement certificate to determine which version of Smartsem to install | 5849
252 | Configure PC ASAP | Install Smart SEM (See WI 9693) Record Version | 5850
252 | Configure PC ASAP | Check for available SmartSEMService Packs and install them. (N/A if none) | 5851
252 | Configure PC ASAP | Check for available SmartSEM hotfixes and install them. (N/A if none) | 5852
252 | Configure PC ASAP | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 5853
252 | Configure PC ASAP | Install CZ Support Tools. Ver | 5854
252 | Configure PC ASAP | Disable Automatic Updates | 5855
252 | Configure PC ASAP | Run ASAP Screening Program | 5856
252 | Configure PC ASAP | System Specific Datazone Loaded | 5857
252 | Configure PC ASAP | In SEM status set 'Mains Sync' to off | 5858
3242 | PC Installation | Select PC | 5859
3242 | PC Installation | PC S/No. | 5860
3242 | PC Installation | Operating System | 5861
3242 | PC Installation | Change Desktop Settings | 5862
3242 | PC Installation | Install SmartSEM and record Version | 5863
3242 | PC Installation | Check for Service Packs and install them | 5864
3242 | PC Installation | Upgrade ASAP Firmware (check Service Pack for latest version) and API | 5865
3242 | PC Installation | Copy the SmartSEM sotware into a Folder in C:Support and Name it SmartSEM X.x (i.e 8.2) Software | 5866
3242 | PC Installation | Installation of Zen Core EM if Hasap is fitted - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 5867
3242 | PC Installation | Check That AC Loss Policy Depend On, is Set To Power On | 5868
3242 | PC Installation | Set the Time and Date on the PC to current UK | 5869
3242 | PC Installation | Disable Windows Fast Start UP | 5870
3242 | PC Installation | Disable USB Power Save | 5871
3242 | PC Installation | Check PC Part Number | 5872
3242 | PC Installation | Check Windows 11 Certification label is fitted | 5873
3242 | PC Installation | Attach Cable Markers to USB Cables | 5874
347 | Customer Images | 1kV Inlens | 5875
347 | Customer Images | Store 1kV Inlens Customer Image | 5876
347 | Customer Images | 15kV Inlens | 5877
347 | Customer Images | Store 15kV Inlens Customer Image | 5878
11341 | Customer Images (ESB) | 1kV ESB | 5879
11341 | Customer Images (ESB) | Store 1kV ESB Customer Image | 5880
346 | Resolution | 1kV = | 5881
346 | Resolution | Store 1KV Resolution Images and Pate image | 5882
346 | Resolution | 15kV = | 5883
346 | Resolution | Store 1KV Resolution Images and Pate image | 5884
10473 | Administrator Factory - Column/Stage | Select Stage Type | 5885
10473 | Administrator Factory - Column/Stage | Select Stage Type | 5886
10473 | Administrator Factory - Column/Stage | Tick XYZTR to be all enabled | 5887
10473 | Administrator Factory - Column/Stage | Tick Rotate Limit | 5888
10473 | Administrator Factory - Column/Stage | Tick Joystick | 5889
10473 | Administrator Factory - Column/Stage | Tick 'Post init Posn Valid | 5890
10473 | Administrator Factory - Column/Stage | Do Not Tick T init end at zero | 5891
10473 | Administrator Factory - Column/Stage | Select Stage Type | 5892
10473 | Administrator Factory - Column/Stage | Tick XYZTR to be all enabled | 5893
10473 | Administrator Factory - Column/Stage | Tick Rotate Limit | 5894
10473 | Administrator Factory - Column/Stage | Tick Joystick | 5895
10473 | Administrator Factory - Column/Stage | Tick 'Post init Posn Valid | 5896
10473 | Administrator Factory - Column/Stage | Do Not Tick T init end at zero | 5897
10473 | Administrator Factory - Column/Stage | Select Column Pumping If Presenr | 5898
10473 | Administrator Factory - Column/Stage | Select Stage Type | 5899
10473 | Administrator Factory - Column/Stage | Tick XYZTR to be all enabled | 5900
10473 | Administrator Factory - Column/Stage | Tick Rotate Limit | 5901
10473 | Administrator Factory - Column/Stage | Tick Joystick | 5902
10473 | Administrator Factory - Column/Stage | Tick 'Post init Posn Valid | 5903
10473 | Administrator Factory - Column/Stage | Do Not Tick T init end at zero | 5904
10473 | Administrator Factory - Column/Stage | Select Column Pumping If Present | 5905
10473 | Administrator Factory - Column/Stage | Stage Chamberscope - Set Joystick TV Angle to 270deg | 5906
3257 | SmartSEM Configuration | In SEM status set 'Mains Sync' to off | 5907
3257 | SmartSEM Configuration | Stage Chamberscope - Set Joystick TV Angle to 270deg | 5908
91 | Section Type | Housing Type | 5909
91 | Section Type | Inspect insulation disc using white light for cracks | 5910
256 | Run System checker and upgrade board FW to latest versions. | Ensure the Stage Controller version is compatible with the Stage Axes version (See compatible verions in drop down list) | 5911
256 | Run System checker and upgrade board FW to latest versions. | Store System checker full report | 5912
506 | Additional Raith Aperture Measurments | Measure aperture size, 7.5um (+/-2um) | 5913
506 | Additional Raith Aperture Measurments | Multi_TS_7.5um - Passed inspection criteria, Image Saved | 5914
506 | Additional Raith Aperture Measurments | Measure aperture size, 10um (+/-2um) | 5915
506 | Additional Raith Aperture Measurments | Multi_TS_10um - Passed inspection criteria, Image Saved | 5916
506 | Additional Raith Aperture Measurments | Measure aperture size, 15um (+/-2um) | 5917
506 | Additional Raith Aperture Measurments | Multi_TS_15um - Passed inspection criteria, Image Saved | 5918
506 | Additional Raith Aperture Measurments | Measure aperture size, 20um (+/-2um) | 5919
506 | Additional Raith Aperture Measurments | Multi_TS_20um - Passed inspection criteria, Image Saved | 5920
506 | Additional Raith Aperture Measurments | Measure aperture size, 60um (+/-2um) | 5921
506 | Additional Raith Aperture Measurments | Multi_TS_60um - Passed inspection criteria, Image Saved | 5922
10489 | Check Function of SE Detector | Scintillator | 5923
10489 | Check Function of SE Detector | Collector Volts | 5924
10489 | Check Function of SE Detector | Check that the Signal Level in Line Mode is at Black Level with Brightness=50% and Contrast =0% | 5925
10489 | Check Function of SE Detector | Check that the Signal Level in Line Mode is at White Level with Brigfhtness =62% and Contrast =0% | 5926
1278 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 5927
1278 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 5928
1278 | ZEN Core EM | Install Customer's ZEN Licence | 5929
1278 | ZEN Core EM | Setup SmartSEM Toggle Button | 5930
1278 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 5931
1278 | ZEN Core EM | Mains Sync Set To OFF | 5932
1278 | ZEN Core EM | Select Your ZEN core EM Runtime Profile | 5933
1278 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 5934
1278 | ZEN Core EM | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems. | 5935
1278 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 5936
1278 | ZEN Core EM | Configue Airlock Cam and Setup Exchange Position $Exchange (use Upper case E in ZEN) | 5937
1278 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 5938
1278 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 5939
1278 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 5940
1278 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 5941
1278 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 5942
1278 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 5943
1279 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 5944
1279 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 5945
1279 | ZEN Core EM | Install Customer's ZEN Licence | 5946
1279 | ZEN Core EM | Setup SmartSEM Toggle Button | 5947
1279 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 5948
1279 | ZEN Core EM | Mains Sync Set To OFF | 5949
1279 | ZEN Core EM | Choose your ZEN core EM runtime profile | 5950
1279 | ZEN Core EM | Create User Supervisor Volutome (only applicable for Volutome systems) | 5951
1279 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 5952
1279 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 5953
1279 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 5954
1279 | ZEN Core EM | Configue $NavCam (Capital N and C) | 5955
1279 | ZEN Core EM | Configue $Exchange (capital E) | 5956
1279 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 5957
1279 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 5958
1279 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 5959
1279 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 5960
1279 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 5961
1279 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 5962
1280 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 5963
1280 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 5964
1280 | ZEN Core EM | Install Customer's ZEN Licence | 5965
1280 | ZEN Core EM | Setup SmartSEM Toggle Button | 5966
1280 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 5967
1280 | ZEN Core EM | Mains Sync Set To OFF | 5968
1280 | ZEN Core EM | Choose your ZEN core EM runtime profile | 5969
1280 | ZEN Core EM | Create User Supervisor Volutome (only applicable for Volutome systems) | 5970
1280 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 5971
1280 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 5972
1280 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 5973
1280 | ZEN Core EM | Configue $NavCam (Capital N and C) | 5974
1280 | ZEN Core EM | Configue $Exchange (capital E) | 5975
1280 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 5976
1280 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 5977
1280 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 5978
1280 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 5979
1280 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 5980
1280 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 5981
1280 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 5982
1280 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 5983
1280 | ZEN Core EM | Install Customer's ZEN Licence | 5984
1280 | ZEN Core EM | Setup SmartSEM Toggle Button | 5985
1280 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 5986
1280 | ZEN Core EM | Mains Sync Set To OFF | 5987
1280 | ZEN Core EM | Choose your ZEN core EM runtime profile | 5988
1280 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 5989
1280 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 5990
1280 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 5991
1280 | ZEN Core EM | Configue $NavCam (Capital N and C) | 5992
1280 | ZEN Core EM | Configue $Exchange (capital E) | 5993
1280 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 5994
1280 | ZEN Core EM | Enable Project Folder Mode | 5995
1280 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 5996
1280 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 5997
1280 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 5998
1280 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 5999
1280 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6000
1278 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6001
1278 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6002
1278 | ZEN Core EM | Install Customer's ZEN Licence | 6003
1278 | ZEN Core EM | Setup SmartSEM Toggle Button | 6004
1278 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6005
1278 | ZEN Core EM | Mains Sync Set To OFF | 6006
1278 | ZEN Core EM | Select Your ZEN core EM Runtime Profile | 6007
1278 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6008
1278 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 6009
1278 | ZEN Core EM | Configue Airlock Cam and Setup Exchange Position $Exchange (use Upper case E in ZEN) | 6010
1278 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6011
1278 | ZEN Core EM | Enable Project Folder Mode | 6012
1278 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6013
1278 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6014
1278 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6015
1278 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6016
1278 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6017
94 | Fit HE Inlens Detector | Serial: | 6018
94 | Fit HE Inlens Detector | Remove metal protective cover from detector and inspect detector head | 6019
1645 | Fit ESB Detector | Serial: | 6020
1645 | Fit ESB Detector | Remove metal protective cover from detector and inspect detector head | 6021
11336 | Fit InLens Duo Detector | Serial: | 6022
11336 | Fit InLens Duo Detector | Remove metal protective cover from detector and inspect detector head | 6023
3295 | Fine positioning Inlens Detector | Serial: | 6024
3295 | Fine positioning Inlens Detector | Remove metal protective cover from detector and inspect detector head | 6025
168 | 250V Isolation Pre-fitment Checks | Driver A – Liner (>1000MΩ) | 6026
168 | 250V Isolation Pre-fitment Checks | Driver B – Liner (>1000MΩ) | 6027
168 | 250V Isolation Pre-fitment Checks | Driver A – Driver B (>1000MΩ) | 6028
11190 | 250V Isolation Pre-fitment Checks (Raith) | Driver A – Liner (>1000MΩ) | 6029
171 | 250V Isolation Post-fitment Checks | Driver A – Liner (>1000MΩ) | 6030
171 | 250V Isolation Post-fitment Checks | Driver B – Liner (>1000MΩ) | 6031
171 | 250V Isolation Post-fitment Checks | Driver A – Driver B (>1000MΩ) | 6032
171 | 250V Isolation Post-fitment Checks | Driver A – GROUND (>1000MΩ) | 6033
171 | 250V Isolation Post-fitment Checks | Driver B – GROUND (>1000MΩ) | 6034
171 | 250V Isolation Post-fitment Checks | Liner – GROUND (>1000MΩ) | 6035
172 | EBBU Continuity Check | Driver A – Driver A In | 6036
172 | EBBU Continuity Check | Driver B – Driver B In | 6037
172 | EBBU Continuity Check | Anode – Liner In | 6038
1459 | Set default Aligning System Offset Angle (Gemini 1) | Align offset angle set to -140 degrees. | 6039
1459 | Set default Aligning System Offset Angle (Gemini 1) | Align offset angle set to -140 degrees. | 6040
1459 | Set default Aligning System Offset Angle (Gemini 360 /560) | Align offset angle set to -140 degrees. | 6041
1459 | Set default Aligning System Offset Angle (Gemini 360 /560) | Align offset angle set to -140 degrees. | 6042
1279 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6043
1279 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6044
1279 | ZEN Core EM | Install Customer's ZEN Licence | 6045
1279 | ZEN Core EM | Setup SmartSEM Toggle Button | 6046
1279 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6047
1279 | ZEN Core EM | Mains Sync Set To OFF | 6048
1279 | ZEN Core EM | Choose your ZEN core EM runtime profile | 6049
1279 | ZEN Core EM | Create User Supervisor Volutome (only applicable for Volutome systems) | 6050
1279 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6051
1279 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 6052
1279 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 6053
1279 | ZEN Core EM | Configue $NavCam (Capital N and C) | 6054
1279 | ZEN Core EM | Configue $Exchange (capital E) | 6055
1279 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6056
1279 | ZEN Core EM | Enable Project Folder Mode | 6057
1279 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6058
1279 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6059
1279 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6060
1279 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6061
1279 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6062
1459 | Set default Aligning System Offset Angle (Gemini 360 /560) | Align offset angle set to -51 degrees. | 6063
1459 | Set default Aligning System Offset Angle (Gemini 360 /560) | Align offset angle set to -51 degrees. | 6064
1428 | Set default Aligning System Offset Angle (Gemini 360 OEM7) | Align offset angle set to -140 degrees. | 6065
1460 | Set default Aligning System Offset Angle (Gemini 360 /560) | Align offset angle set to -51 degrees. | 6066
1460 | Set default Aligning System Offset Angle (Crossbeam 350) | Align offset angle set to -51 degrees. | 6067
1460 | Set default Aligning System Offset Angle (Crossbeam 350) | Align offset angle set to -51 degrees. | 6068
1460 | Set default Aligning System Offset Angle (Crossbeam 350) | Align offset angle set to -21 degrees. | 6069
1460 | Set default Aligning System Offset Angle (Crossbeam 350) | Align offset angle set to -21 degrees. | 6070
10957 | Column Pumping Jig | Clean and Prepare the Jig and Fit the O ring | 6071
10968 | C3 Outer Cone | Fit C3 Cone into Jig (Vac seal facing you) | 6072
10969 | C3 Inner Cone | Fit the Sealing Plate onto the Inner C3 Cone | 6073
10969 | C3 Inner Cone | Fit Handles and Lower into the Outer Cone | 6074
10969 | C3 Inner Cone | Secure the Outer C3 Cone to the Jig | 6075
10969 | C3 Inner Cone | Secure the Inner C£ Cone to the Outer Cone | 6076
10969 | C3 Inner Cone | Fit Scan Coils | 6077
10969 | C3 Inner Cone | Fit FoV Coils | 6078
10969 | C3 Inner Cone | Fit O Ring to Top of Scan Coil | 6079
10969 | C3 Inner Cone | Perform Vacuum Leak Check | 6080
10970 | Column Pumping Body | Fit Column Pumping Body onto the C3 Assembly | 6081
10970 | Column Pumping Body | Fit Test Square Blanking Plate | 6082
10970 | Column Pumping Body | Fit Mu-metal Shield | 6083
10970 | Column Pumping Body | Fit Aperture Changer | 6084
10970 | Column Pumping Body | Fit Circular Blank (PCM) | 6085
10970 | Column Pumping Body | Fit Rectangular Blank (CIV) | 6086
10974 | C1/C2 Body | Fit C1/C2 Body | 6087
10974 | C1/C2 Body | Perform Vacuum Leak Check | 6088
10975 | C3 Continuity Test Box | Scan Coil - LED is On | 6089
10975 | C3 Continuity Test Box | Press Button and LED Goes Off | 6090
10976 | Gun Head | Fit Firing Unit into Gun Head | 6091
10976 | Gun Head | Fit O Ring into Top of C!/C2 Body | 6092
10976 | Gun Head | Fit Gun Head to C1/C2 Body | 6093
10977 | C1/C2 Body External Parts | Fit Gun Safety Switch | 6094
10977 | C1/C2 Body External Parts | Fit Gun Align Cable To Bracket | 6095
10977 | C1/C2 Body External Parts | Fit Assembled Bracket To C1/C2 Body | 6096
10978 | C1/C2 Body Internal Parts | Fit Pole Pieces | 6097
10978 | C1/C2 Body Internal Parts | Fit Bottom Plate | 6098
10978 | C1/C2 Body Internal Parts | Fit Gun Align Coils | 6099
10978 | C1/C2 Body Internal Parts | Fit Anode | 6100
10978 | C1/C2 Body Internal Parts | Check for Cleanliness Using UV Torch | 6101
10980 | Final Vacuum Check | Complete Column Vauum Checked | 6102
10980 | Final Vacuum Check | Vacuum Isolation Valve Closed | 6103
11246 | Production and Supplier Paperwork | All Paperwork Sorted and Stapled Together | 6104
10942 | Circular Blank | Clean and Inspect O Ring | 6105
10942 | Circular Blank | Fit O Ring into Blank | 6106
10979 | System Type | Select System Type | 6107
10979 | System Type | Column Type | 6108
10943 | Rectangular Blank | Clean and Inspect O Ring | 6109
10943 | Rectangular Blank | Fit O Ring into Blank | 6110
10944 | Sealing Plate | Clean and Inspect O Rings | 6111
10944 | Sealing Plate | Fit O Rings into Plate | 6112
10945 | C3 Cone Vac Seal | Clean and Inspect O Ring | 6113
10945 | C3 Cone Vac Seal | Fit O Ring into Vac Seal | 6114
10945 | C3 Cone Vac Seal | Insert Vac Seal into C3 Cone | 6115
10946 | Lens End Cap Assembly | Clean and Inspect O Ring | 6116
10946 | Lens End Cap Assembly | Fit O Ring it to the Lens Cap | 6117
10946 | Lens End Cap Assembly | Mount the End Cap Onto the C3 Cone | 6118
10946 | Lens End Cap Assembly | Tighten the Screws | 6119
10946 | Lens End Cap Assembly | End Cap and C3 Cone Gap Checked | 6120
10947 | Gun Align Coil | Liner Tube Inspected | 6121
10947 | Gun Align Coil | Inspect O Ring Sealing Surfaces | 6122
10947 | Gun Align Coil | Liner Tube Cleaned | 6123
10947 | Gun Align Coil | Clean and Inspect O Rings | 6124
10947 | Gun Align Coil | Fit O Rings onto Coils | 6125
479 | Scan Coil Auto Test | Auto Inductance Test Passed | 6126
479 | Scan Coil Auto Test | Auto Flux Test Passed | 6127
10948 | Scan Coils | Liner Tube Inspected | 6128
10948 | Scan Coils | Inspect O Ring Sealing Surfaces | 6129
10948 | Scan Coils | Liner Tube Cleaned | 6130
10948 | Scan Coils | Clean and Inspect O Rings | 6131
10948 | Scan Coils | Fit O Rings onto Coils | 6132
10949 | Pumping Body | Clean and Inspect O Ring | 6133
10949 | Pumping Body | Fit O Ring to the Underside of the Body | 6134
10949 | Pumping Body | Fit Pumping Guide | 6135
10949 | Pumping Body | Check Pumping Guide Alignment | 6136
10949 | Pumping Body | Fit M3x10 Countersunk Screws | 6137
10949 | Pumping Body | Fit Chamber Pumping Shield | 6138
10949 | Pumping Body | Clean and Inspect O Ring | 6139
10949 | Pumping Body | Fit O Ring to Upperside of the Body | 6140
11198 | Fit Ski Jump and CIV | Fit Ski Jump Onto Plate | 6141
11198 | Fit Ski Jump and CIV | Using Jig Align Ski Slope Bore With Plate Bore. | 6142
11198 | Fit Ski Jump and CIV | Fit Bottom Plate Using 4x M4X8 CSK Screws | 6143
11198 | Fit Ski Jump and CIV | Fit Metal Shield | 6144
11198 | Fit Ski Jump and CIV | Insert Ski Jump Plate and Secure | 6145
11198 | Fit Ski Jump and CIV | Fit O Ring | 6146
11198 | Fit Ski Jump and CIV | Clean and Check CIV Sealing O Ring | 6147
11198 | Fit Ski Jump and CIV | Clean, Check and Fit CIV Body O Ring | 6148
11198 | Fit Ski Jump and CIV | Check CIV Height Using Letterbox Jig | 6149
11198 | Fit Ski Jump and CIV | Fit CIV Onto Pumping Body | 6150
11198 | Fit Ski Jump and CIV | Check CIV Can Open and Close Without Touching Insert Plate | 6151
11199 | Ion Pump | Fit Heater and Cover Plate | 6152
11199 | Ion Pump | Fit the Large Cover Plate and Temperature Label | 6153
11199 | Ion Pump | Remove Ion Pump Transit Blank | 6154
11199 | Ion Pump | Fit Copper Seal | 6155
11199 | Ion Pump | Fit and Secure Baffle | 6156
11199 | Ion Pump | Align and Secure T Pipe | 6157
11199 | Ion Pump | Fit Ion Pump Assy. Onto Column | 6158
10950 | Aperture Changer | Inspect Aperture Blade | 6159
10950 | Aperture Changer | Install the Apertures | 6160
10951 | Firing Unit | Disassemble the Firing Unit | 6161
10951 | Firing Unit | Clean Listed Parts in Ultrasonic Tank | 6162
10951 | Firing Unit | Reassemble Firing Unit and Fit Filament | 6163
10951 | Firing Unit | Check Filament Height and Centre | 6164
10952 | Bottom Plate and Aperture Clamp (W) | Inspect Bores of Plate and Clamp | 6165
10952 | Bottom Plate and Aperture Clamp (W) | Fit 600um Aperture | 6166
10952 | Bottom Plate and Aperture Clamp (W) | Clean and Inspect O Ring and Fit to the Underside of Plate | 6167
10954 | Bottom Plate and Aperture Clamp (LaB6) | Inspect Bores of Plate and Clamp | 6168
10954 | Bottom Plate and Aperture Clamp (LaB6) | Fit 300um Aperture | 6169
10954 | Bottom Plate and Aperture Clamp (LaB6) | Clean and Inspect O Ring and Fit to the Underside of Plate | 6170
10953 | C1/C2 Pole Pieces | Fit an Aperture Insert Into Each Pole Piece | 6171
10953 | C1/C2 Pole Pieces | Check Aperture Inserts are Securely Fitted | 6172
10464 | Setup Bios & More Windows Setup | Check That AC Loss Policy Depend On, is Set To Power On | 6173
10464 | Setup Bios & More Windows Setup | Set the Time and Date on the PC to current UK | 6174
10464 | Setup Bios & More Windows Setup | Disable Windows Fast Start UP | 6175
10464 | Setup Bios & More Windows Setup | Set Dashclientservice to Manual | 6176
10464 | Setup Bios & More Windows Setup | Disable USB Power Save | 6177
3242 | PC Installation | Select PC | 6178
3242 | PC Installation | PC S/No. | 6179
3242 | PC Installation | Operating System | 6180
3242 | PC Installation | Change Desktop Settings | 6181
3242 | PC Installation | Install SmartSEM and record Version | 6182
3242 | PC Installation | Check for Service Packs and install them | 6183
3242 | PC Installation | Upgrade ASAP Firmware (check Service Pack for latest version) and API | 6184
3242 | PC Installation | Copy the SmartSEM sotware into a Folder in C:Support and Name it SmartSEM X.x (i.e 8.2) Software | 6185
3242 | PC Installation | Installation of Zen Core EM if Hasap is fitted - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 6186
3242 | PC Installation | Check That AC Loss Policy Depend On, is Set To Power On | 6187
3242 | PC Installation | Set the Time and Date on the PC to current UK | 6188
3242 | PC Installation | Disable Windows Fast Start UP | 6189
3242 | PC Installation | Set Dashclientservice to Manual | 6190
3242 | PC Installation | Disable USB Power Save | 6191
3242 | PC Installation | Check PC Part Number | 6192
3242 | PC Installation | Check Windows 11 Certification label is fitted | 6193
3242 | PC Installation | Attach Cable Markers to USB Cables | 6194
1278 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6195
1278 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6196
1278 | ZEN Core EM | Install Customer's ZEN Licence | 6197
1278 | ZEN Core EM | Setup SmartSEM Toggle Button | 6198
1278 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6199
1278 | ZEN Core EM | Mains Sync Set To OFF | 6200
1278 | ZEN Core EM | Select Your ZEN core EM Runtime Profile | 6201
1278 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6202
1278 | ZEN Core EM | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems. | 6203
1278 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 6204
1278 | ZEN Core EM | Configue Airlock Cam and Setup Exchange Position $Exchange (use Upper case E in ZEN) | 6205
1278 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6206
1278 | ZEN Core EM | Enable Project Folder Mode | 6207
1278 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6208
1278 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6209
1278 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6210
1278 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6211
1278 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6212
1280 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6213
1280 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6214
1280 | ZEN Core EM | Install Customer's ZEN Licence | 6215
1280 | ZEN Core EM | Setup SmartSEM Toggle Button | 6216
1280 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6217
1280 | ZEN Core EM | Mains Sync Set To OFF | 6218
1280 | ZEN Core EM | Choose your ZEN core EM runtime profile | 6219
1280 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6220
1280 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 6221
1280 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 6222
1280 | ZEN Core EM | Configue $NavCam (Capital N and C) | 6223
1280 | ZEN Core EM | Configue $Exchange (capital E) | 6224
1280 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6225
1280 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6226
1280 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6227
1280 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6228
1280 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6229
1280 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6230
1279 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6231
1279 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6232
1279 | ZEN Core EM | Install Customer's ZEN Licence | 6233
1279 | ZEN Core EM | Setup SmartSEM Toggle Button | 6234
1279 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6235
1279 | ZEN Core EM | Mains Sync Set To OFF | 6236
1279 | ZEN Core EM | Choose your ZEN core EM runtime profile | 6237
1279 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6238
1279 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 6239
1279 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 6240
1279 | ZEN Core EM | Configue $NavCam (Capital N and C) | 6241
1279 | ZEN Core EM | Configue $Exchange (capital E) | 6242
1279 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6243
1279 | ZEN Core EM | Enable Project Folder Mode | 6244
1279 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6245
1279 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6246
1279 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6247
1279 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6248
1279 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6249
3309 | Atlas 5 | Connect Silver Box and Cables | 6250
3309 | Atlas 5 | Download latest Version Of Software (see link in docs) | 6251
3309 | Atlas 5 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info | 6252
3309 | Atlas 5 | Request Licence File from Fibics | 6253
3309 | Atlas 5 | Install Atlas Software | 6254
3309 | Atlas 5 | Install Licence File | 6255
3309 | Atlas 5 | Calibrate SEM | 6256
3309 | Atlas 5 | Obtain Mosaic | 6257
3309 | Atlas 5 | Set ATLAS Rotation Offset | 6258
3309 | Atlas 5 | Save ATLAS Licence Document | 6259
10526 | Chamberscope | Number of Chamberscope's on Order | 6260
10526 | Chamberscope | Chamberscope 1 - Select Port | 6261
10526 | Chamberscope | Check Camera for focus | 6262
10526 | Chamberscope | Functional Check-Black/White + Colour Mode | 6263
10526 | Chamberscope | Chamberscope 2 - Select Port or Not Fitted | 6264
10526 | Chamberscope | Chamberscope 2 - Check Camera 2 for Focus | 6265
10526 | Chamberscope | Chamberscope 2 - Functional Check-Black/White + Colour Mode | 6266
10526 | Chamberscope | Chamberscope 2 - Add Icon to Toolbar | 6267
10526 | Chamberscope | Number of Chamberscope's on Order | 6268
10526 | Chamberscope | Chamberscope 1 - Select Port | 6269
10526 | Chamberscope | Check Camera for focus | 6270
10526 | Chamberscope | Functional Check-Black/White + Colour Mode | 6271
10526 | Chamberscope | Chamberscope 2 - Select Port or Not Fitted | 6272
10526 | Chamberscope | Chamberscope 2 - Check Camera 2 for Focus | 6273
10526 | Chamberscope | Chamberscope 2 - Functional Check-Black/White + Colour Mode | 6274
10526 | Chamberscope | Chamberscope 2 - Add Icon to Toolbar | 6275
330 | Stigmation / Aperture Checks | Stig: 20μm Aper, 10kV is ≤ 35% | 6276
330 | Stigmation / Aperture Checks | Stig: Off Axis Aper, 10kV is ≤ 45% | 6277
330 | Stigmation / Aperture Checks | Stig: 20μm Aper, 1kV is ≤ 20% | 6278
330 | Stigmation / Aperture Checks | Stig: Off Axis Aper, 1kV is ≤ 30% | 6279
330 | Stigmation / Aperture Checks | Aper/Gun: 20μm, All EHT’s is ≤ 60% | 6280
330 | Stigmation / Aperture Checks | User align wobble <10% on all entries at 2mm WD (1kV & 10kV) | 6281
330 | Stigmation / Aperture Checks | User align wobble <10% on all entries at 5mm WD (1kV & 10kV) | 6282
330 | Stigmation / Aperture Checks | Store table editor Stigmation Image | 6283
330 | Stigmation / Aperture Checks | Store table editor Gun Alignment Image | 6284
330 | Stigmation / Aperture Checks | Store table editor Aperture Alignment Image | 6285
1474 | ZEN core EM | ZEN Software Installed | 6286
1474 | ZEN core EM | Obtain a SE Image in ZEN | 6287
1474 | ZEN core EM | Obtain a Nav Cam Image and Check that a POI can be Navigated to. | 6288
1475 | ZEN core EM | ZEN Software Installed | 6289
1475 | ZEN core EM | Obtain a SE Image in ZEN | 6290
1475 | ZEN core EM | Obtain a Nav Cam Image and Check that a POI can be Navigated to. | 6291
10567 | Install Customer Licences | Delete the Production Test Licences | 6292
10567 | Install Customer Licences | Obtain the customer Entitlement Certificate | 6293
10567 | Install Customer Licences | Download into the C:Support/Data/Licence folder | 6294
10567 | Install Customer Licences | Install the Customer Licences | 6295
10567 | Install Customer Licences | Copy Contents of Image Folder to the Product Folder | 6296
10567 | Install Customer Licences | Copy EMServer.db to the Product Folder | 6297
10567 | Install Customer Licences | Copy EMServer.db to the C: Support/Data/Log Files folder | 6298
10567 | Install Customer Licences | Print out the Customer Entitlement Certicate - To be sent to the customer | 6299
10567 | Install Customer Licences | Store EM Server.db | 6300
1278 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6301
1278 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6302
1278 | ZEN Core EM | Install Customer's ZEN Licence | 6303
1278 | ZEN Core EM | Setup SmartSEM Toggle Button | 6304
1278 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6305
1278 | ZEN Core EM | Mains Sync Set To OFF | 6306
1278 | ZEN Core EM | Select Your ZEN core EM Runtime Profile | 6307
1278 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6308
1278 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 6309
1278 | ZEN Core EM | Configue Airlock Cam and Setup Exchange Position $Exchange (use Upper case E in ZEN) | 6310
1278 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6311
1278 | ZEN Core EM | Enable Project Folder Mode | 6312
1278 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6313
1278 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6314
1278 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6315
1278 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6316
1278 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6317
3309 | Atlas 5 | Connect Silver Box and Cables | 6318
3309 | Atlas 5 | Download latest Version Of Software (see link in docs) | 6319
3309 | Atlas 5 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info | 6320
3309 | Atlas 5 | Request Licence File from Fibics | 6321
3309 | Atlas 5 | Install Atlas Software | 6322
3309 | Atlas 5 | Install Licence File | 6323
3309 | Atlas 5 | Calibrate SEM | 6324
3309 | Atlas 5 | Obtain Mosaic | 6325
3309 | Atlas 5 | Set ATLAS Rotation Offset | 6326
3309 | Atlas 5 | Save ATLAS Licence Document | 6327
1280 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6328
1280 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6329
1280 | ZEN Core EM | Install Customer's ZEN Licence | 6330
1280 | ZEN Core EM | Setup SmartSEM Toggle Button | 6331
1280 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6332
1280 | ZEN Core EM | Mains Sync Set To OFF | 6333
1280 | ZEN Core EM | Choose your ZEN core EM runtime profile | 6334
1280 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6335
1280 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 6336
1280 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 6337
1280 | ZEN Core EM | Configue $NavCam (Capital N and C) | 6338
1280 | ZEN Core EM | Configue $Exchange (capital E) | 6339
1280 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6340
1280 | ZEN Core EM | Enable Project Folder Mode | 6341
1280 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6342
1280 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6343
1280 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6344
1280 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6345
1280 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6346
1279 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6347
1279 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6348
1279 | ZEN Core EM | Install Customer's ZEN Licence | 6349
1279 | ZEN Core EM | Setup SmartSEM Toggle Button | 6350
1279 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6351
1279 | ZEN Core EM | Mains Sync Set To OFF | 6352
1279 | ZEN Core EM | Choose your ZEN core EM runtime profile | 6353
1279 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6354
1279 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 6355
1279 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 6356
1279 | ZEN Core EM | Configue $NavCam (Capital N and C) | 6357
1279 | ZEN Core EM | Configue $Exchange (capital E) | 6358
1279 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6359
1279 | ZEN Core EM | Enable Project Folder Mode | 6360
1279 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6361
1279 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6362
1279 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6363
1279 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6364
1279 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6365
11029 | Prepare Open X Configuration | Clear EPROM on VAC Extension | 6366
10548 | Atlas 5 | Connect Silver Box and Cables | 6367
10548 | Atlas 5 | Download latest Version Of Software (see link in docs) | 6368
10548 | Atlas 5 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info | 6369
10548 | Atlas 5 | Request Licence File from Fibics | 6370
10548 | Atlas 5 | Install Atlas Software | 6371
10548 | Atlas 5 | Install Licence File | 6372
10548 | Atlas 5 | Calibrate SEM | 6373
10548 | Atlas 5 | Obtain Mosaic | 6374
10548 | Atlas 5 | Set ATLAS Rotation Offset | 6375
10548 | Atlas 5 | Save ATLAS Licence Document | 6376
10868 | Atlas 5 | Connect Silver Box and Cables | 6377
10868 | Atlas 5 | Download latest Version Of Software | 6378
10868 | Atlas 5 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info | 6379
10868 | Atlas 5 | Request Licence File from Fibics | 6380
10868 | Atlas 5 | Install Atlas Software | 6381
10868 | Atlas 5 | Install Licence File | 6382
10868 | Atlas 5 | Calibrate SEM | 6383
10868 | Atlas 5 | Obtain Mosaic | 6384
10868 | Atlas 5 | Set ATLAS Rotation Offset | 6385
10868 | Atlas 5 | Save ATLAS Licence Document | 6386
8699 | Cable Up ( Lab6 ) | System Mains Voltage | 6387
8699 | Cable Up ( Lab6 ) | IF LaB6 and 120V - Transformer Installed | 6388
8699 | Cable Up ( Lab6 ) | Cable Up | 6389
8699 | Cable Up ( Lab6 ) | Label Cables | 6390
1472 | ZEN core EM | ZEN Software Installed | 6391
10906 | ZEN | ZEN Software Installed | 6392
1472 | ZEN core EM | ZEN Software Installed | 6393
1472 | ZEN core EM | Obtain a SE Image in ZEN | 6394
1472 | ZEN core EM | ZEN Software Installed | 6395
1472 | ZEN core EM | Obtain a SE Image in ZEN | 6396
1472 | ZEN core EM | Obtain a Nav Cam Image and Check that a POI can be Navigated to. | 6397
1472 | ZEN core EM | ZEN Software Installed | 6398
1472 | ZEN core EM | Obtain a SE Image in ZEN | 6399
1472 | ZEN core EM | Obtain a Nav Cam Image and Check that a POI can be Navigated to. | 6400
1473 | ZEN core EM | ZEN Software Installed | 6401
1473 | ZEN core EM | Obtain a SE Image in ZEN | 6402
1473 | ZEN core EM | Obtain a Nav Cam Image and Check that a POI can be Navigated to. | 6403
10926 | ZEN | ZEN Software Installed | 6404
1280 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6405
1280 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6406
1280 | ZEN Core EM | Install Customer's ZEN Licence | 6407
1280 | ZEN Core EM | Setup SmartSEM Toggle Button | 6408
1280 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6409
1280 | ZEN Core EM | Mains Sync Set To OFF | 6410
1280 | ZEN Core EM | Choose your ZEN core EM runtime profile | 6411
1280 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6412
1280 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 6413
1280 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 6414
1280 | ZEN Core EM | Configue $NavCam (Capital N and C) | 6415
1280 | ZEN Core EM | Configue $Exchange (capital E) | 6416
1280 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6417
1280 | ZEN Core EM | Enable Project Folder Mode | 6418
1280 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6419
1280 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6420
1280 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6421
1280 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6422
1280 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6423
344 | Ripple | Actual X-Axis | 6424
344 | Ripple | Store X ripple shots and results | 6425
344 | Ripple | Actual Y-Axis | 6426
344 | Ripple | Store Y ripple shots and results | 6427
1476 | Configure PC ASAP | PC S/N | 6428
1476 | Configure PC ASAP | Hard disk Image (Version) | 6429
1476 | Configure PC ASAP | Record Operating System | 6430
1476 | Configure PC ASAP | Check the Entitlement certificate to determine which version of Smartsem to install | 6431
1476 | Configure PC ASAP | Install Smart SEM (See WI 9693) Record Version | 6432
1476 | Configure PC ASAP | Check for available SmartSEMService Packs and install them. (N/A if none) | 6433
1476 | Configure PC ASAP | Check for available SmartSEM hotfixes and install them. (N/A if none) | 6434
1476 | Configure PC ASAP | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 6435
1476 | Configure PC ASAP | Install CZ Support Tools. Ver | 6436
1476 | Configure PC ASAP | Disable Automatic Updates | 6437
1476 | Configure PC ASAP | Run ASAP Screening Program | 6438
1476 | Configure PC ASAP | System Specific Datazone Loaded | 6439
1476 | Configure PC ASAP | In SEM status set 'Mains Sync' to off | 6440
1476 | Configure PC ASAP | PC S/N | 6441
1476 | Configure PC ASAP | Hard disk Image (Version) | 6442
1476 | Configure PC ASAP | Record Operating System | 6443
1476 | Configure PC ASAP | Check the Entitlement certificate to determine which version of Smartsem to install | 6444
1476 | Configure PC ASAP | Install Smart SEM (See WI 9693) Record Version | 6445
1476 | Configure PC ASAP | Check for available SmartSEMService Packs and install them. (N/A if none) | 6446
1476 | Configure PC ASAP | Check for available SmartSEM hotfixes and install them. (N/A if none) | 6447
1476 | Configure PC ASAP | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 6448
1476 | Configure PC ASAP | Install CZ Support Tools. Ver | 6449
1476 | Configure PC ASAP | Disable Automatic Updates | 6450
1476 | Configure PC ASAP | Run ASAP Screening Program | 6451
1476 | Configure PC ASAP | System Specific Datazone Loaded | 6452
1476 | Configure PC ASAP | In SEM status set 'Mains Sync' to off | 6453
1476 | Configure PC ASAP (Crossbeam 750) | PC S/N | 6454
1476 | Configure PC ASAP (Crossbeam 750) | Hard disk Image (Version) | 6455
1476 | Configure PC ASAP (Crossbeam 750) | Record Operating System | 6456
1476 | Configure PC ASAP (Crossbeam 750) | Check the Entitlement certificate to determine which version of Smartsem to install | 6457
1476 | Configure PC ASAP (Crossbeam 750) | Install Smart SEM (See WI 9693) Record Version | 6458
1476 | Configure PC ASAP (Crossbeam 750) | Check for available SmartSEMService Packs and install them. (N/A if none) | 6459
1476 | Configure PC ASAP (Crossbeam 750) | Check for available SmartSEM hotfixes and install them. (N/A if none) | 6460
1476 | Configure PC ASAP (Crossbeam 750) | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 6461
1476 | Configure PC ASAP (Crossbeam 750) | Install CZ Support Tools. Ver | 6462
1476 | Configure PC ASAP (Crossbeam 750) | Disable Automatic Updates | 6463
1476 | Configure PC ASAP (Crossbeam 750) | Run ASAP Screening Program | 6464
1476 | Configure PC ASAP (Crossbeam 750) | System Specific Datazone Loaded | 6465
1476 | Configure PC ASAP (Crossbeam 750) | In SEM status set 'Mains Sync' to off | 6466
252 | Configure PC ASAP | PC S/N | 6467
252 | Configure PC ASAP | Hard disk Image (Version) | 6468
252 | Configure PC ASAP | Record Operating System | 6469
252 | Configure PC ASAP | Check the Entitlement certificate to determine which version of Smartsem to install | 6470
252 | Configure PC ASAP | Install Smart SEM (See WI 9693) Record Version | 6471
252 | Configure PC ASAP | Check for available SmartSEMService Packs and install them. (N/A if none) | 6472
252 | Configure PC ASAP | Check for available SmartSEM hotfixes and install them. (N/A if none) | 6473
252 | Configure PC ASAP | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 6474
252 | Configure PC ASAP | Install CZ Support Tools. Ver | 6475
252 | Configure PC ASAP | Disable Automatic Updates | 6476
252 | Configure PC ASAP | Run ASAP Screening Program | 6477
252 | Configure PC ASAP | System Specific Datazone Loaded | 6478
252 | Configure PC ASAP | In SEM status set 'Mains Sync' to off | 6479
252 | Configure PC ASAP | PC S/N | 6480
252 | Configure PC ASAP | Hard disk Image (Version) | 6481
252 | Configure PC ASAP | Record Operating System | 6482
252 | Configure PC ASAP | Check the Entitlement certificate to determine which version of Smartsem to install | 6483
252 | Configure PC ASAP | Install Smart SEM (See WI 9693) Record Version | 6484
252 | Configure PC ASAP | Check for available SmartSEMService Packs and install them. (N/A if none) | 6485
252 | Configure PC ASAP | Check for available SmartSEM hotfixes and install them. (N/A if none) | 6486
252 | Configure PC ASAP | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 6487
252 | Configure PC ASAP | Install CZ Support Tools. Ver | 6488
252 | Configure PC ASAP | Disable Automatic Updates | 6489
252 | Configure PC ASAP | Run ASAP Screening Program | 6490
252 | Configure PC ASAP | System Specific Datazone Loaded | 6491
252 | Configure PC ASAP | In SEM status set 'Mains Sync' to off | 6492
1476 | Configure PC ASAP (Crossbeam 750) | PC S/N | 6493
1476 | Configure PC ASAP (Crossbeam 750) | Hard disk Image (Version) | 6494
1476 | Configure PC ASAP (Crossbeam 750) | Record Operating System | 6495
1476 | Configure PC ASAP (Crossbeam 750) | Check the Entitlement certificate to determine which version of Smartsem to install | 6496
1476 | Configure PC ASAP (Crossbeam 750) | Install Smart SEM (Crossbeam 750) (See WI 9693) Record Version | 6497
1476 | Configure PC ASAP (Crossbeam 750) | Check for available SmartSEMService Packs and install them. (N/A if none) | 6498
1476 | Configure PC ASAP (Crossbeam 750) | Check for available SmartSEM hotfixes and install them. (N/A if none) | 6499
1476 | Configure PC ASAP (Crossbeam 750) | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 6500
1476 | Configure PC ASAP (Crossbeam 750) | Install CZ Support Tools. Ver | 6501
1476 | Configure PC ASAP (Crossbeam 750) | Disable Automatic Updates | 6502
1476 | Configure PC ASAP (Crossbeam 750) | Run ASAP Screening Program | 6503
1476 | Configure PC ASAP (Crossbeam 750) | System Specific Datazone Loaded | 6504
1476 | Configure PC ASAP (Crossbeam 750) | In SEM status set 'Mains Sync' to off | 6505
1476 | Configure PC ASAP (Crossbeam 750) | PC S/N | 6506
1476 | Configure PC ASAP (Crossbeam 750) | Hard disk Image (Version) | 6507
1476 | Configure PC ASAP (Crossbeam 750) | Record Operating System | 6508
1476 | Configure PC ASAP (Crossbeam 750) | Check the Entitlement certificate to determine which version of Smartsem to install | 6509
1476 | Configure PC ASAP (Crossbeam 750) | Install Smart SEM - This system has a Gemini 4 column and must be installled as a Crossbeam 750 | 6510
1476 | Configure PC ASAP (Crossbeam 750) | Check for available SmartSEMService Packs and install them. (N/A if none) | 6511
1476 | Configure PC ASAP (Crossbeam 750) | Check for available SmartSEM hotfixes and install them. (N/A if none) | 6512
1476 | Configure PC ASAP (Crossbeam 750) | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 6513
1476 | Configure PC ASAP (Crossbeam 750) | Install CZ Support Tools. Ver | 6514
1476 | Configure PC ASAP (Crossbeam 750) | Disable Automatic Updates | 6515
1476 | Configure PC ASAP (Crossbeam 750) | Run ASAP Screening Program | 6516
1476 | Configure PC ASAP (Crossbeam 750) | System Specific Datazone Loaded | 6517
1476 | Configure PC ASAP (Crossbeam 750) | In SEM status set 'Mains Sync' to off | 6518
1476 | Configure PC ASAP (Crossbeam 750) | PC S/N | 6519
1476 | Configure PC ASAP (Crossbeam 750) | Hard disk Image (Version) | 6520
1476 | Configure PC ASAP (Crossbeam 750) | Record Operating System | 6521
1476 | Configure PC ASAP (Crossbeam 750) | Check the Entitlement certificate to determine which version of Smartsem to install | 6522
1476 | Configure PC ASAP (Crossbeam 750) | Install Smart SEM - This system has a Gemini 4 column and must be installled as a Crossbeam 750 | 6523
1476 | Configure PC ASAP (Crossbeam 750) | Install Smart SEM - This system has a Gemini 4 column and must be installled as a Crossbeam 750 | 6524
1476 | Configure PC ASAP (Crossbeam 750) | Check for available SmartSEMService Packs and install them. (N/A if none) | 6525
1476 | Configure PC ASAP (Crossbeam 750) | Check for available SmartSEM hotfixes and install them. (N/A if none) | 6526
1476 | Configure PC ASAP (Crossbeam 750) | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 6527
1476 | Configure PC ASAP (Crossbeam 750) | Install CZ Support Tools. Ver | 6528
1476 | Configure PC ASAP (Crossbeam 750) | Disable Automatic Updates | 6529
1476 | Configure PC ASAP (Crossbeam 750) | Run ASAP Screening Program | 6530
1476 | Configure PC ASAP (Crossbeam 750) | System Specific Datazone Loaded | 6531
1476 | Configure PC ASAP (Crossbeam 750) | In SEM status set 'Mains Sync' to off | 6532
1476 | Configure PC ASAP (Crossbeam 750) | PC S/N | 6533
1476 | Configure PC ASAP (Crossbeam 750) | Hard disk Image (Version) | 6534
1476 | Configure PC ASAP (Crossbeam 750) | Record Operating System | 6535
1476 | Configure PC ASAP (Crossbeam 750) | Check the Entitlement certificate to determine which version of Smartsem to install | 6536
1476 | Configure PC ASAP (Crossbeam 750) | Install Smart SEM - This system has a Gemini 4 column and must be installled as a Crossbeam 750 | 6537
1476 | Configure PC ASAP (Crossbeam 750) | Confirm SmartSEM is configured as Crossbeam 750 | 6538
1476 | Configure PC ASAP (Crossbeam 750) | Check for available SmartSEMService Packs and install them. (N/A if none) | 6539
1476 | Configure PC ASAP (Crossbeam 750) | Check for available SmartSEM hotfixes and install them. (N/A if none) | 6540
1476 | Configure PC ASAP (Crossbeam 750) | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 6541
1476 | Configure PC ASAP (Crossbeam 750) | Install CZ Support Tools. Ver | 6542
1476 | Configure PC ASAP (Crossbeam 750) | Disable Automatic Updates | 6543
1476 | Configure PC ASAP (Crossbeam 750) | Run ASAP Screening Program | 6544
1476 | Configure PC ASAP (Crossbeam 750) | System Specific Datazone Loaded | 6545
1476 | Configure PC ASAP (Crossbeam 750) | In SEM status set 'Mains Sync' to off | 6546
1483 | Configure PC ASAP (Crossbeam 330) | PC S/N | 6547
1483 | Configure PC ASAP (Crossbeam 330) | Hard disk Image (Version) | 6548
1483 | Configure PC ASAP (Crossbeam 330) | Record Operating System | 6549
1483 | Configure PC ASAP (Crossbeam 330) | Check the Entitlement certificate to determine which version of Smartsem to install | 6550
1483 | Configure PC ASAP (Crossbeam 330) | Install Smart SEM - This system has a Gemini 4 column and must be installled as a Crossbeam 750 | 6551
1483 | Configure PC ASAP (Crossbeam 330) | Confirm SmartSEM is configured as Crossbeam 750 | 6552
1483 | Configure PC ASAP (Crossbeam 330) | Check for available SmartSEMService Packs and install them. (N/A if none) | 6553
1483 | Configure PC ASAP (Crossbeam 330) | Check for available SmartSEM hotfixes and install them. (N/A if none) | 6554
1483 | Configure PC ASAP (Crossbeam 330) | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 6555
1483 | Configure PC ASAP (Crossbeam 330) | Install CZ Support Tools. Ver | 6556
1483 | Configure PC ASAP (Crossbeam 330) | Disable Automatic Updates | 6557
1483 | Configure PC ASAP (Crossbeam 330) | Run ASAP Screening Program | 6558
1483 | Configure PC ASAP (Crossbeam 330) | System Specific Datazone Loaded | 6559
1483 | Configure PC ASAP (Crossbeam 330) | In SEM status set 'Mains Sync' to off | 6560
1483 | Configure PC ASAP (Crossbeam 330) | PC S/N | 6561
1483 | Configure PC ASAP (Crossbeam 330) | Hard disk Image (Version) | 6562
1483 | Configure PC ASAP (Crossbeam 330) | Record Operating System | 6563
1483 | Configure PC ASAP (Crossbeam 330) | Check the Entitlement certificate to determine which version of Smartsem to install | 6564
1483 | Configure PC ASAP (Crossbeam 330) | Install Smart SEM - This system should be configured as a Crossbeam 330 (Early systems will use Beta version of Smartsem V8.04) | 6565
1483 | Configure PC ASAP (Crossbeam 330) | Confirm SmartSEM is configured as Crossbeam 330 | 6566
1483 | Configure PC ASAP (Crossbeam 330) | Check for available SmartSEMService Packs and install them. (N/A if none) | 6567
1483 | Configure PC ASAP (Crossbeam 330) | Check for available SmartSEM hotfixes and install them. (N/A if none) | 6568
1483 | Configure PC ASAP (Crossbeam 330) | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 6569
1483 | Configure PC ASAP (Crossbeam 330) | Install CZ Support Tools. Ver | 6570
1483 | Configure PC ASAP (Crossbeam 330) | Disable Automatic Updates | 6571
1483 | Configure PC ASAP (Crossbeam 330) | Run ASAP Screening Program | 6572
1483 | Configure PC ASAP (Crossbeam 330) | System Specific Datazone Loaded | 6573
1483 | Configure PC ASAP (Crossbeam 330) | In SEM status set 'Mains Sync' to off | 6574
252 | Configure PC ASAP | PC S/N | 6575
252 | Configure PC ASAP | Hard disk Image (Version) | 6576
252 | Configure PC ASAP | Record Operating System | 6577
252 | Configure PC ASAP | Check the Entitlement certificate to determine which version of Smartsem to install | 6578
252 | Configure PC ASAP | Install Smart SEM (See WI 9693) Record Version | 6579
252 | Configure PC ASAP | Check for available SmartSEMService Packs and install them. (N/A if none) | 6580
252 | Configure PC ASAP | Check for available SmartSEM hotfixes and install them. (N/A if none) | 6581
252 | Configure PC ASAP | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 6582
252 | Configure PC ASAP | Install CZ Support Tools. Ver | 6583
252 | Configure PC ASAP | Disable Automatic Updates | 6584
252 | Configure PC ASAP | Run ASAP Screening Program | 6585
252 | Configure PC ASAP | System Specific Datazone Loaded | 6586
252 | Configure PC ASAP | In SEM status set 'Mains Sync' to off | 6587
252 | Configure PC ASAP | PC S/N | 6588
252 | Configure PC ASAP | Hard disk Image (Version) | 6589
252 | Configure PC ASAP | Record Operating System | 6590
252 | Configure PC ASAP | Check the Entitlement certificate to determine which version of Smartsem to install | 6591
252 | Configure PC ASAP | Install Smart SEM (See WI 9693) Record Version | 6592
252 | Configure PC ASAP | Check for available SmartSEMService Packs and install them. (N/A if none) | 6593
252 | Configure PC ASAP | Check for available SmartSEM hotfixes and install them. (N/A if none) | 6594
252 | Configure PC ASAP | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 6595
252 | Configure PC ASAP | Install CZ Support Tools. Ver | 6596
252 | Configure PC ASAP | Disable Automatic Updates | 6597
252 | Configure PC ASAP | Run ASAP Screening Program | 6598
252 | Configure PC ASAP | System Specific Datazone Loaded | 6599
252 | Configure PC ASAP | In SEM status set 'Mains Sync' to off | 6600
1484 | Print and Fit SN Label for XB330 | ✘ 𝗡𝗢 CSA marking on SN label.. | 6601
1484 | Print and Fit SN Label for XB330 | ✘ 𝗡𝗢 SEMI marking on SN label. | 6602
1484 | Print and Fit SN Label for XB330 | ✘ 𝗡𝗢 CSA marking on SN label.. | 6603
1484 | Print and Fit SN Label for XB330 | ✘ 𝗡𝗢 SEMI marking on SN label. | 6604
1484 | Print and Fit SN Label for XB330 | Confirm that the only marking on the label is the CE mark | 6605
1236 | Print and Fit SN Label (NO CSA, NO SEMI) | ✘ 𝗡𝗢 CSA marking on SN label.. | 6606
1236 | Print and Fit SN Label (NO CSA, NO SEMI) | ✘ 𝗡𝗢 SEMI marking on SN label. | 6607
1546 | Heater HiPot Tests HV | HAL SN: | 6608
1546 | Heater HiPot Tests HV | HAL Calibration Date: | 6609
1546 | Heater HiPot Tests HV | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.1Ω) | 6610
1546 | Heater HiPot Tests HV | IGP Heaters (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6611
1546 | Heater HiPot Tests HV | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6612
1546 | Heater HiPot Tests HV | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6613
625 | Heater HiPot Tests VP | HAL SN: | 6614
625 | Heater HiPot Tests VP | HAL Calibration Date: | 6615
625 | Heater HiPot Tests VP | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.1Ω) | 6616
625 | Heater HiPot Tests VP | IGP Heaters (IGP#1) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6617
625 | Heater HiPot Tests VP | IGP Heaters (IGP#2) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6618
625 | Heater HiPot Tests VP | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6619
625 | Heater HiPot Tests VP | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6620
625 | Heater HiPot Tests VP | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6621
1204 | Heater HiPot Tests HV | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.1Ω) | 6622
1204 | Heater HiPot Tests HV | IGP Heaters (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6623
1204 | Heater HiPot Tests HV | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6624
1204 | Heater HiPot Tests HV | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6625
1203 | Heater HiPot Tests VP | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.1Ω) | 6626
1203 | Heater HiPot Tests VP | IGP Heaters (IGP#1) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6627
1203 | Heater HiPot Tests VP | IGP Heaters (IGP#2) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6628
1203 | Heater HiPot Tests VP | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6629
1203 | Heater HiPot Tests VP | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6630
1203 | Heater HiPot Tests VP | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6631
1270 | Heater HiPot Test | HAL SN: | 6632
1270 | Heater HiPot Test | HAL Calibration Date: | 6633
1270 | Heater HiPot Test | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.1Ω) | 6634
1270 | Heater HiPot Test | Column Heater (354920-9179-000). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6635
10470 | Administrator Factory Setup-Serial Number | Enter System Serial number into Administrator. | 6636
1546 | Heater HiPot Tests HV | HAL SN: | 6637
1546 | Heater HiPot Tests HV | HAL Calibration Date: | 6638
1546 | Heater HiPot Tests HV | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.1Ω) | 6639
1546 | Heater HiPot Tests HV | IGP Heaters (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6640
1546 | Heater HiPot Tests HV | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6641
1546 | Heater HiPot Tests HV | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6642
625 | Heater HiPot Tests VP | HAL SN: | 6643
625 | Heater HiPot Tests VP | HAL Calibration Date: | 6644
625 | Heater HiPot Tests VP | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.1Ω) | 6645
625 | Heater HiPot Tests VP | IGP Heaters (IGP#1) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6646
625 | Heater HiPot Tests VP | IGP Heaters (IGP#2) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6647
625 | Heater HiPot Tests VP | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6648
625 | Heater HiPot Tests VP | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6649
625 | Heater HiPot Tests VP | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6650
625 | Heater HiPot Tests VP | HAL SN: | 6651
625 | Heater HiPot Tests VP | HAL Calibration Date: | 6652
625 | Heater HiPot Tests VP | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.1Ω) | 6653
625 | Heater HiPot Tests VP | IGP Heaters (IGP#1) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6654
625 | Heater HiPot Tests VP | IGP Heaters (IGP#2) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6655
625 | Heater HiPot Tests VP | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6656
625 | Heater HiPot Tests VP | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6657
625 | Heater HiPot Tests VP | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6658
1204 | Heater HiPot Tests HV | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.1Ω) | 6659
1204 | Heater HiPot Tests HV | IGP Heaters (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6660
1204 | Heater HiPot Tests HV | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6661
1204 | Heater HiPot Tests HV | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6662
1204 | Heater HiPot Tests HV | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.1Ω) | 6663
1204 | Heater HiPot Tests HV | IGP Heaters (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6664
1204 | Heater HiPot Tests HV | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6665
1204 | Heater HiPot Tests HV | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6666
1204 | Heater HiPot Tests HV | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.1Ω) | 6667
1204 | Heater HiPot Tests HV | IGP Heaters (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6668
1204 | Heater HiPot Tests HV | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6669
1204 | Heater HiPot Tests HV | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6670
1203 | Heater HiPot Tests VP | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.1Ω) | 6671
1203 | Heater HiPot Tests VP | IGP Heaters (IGP#1) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6672
1203 | Heater HiPot Tests VP | IGP Heaters (IGP#2) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6673
1203 | Heater HiPot Tests VP | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6674
1203 | Heater HiPot Tests VP | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6675
1203 | Heater HiPot Tests VP | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA) | 6676
1203 | Heater HiPot Tests VP | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.1Ω) | 6677
1203 | Heater HiPot Tests VP | IGP Heaters (IGP#1) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6678
1203 | Heater HiPot Tests VP | IGP Heaters (IGP#2) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6679
1203 | Heater HiPot Tests VP | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6680
1203 | Heater HiPot Tests VP | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6681
1203 | Heater HiPot Tests VP | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6682
1270 | Heater HiPot Test | HAL SN: | 6683
1270 | Heater HiPot Test | HAL Calibration Date: | 6684
1270 | Heater HiPot Test | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.1Ω) | 6685
1270 | Heater HiPot Test | Column Heater (354920-9179-000). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.1 mA) | 6686
10565 | Copy Standard Rersolution Images | Copy the standard resolution images into the customer images folder C:Program Data/Carl Zeiss/SmartSEM/images/CusrtomerImagaes | 6687
10565 | Copy Standard Rersolution Images | Copy the standard resolution images into the customer images folder C:Program Data/Carl Zeiss/SmartSEM/Images/CusrtomerImagaes | 6688
1278 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6689
1278 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6690
1278 | ZEN Core EM | Install Customer's ZEN Licence | 6691
1278 | ZEN Core EM | Setup SmartSEM Toggle Button in Toolbar | 6692
1278 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6693
1278 | ZEN Core EM | Mains Sync Set To OFF | 6694
1278 | ZEN Core EM | Select Your ZEN core EM Runtime Profile | 6695
1278 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6696
1278 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 6697
1278 | ZEN Core EM | Configue Airlock Cam and Setup Exchange Position $Exchange (use Upper case E in ZEN) | 6698
1278 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6699
1278 | ZEN Core EM | Enable Project Folder Mode | 6700
1278 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6701
1278 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6702
1278 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6703
1278 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6704
1278 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6705
1546 | Heater HiPot Tests HV | HAL SN: | 6706
1546 | Heater HiPot Tests HV | HAL Calibration Date: | 6707
1546 | Heater HiPot Tests HV | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.5Ω) | 6708
1546 | Heater HiPot Tests HV | IGP Heaters (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6709
1546 | Heater HiPot Tests HV | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6710
1546 | Heater HiPot Tests HV | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6711
625 | Heater HiPot Tests VP | HAL SN: | 6712
625 | Heater HiPot Tests VP | HAL Calibration Date: | 6713
625 | Heater HiPot Tests VP | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.5Ω) | 6714
625 | Heater HiPot Tests VP | IGP Heaters (IGP#1) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6715
625 | Heater HiPot Tests VP | IGP Heaters (IGP#2) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6716
625 | Heater HiPot Tests VP | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6717
625 | Heater HiPot Tests VP | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6718
625 | Heater HiPot Tests VP | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6719
1204 | Heater HiPot Tests HV | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.5Ω) | 6720
1204 | Heater HiPot Tests HV | IGP Heaters (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6721
1204 | Heater HiPot Tests HV | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6722
1204 | Heater HiPot Tests HV | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6723
1203 | Heater HiPot Tests VP | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.5Ω) | 6724
1203 | Heater HiPot Tests VP | IGP Heaters (IGP#1) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6725
1203 | Heater HiPot Tests VP | IGP Heaters (IGP#2) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6726
1203 | Heater HiPot Tests VP | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6727
1203 | Heater HiPot Tests VP | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6728
1203 | Heater HiPot Tests VP | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 6729
1270 | Heater HiPot Test | HAL SN: | 6730
1270 | Heater HiPot Test | HAL Calibration Date: | 6731
1270 | Heater HiPot Test | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.5Ω) | 6732
1270 | Heater HiPot Test | Column Heater (354920-9179-000). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.1 mA) | 6733
1278 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6734
1278 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6735
1278 | ZEN Core EM | Install Customer's ZEN Licence | 6736
1278 | ZEN Core EM | Setup SmartSEM Toggle Button in Toolbar | 6737
1278 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6738
1278 | ZEN Core EM | Mains Sync Set To OFF | 6739
1278 | ZEN Core EM | Select Your ZEN core EM Runtime Profile | 6740
1278 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6741
1278 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 6742
1278 | ZEN Core EM | Configue Airlock Cam and Setup Exchange Position $Exchange (use Upper case E in ZEN) | 6743
1278 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6744
1278 | ZEN Core EM | Configue NavCam Position $NavCam | 6745
1278 | ZEN Core EM | Snap Nav Cam Image | 6746
1278 | ZEN Core EM | Enable Project Folder Mode | 6747
1278 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6748
1278 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6749
1278 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6750
1278 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6751
1278 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6752
1278 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6753
1278 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6754
1278 | ZEN Core EM | Install Customer's ZEN Licence | 6755
1278 | ZEN Core EM | Setup SmartSEM Toggle Button in Toolbar | 6756
1278 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6757
1278 | ZEN Core EM | Mains Sync Set To OFF | 6758
1278 | ZEN Core EM | Select Your ZEN core EM Runtime Profile | 6759
1278 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6760
1278 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 6761
1278 | ZEN Core EM | Configue Airlock Cam and Setup Exchange Position $Exchange (use Upper case E in ZEN) | 6762
1278 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6763
1278 | ZEN Core EM | Configue NavCam Position $NavCam | 6764
1278 | ZEN Core EM | Snap Nav Cam Image | 6765
1278 | ZEN Core EM | Set Up NavCam Registration (3 points) | 6766
1278 | ZEN Core EM | Enable Project Folder Mode | 6767
1278 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6768
1278 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6769
1278 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6770
1278 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6771
1278 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6772
1278 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6773
1278 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6774
1278 | ZEN Core EM | Install Customer's ZEN Licence | 6775
1278 | ZEN Core EM | Setup SmartSEM Toggle Button in Toolbar | 6776
1278 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6777
1278 | ZEN Core EM | Mains Sync Set To OFF | 6778
1278 | ZEN Core EM | Select Your ZEN core EM Runtime Profile | 6779
1278 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6780
1278 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 6781
1278 | ZEN Core EM | Configue Airlock Cam and Setup Exchange Position $Exchange (use Upper case E in ZEN) | 6782
1278 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6783
1278 | ZEN Core EM | Configue NavCam Position $NavCam | 6784
1278 | ZEN Core EM | Snap Nav Cam Image | 6785
1278 | ZEN Core EM | Set Up NavCam Registration (3 points) | 6786
1278 | ZEN Core EM | Enable Project Folder Mode | 6787
1278 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6788
1278 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6789
1278 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6790
1278 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6791
1278 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6792
1494 | Nano-probing Manual Tablefiller Entries | Nano-probing Tablefiller Entries successfully completed. | 6793
1494 | VP Table filler Entries | VP Tablefiller Entries successfully completed. | 6794
1278 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6795
1278 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6796
1278 | ZEN Core EM | Install Customer's ZEN Licence | 6797
1278 | ZEN Core EM | Setup SmartSEM Toggle Button in Toolbar | 6798
1278 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6799
1278 | ZEN Core EM | Mains Sync Set To OFF | 6800
1278 | ZEN Core EM | Select Your ZEN core EM Runtime Profile | 6801
1278 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6802
1278 | ZEN Core EM | Configue Airlock Cam and Setup Exchange Position $Exchange (use Upper case E in ZEN) | 6803
1278 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6804
1278 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 6805
1278 | ZEN Core EM | Snap Nav Cam Image | 6806
1278 | ZEN Core EM | Set Up NavCam Registration (3 points) | 6807
1278 | ZEN Core EM | Enable Project Folder Mode | 6808
1278 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6809
1278 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6810
1278 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6811
1278 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6812
1278 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6813
1278 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6814
1278 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6815
1278 | ZEN Core EM | Install Customer's ZEN Licence | 6816
1278 | ZEN Core EM | Setup SmartSEM Toggle Button in Toolbar | 6817
1278 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6818
1278 | ZEN Core EM | Mains Sync Set To OFF | 6819
1278 | ZEN Core EM | Select Your ZEN core EM Runtime Profile | 6820
1278 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6821
1278 | ZEN Core EM | Configue Airlock Cam and Setup Exchange Position $Exchange (use Upper case E in ZEN) | 6822
1278 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6823
1278 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 6824
1278 | ZEN Core EM | Snap NavCam Image | 6825
1278 | ZEN Core EM | Set Up NavCam Registration (3 points) | 6826
1278 | ZEN Core EM | Enable Project Folder Mode | 6827
1278 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6828
1278 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6829
1278 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6830
1278 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6831
1278 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6832
1279 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6833
1279 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6834
1279 | ZEN Core EM | Install Customer's ZEN Licence | 6835
1279 | ZEN Core EM | Setup SmartSEM Toggle Button | 6836
1279 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6837
1279 | ZEN Core EM | Mains Sync Set To OFF | 6838
1279 | ZEN Core EM | Choose your ZEN core EM runtime profile | 6839
1279 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6840
1279 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 6841
1279 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 6842
1279 | ZEN Core EM | Configue $Exchange (capital E) | 6843
1279 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6844
1279 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 6845
1279 | ZEN Core EM | Snap Nav Cam Image | 6846
1279 | ZEN Core EM | Set Up NavCam Registration (3 points) | 6847
1279 | ZEN Core EM | Enable Project Folder Mode | 6848
1279 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6849
1279 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6850
1279 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6851
1279 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6852
1279 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6853
1280 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6854
1280 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6855
1280 | ZEN Core EM | Install Customer's ZEN Licence | 6856
1280 | ZEN Core EM | Setup SmartSEM Toggle Button | 6857
1280 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6858
1280 | ZEN Core EM | Mains Sync Set To OFF | 6859
1280 | ZEN Core EM | Choose your ZEN core EM runtime profile | 6860
1280 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6861
1280 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 6862
1280 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 6863
1280 | ZEN Core EM | Configue $Exchange (capital E) | 6864
1280 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6865
1280 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 6866
1280 | ZEN Core EM | Snap Nav Cam Image | 6867
1280 | ZEN Core EM | Set Up NavCam Registration (3 points) | 6868
1280 | ZEN Core EM | Enable Project Folder Mode | 6869
1280 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6870
1280 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6871
1280 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6872
1280 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6873
1280 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6874
1494 | Nano-probing Manual Tablefiller Entries | Nano-probing Tablefiller Entries successfully completed. | 6875
1280 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6876
1280 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6877
1280 | ZEN Core EM | Install Customer's ZEN Licence | 6878
1280 | ZEN Core EM | Setup SmartSEM Toggle Button | 6879
1280 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6880
1280 | ZEN Core EM | Mains Sync Set To OFF | 6881
1280 | ZEN Core EM | Choose your ZEN core EM runtime profile | 6882
1280 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6883
1280 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 6884
1280 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 6885
1280 | ZEN Core EM | Configue $Exchange (capital E) | 6886
1280 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6887
1280 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 6888
1280 | ZEN Core EM | Snap Nav Cam Image | 6889
1280 | ZEN Core EM | Set Up NavCam Registration (3 points) | 6890
1280 | ZEN Core EM | Enable Project Folder Mode | 6891
1280 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6892
1280 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6893
1280 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6894
1280 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6895
1280 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6896
1280 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6897
1280 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6898
1280 | ZEN Core EM | Install Customer's ZEN Licence | 6899
1280 | ZEN Core EM | Setup SmartSEM Toggle Button | 6900
1280 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6901
1280 | ZEN Core EM | Mains Sync Set To OFF | 6902
1280 | ZEN Core EM | Choose your ZEN core EM runtime profile | 6903
1280 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6904
1280 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 6905
1280 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 6906
1280 | ZEN Core EM | Configue $Exchange (capital E) | 6907
1280 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6908
1280 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 6909
1280 | ZEN Core EM | Snap Nav Cam Image | 6910
1280 | ZEN Core EM | Set Up NavCam Registration (3 points) | 6911
1280 | ZEN Core EM | Check That NavCam Registration is Set Up Accurately | 6912
1280 | ZEN Core EM | Enable Project Folder Mode | 6913
1280 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6914
1280 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6915
1280 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6916
1280 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6917
1280 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6918
1278 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6919
1278 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6920
1278 | ZEN Core EM | Install Customer's ZEN Licence | 6921
1278 | ZEN Core EM | Setup SmartSEM Toggle Button in Toolbar | 6922
1278 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6923
1278 | ZEN Core EM | Mains Sync Set To OFF | 6924
1278 | ZEN Core EM | Select Your ZEN core EM Runtime Profile | 6925
1278 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6926
1278 | ZEN Core EM | Configue Airlock Cam and Setup Exchange Position $Exchange (use Upper case E in ZEN) | 6927
1278 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6928
1278 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 6929
1278 | ZEN Core EM | Snap NavCam Image | 6930
1278 | ZEN Core EM | Set Up NavCam Registration (3 points) | 6931
1278 | ZEN Core EM | Check That NavCam Registration is Set Up Accurately | 6932
1278 | ZEN Core EM | Enable Project Folder Mode | 6933
1278 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6934
1278 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6935
1278 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6936
1278 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6937
1278 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6938
1278 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6939
1278 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6940
1278 | ZEN Core EM | Install Customer's ZEN Licence | 6941
1278 | ZEN Core EM | Setup SmartSEM Toggle Button in Toolbar | 6942
1278 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6943
1278 | ZEN Core EM | Mains Sync Set To OFF | 6944
1278 | ZEN Core EM | Select Your ZEN core EM Runtime Profile | 6945
1278 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6946
1278 | ZEN Core EM | Configue Airlock Cam and Setup Exchange Position $Exchange (use Upper case E in ZEN) | 6947
1278 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6948
1278 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 6949
1278 | ZEN Core EM | Snap NavCam Image | 6950
1278 | ZEN Core EM | Set Up NavCam Registration (3 points) | 6951
1278 | ZEN Core EM | Check That NavCam Registration is Set Up Accurately | 6952
1278 | ZEN Core EM | Enable Project Folder Mode | 6953
1278 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6954
1278 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6955
1278 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6956
1278 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6957
1278 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6958
1278 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6959
1278 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6960
1278 | ZEN Core EM | Install Customer's ZEN Licence | 6961
1278 | ZEN Core EM | Setup SmartSEM Toggle Button in Toolbar | 6962
1278 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6963
1278 | ZEN Core EM | Mains Sync Set To OFF | 6964
1278 | ZEN Core EM | Select Your ZEN core EM Runtime Profile | 6965
1278 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6966
1278 | ZEN Core EM | Configue Airlock Cam and Setup Exchange Position $Exchange (use Upper case E in ZEN) | 6967
1278 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6968
1278 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 6969
1278 | ZEN Core EM | Snap NavCam Image | 6970
1278 | ZEN Core EM | Setup NavCam Registration (3 points) | 6971
1278 | ZEN Core EM | Check That NavCam Registration is Set Up Accurately | 6972
1278 | ZEN Core EM | Enable Project Folder Mode | 6973
1278 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6974
1278 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6975
1278 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6976
1278 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6977
1278 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 6978
1280 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 6979
1280 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 6980
1280 | ZEN Core EM | Install Customer's ZEN Licence | 6981
1280 | ZEN Core EM | Setup SmartSEM Toggle Button | 6982
1280 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 6983
1280 | ZEN Core EM | Mains Sync Set To OFF | 6984
1280 | ZEN Core EM | Choose your ZEN core EM runtime profile | 6985
1280 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 6986
1280 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 6987
1280 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 6988
1280 | ZEN Core EM | Configue $Exchange (capital E) | 6989
1280 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 6990
1280 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 6991
1280 | ZEN Core EM | Snap NavCam Image | 6992
1280 | ZEN Core EM | Setup NavCam Registration (3 points) | 6993
1280 | ZEN Core EM | Check That NavCam Registration is Set Up Accurately | 6994
1280 | ZEN Core EM | Enable Project Folder Mode | 6995
1280 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 6996
1280 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 6997
1280 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 6998
1280 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 6999
1280 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 7000
1280 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 7001
1280 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 7002
1280 | ZEN Core EM | Install Customer's ZEN Licence | 7003
1280 | ZEN Core EM | Setup SmartSEM Toggle Button | 7004
1280 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 7005
1280 | ZEN Core EM | Mains Sync Set To OFF | 7006
1280 | ZEN Core EM | Choose your ZEN core EM runtime profile | 7007
1280 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 7008
1280 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 7009
1280 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 7010
1280 | ZEN Core EM | Configue $Exchange (capital E) | 7011
1280 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 7012
1280 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 7013
1280 | ZEN Core EM | Snap NavCam Image | 7014
1280 | ZEN Core EM | Setup NavCam Registration (3 points) | 7015
1280 | ZEN Core EM | Check That NavCam Registration is Set Up Accurately | 7016
1280 | ZEN Core EM | Enable Project Folder Mode | 7017
1280 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 7018
1280 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 7019
1280 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 7020
1280 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 7021
1280 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 7022
1279 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 7023
1279 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 7024
1279 | ZEN Core EM | Install Customer's ZEN Licence | 7025
1279 | ZEN Core EM | Setup SmartSEM Toggle Button | 7026
1279 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 7027
1279 | ZEN Core EM | Mains Sync Set To OFF | 7028
1279 | ZEN Core EM | Choose your ZEN core EM runtime profile | 7029
1279 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 7030
1279 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 7031
1279 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 7032
1279 | ZEN Core EM | Configue $Exchange (capital E) | 7033
1279 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 7034
1279 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 7035
1279 | ZEN Core EM | Snap NavCam Image | 7036
1279 | ZEN Core EM | Setup NavCam Registration (3 points) | 7037
1279 | ZEN Core EM | Check That NavCam Registration is Set Up Accurately | 7038
1279 | ZEN Core EM | Enable Project Folder Mode | 7039
1279 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 7040
1279 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 7041
1279 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 7042
1279 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 7043
1279 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 7044
1279 | ZEN Core EM | Install Latest Opal Kelly Driver (5.3.3) | 7045
1279 | ZEN Core EM | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure | 7046
1279 | ZEN Core EM | Install Customer's ZEN Licence | 7047
1279 | ZEN Core EM | Setup SmartSEM Toggle Button | 7048
1279 | ZEN Core EM | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50) | 7049
1279 | ZEN Core EM | Mains Sync Set To OFF | 7050
1279 | ZEN Core EM | Choose your ZEN core EM runtime profile | 7051
1279 | ZEN Core EM | Set ZEN API Gateway Port to 5002 | 7052
1279 | ZEN Core EM | Upgrade Pneumatic Detector Firmware to 3.4.0 | 7053
1279 | ZEN Core EM | Calibrate and Configure All Detectors (including legacy detectors) | 7054
1279 | ZEN Core EM | Configue $Exchange (capital E) | 7055
1279 | ZEN Core EM | General Athena Settings - Tick Correct NavCam if Fitted | 7056
1279 | ZEN Core EM | Configue NavCam and Setup Position $NavCam (Capital N and C) | 7057
1279 | ZEN Core EM | Snap NavCam Image | 7058
1279 | ZEN Core EM | Setup NavCam Registration (3 points) | 7059
1279 | ZEN Core EM | Check That NavCam Registration is Set Up Accurately | 7060
1279 | ZEN Core EM | Enable Project Folder Mode | 7061
1279 | ZEN Core EM | Snap an Image and Save To Images Folder in SmartSEM | 7062
1279 | ZEN Core EM | Run a 3 x3 Tile Aquisition | 7063
1279 | ZEN Core EM | Remove the Static and Dynamic Calibration Workbenches | 7064
1279 | ZEN Core EM | Automated imaging - check automated imaging workbench is functional (only if licensed for automated imaging) | 7065
1279 | ZEN Core EM | EDS Workbenches can be Loaded (AZtec on order and AZtec 6.2 installed and setup) | 7066
1494 | Nano-probing Manual Tablefiller Entries | Nano-probing Tablefiller Entries successfully completed. | 7067
1494 | Nano Probing Table filler Entries | VP Tablefiller Entries successfully completed. | 7068
1494 | Nano-Probing Manual Table filler Entries | Nano-probing Tablefiller Entries successfully completed. | 7069
10565 | Copy Standard Rersolution Images | Copy the standard resolution images into the customer images folder C:Program Data/Carl Zeiss/SmartSEM/Images/CusrtomerImages | 7070
303 | Stage Drift Program - 4 & 5 Inch Stage | Actual (Spec: 40nm /6 min) | 7071
303 | Stage Drift Program - 4 & 5 Inch Stage | Actual (Spec: 40nm /6 min) | 7072
303 | Stage Drift Program - 4 & 5 Inch Stage | Store stage drift macro results image | 7073
1495 | Stage Drift Program - 6 & 8 Inch Stages | Actual (Spec: 40nm /6 min) | 7074
1495 | Stage Drift Program - 6 & 8 Inch Stages | Actual (Spec: 40nm /6 min) | 7075
1495 | Stage Drift Program - 6 & 8 Inch Stages | Store stage drift macro results image | 7076
1495 | Stage Drift Program - 6 & 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7077
1495 | Stage Drift Program - 6 & 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7078
1495 | Stage Drift Program - 6 & 8 Inch Stages | Store stage drift macro results image | 7079
1230 | HP Stage Checks | Stage mapping completed successfuly | 7080
3321 | Install Customer Licences | Delete the Production Test Licences | 7081
3321 | Install Customer Licences | Copy EMServer.db to the C: Support/Data/Log Files folder | 7082
3321 | Install Customer Licences | Obtain the customer Entitlement Certificate | 7083
3321 | Install Customer Licences | Download into the C:Support/Data/Licence folder | 7084
3321 | Install Customer Licences | Print out the Customer Entitlement Certificate - To be sent to the customer | 7085
3321 | Install Customer Licences | Load customer licences, SmartSEM & ZEN, Including ZEN Promo and check against order they are correct. | 7086
3321 | Install Customer Licences | Test customer licenses | 7087
354 | Install Customer Licences | Delete the Production Test Licences | 7088
354 | Install Customer Licences | Obtain the customer Entitlement Certificate | 7089
354 | Install Customer Licences | Download into the Engineer folder | 7090
354 | Install Customer Licences | Install the Customer Licences,SmartSEM & ZEN, Including ZEN Promo | 7091
354 | Install Customer Licences | Copy Contents of Image Folder to the Product Folder | 7092
354 | Install Customer Licences | Copy EMServer.db to the Product Folder | 7093
354 | Install Customer Licences | Print out the Customer Entitlement Certicate - To be sent to the customer | 7094
354 | Install Customer Licences | Delete the Production Test Licences | 7095
354 | Install Customer Licences | Obtain the customer Entitlement Certificate | 7096
354 | Install Customer Licences | Download into the Engineer folder | 7097
354 | Install Customer Licences | Install the Customer Licences, SmartSEM & ZEN, Including ZEN Promo | 7098
354 | Install Customer Licences | Copy Contents of Image Folder to the Product Folder | 7099
354 | Install Customer Licences | Copy EMServer.db to the Product Folder | 7100
354 | Install Customer Licences | Print out the Customer Entitlement Certicate - To be sent to the customer | 7101
303 | Stage Drift Program - 4 & 5 Inch Stage | Actual (Spec: 40nm /6 min) | 7102
303 | Stage Drift Program - 4 & 5 Inch Stage | Actual (Spec: 40nm /6 min) | 7103
303 | Stage Drift Program - 4 & 5 Inch Stage | Store stage drift macro results image | 7104
303 | Stage Drift Program - 4 & 5 Inch Stage | Actual (Spec: 40nm /6 min) | 7105
303 | Stage Drift Program - 4 & 5 Inch Stage | Actual (Spec: 40nm /6 min) | 7106
303 | Stage Drift Program - 4 & 5 Inch Stage | Store stage drift macro results image | 7107
1495 | Stage Drift Program - 6 & 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7108
1495 | Stage Drift Program - 6 & 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7109
1495 | Stage Drift Program - 6 & 8 Inch Stages | Store stage drift macro results image | 7110
1495 | Stage Drift Program - 6 & 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7111
1495 | Stage Drift Program - 6 & 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7112
1495 | Stage Drift Program - 6 & 8 Inch Stages | Store stage drift macro results image | 7113
3307 | Airlock | Set Rod Retract Switch | 7114
3307 | Airlock | Edit Airlock Macro | 7115
3307 | Airlock | Set Up Stage Exchange | 7116
3307 | Airlock | If cartesian 2 stage is fitted select exchange braking in Smartsem administrator | 7117
3307 | Airlock | Airlock Function Test | 7118
745 | Material Availability | Are there any shortages? | 7119
1230 | HP Stage Checks | Stage mapping completed successfuly | 7120
1230 | HP Stage Checks | Stage mapping completed successfuly | 7121
299 | Stage Repeatability Macro | Macro Result 1 | 7122
299 | Stage Repeatability Macro | Macro Result 2 | 7123
299 | Stage Repeatability Macro | Macro Result 3 | 7124
299 | Stage Repeatability Macro | Image of Macro Results | 7125
30 | Fit System cable - EHT | Fit EHT Cables | 7126
30 | Fit System cable - EHT | Torque EHT Connections to 3Nm | 7127
748 | Material Availability | Are there any shortages? | 7128
11389 | Preparing and Fitting Chamber | Fit chamber pumping pipe | 7129
11389 | Preparing and Fitting Chamber | Fit chamber to damper | 7130
11389 | Preparing and Fitting Chamber | Fit chamber bolts | 7131
11389 | Preparing and Fitting Chamber | Fit earthquake bolts | 7132
11389 | Preparing and Fitting Chamber | Record Chamber SN | 7133
1190 | Stage Baking Trial (Temp test point) | Was the Stage baked independently? (Chamber not baked) | 7134
1496 | Stage Repeatability Macro | Select Stage Type | 7135
1496 | Stage Repeatability Macro | Macro Result 1 (Max 5) | 7136
1496 | Stage Repeatability Macro | Macro Result 2 (Max 5) | 7137
1496 | Stage Repeatability Macro | Macro Result 3 (Max 5) | 7138
1496 | Stage Repeatability Macro | Image of Macro Results | 7139
1496 | Stage Repeatability Macro | Macro Result 1 (Max 5) | 7140
1496 | Stage Repeatability Macro | Macro Result 2 (Max 5) | 7141
1496 | Stage Repeatability Macro | Macro Result 3 (Max 5) | 7142
1496 | Stage Repeatability Macro | Image of Macro Results | 7143
1496 | Stage Repeatability Macro | Macro Result 1 (Max 5) | 7144
1496 | Stage Repeatability Macro | Macro Result 2 (Max 5) | 7145
1496 | Stage Repeatability Macro | Macro Result 3 (Max 5) | 7146
1496 | Stage Repeatability Macro | Image of Macro Results | 7147
1496 | Stage Repeatability Macro | Macro Result 1 (Max 3um) | 7148
1496 | Stage Repeatability Macro | Macro Result 2 (Max 3um) | 7149
1496 | Stage Repeatability Macro | Macro Result 3 (Max 3um) | 7150
1496 | Stage Repeatability Macro | Image of Macro Results | 7151
1496 | Stage Repeatability Macro | Macro Result 1 (Max 3um) | 7152
1496 | Stage Repeatability Macro | Macro Result 2 (Max 3um) | 7153
1496 | Stage Repeatability Macro | Macro Result 3 (Max 3um) | 7154
1496 | Stage Repeatability Macro | Image of Macro Results | 7155
1496 | Stage Repeatability Macro | Macro Result 1 (Max 3um) | 7156
1496 | Stage Repeatability Macro | Macro Result 2 (Max 3um) | 7157
1496 | Stage Repeatability Macro | Macro Result 3 (Max 3um) | 7158
1496 | Stage Repeatability Macro | Image of Macro Results | 7159
1496 | Stage Repeatability Macro | Macro Result 1 (Max 3um) | 7160
1496 | Stage Repeatability Macro | Macro Result 2 (Max 3um) | 7161
1496 | Stage Repeatability Macro | Macro Result 3 (Max 3um) | 7162
1496 | Stage Repeatability Macro | Image of Macro Results | 7163
1496 | Stage Repeatability Macro | Macro Result 1 (Max 3um) | 7164
1496 | Stage Repeatability Macro | Macro Result 2 (Max 3um) | 7165
1496 | Stage Repeatability Macro | Macro Result 3 (Max 3um) | 7166
1496 | Stage Repeatability Macro | Image of Macro Results | 7167
1496 | Stage Repeatability Macro | Macro Result 1 (Max 3um) | 7168
1496 | Stage Repeatability Macro | Macro Result 2 (Max 3um) | 7169
1496 | Stage Repeatability Macro | Macro Result 3 (Max 3um) | 7170
1496 | Stage Repeatability Macro | Image of Macro Results | 7171
1230 | HP Stage Checks | Macro Result 1 (Max 3um) | 7172
1230 | HP Stage Checks | Macro Result 2 (Max 3um) | 7173
1230 | HP Stage Checks | Macro Result 3 (Max 3um) | 7174
1230 | HP Stage Checks | Image of Macro Results | 7175
1230 | HP Stage Checks | Macro Result 1 (Max 3um) | 7176
1230 | HP Stage Checks | Macro Result 2 (Max 3um) | 7177
1230 | HP Stage Checks | Macro Result 3 (Max 3um) | 7178
1230 | HP Stage Checks | Image of Macro Results | 7179
1230 | HP Stage Checks | Macro Result 1 (Max 3um) | 7180
1230 | HP Stage Checks | Macro Result 2 (Max 3um) | 7181
1230 | HP Stage Checks | Macro Result 3 (Max 3um) | 7182
1230 | HP Stage Checks | Image of Macro Results | 7183
1230 | HP Stage Checks | Macro Result 1 (Max 3um) | 7184
1230 | HP Stage Checks | Macro Result 2 (Max 3um) | 7185
1230 | HP Stage Checks | Macro Result 3 (Max 3um) | 7186
1230 | HP Stage Checks | Image of Macro Results | 7187
1230 | HP Stage Checks | Macro Result 1 (Max 3um) | 7188
1230 | HP Stage Checks | Macro Result 2 (Max 3um) | 7189
1230 | HP Stage Checks | Macro Result 3 (Max 3um) | 7190
1230 | HP Stage Checks | Image of Macro Results | 7191
1230 | HP Stage Checks | Macro Result 1 (Max 0.5um) | 7192
1230 | HP Stage Checks | Macro Result 2 (Max 0.5um) | 7193
1230 | HP Stage Checks | Macro Result 3 (Max 0.5um) | 7194
1230 | HP Stage Checks | Image of Macro Results | 7195
1230 | HP Stage Checks (4/5/6 inch stage) | Macro Result 1 (Max 0.5um) | 7196
1230 | HP Stage Checks (4/5/6 inch stage) | Macro Result 2 (Max 0.5um) | 7197
1230 | HP Stage Checks (4/5/6 inch stage) | Macro Result 3 (Max 0.5um) | 7198
1230 | HP Stage Checks (4/5/6 inch stage) | Image of Macro Results | 7199
1497 | HP Stage Checks (8 inch stage) | Macro Result 1 (Max 0.5um) | 7200
1497 | HP Stage Checks (8 inch stage) | Macro Result 2 (Max 0.5um) | 7201
1497 | HP Stage Checks (8 inch stage) | Macro Result 3 (Max 0.5um) | 7202
1497 | HP Stage Checks (8 inch stage) | Image of Macro Results | 7203
1497 | HP Stage Checks (8 inch stage) | Full Travel Macro Result 1 (Max 0.5um) | 7204
1497 | HP Stage Checks (8 inch stage) | Macro Result 2 (Max 0.5um) | 7205
1497 | HP Stage Checks (8 inch stage) | Macro Result 3 (Max 0.5um) | 7206
1497 | HP Stage Checks (8 inch stage) | Macro Result 1 (Max 0.5um) | 7207
1497 | HP Stage Checks (8 inch stage) | Macro Result 2 (Max 0.5um) | 7208
1497 | HP Stage Checks (8 inch stage) | Macro Result 3 (Max 0.5um) | 7209
1497 | HP Stage Checks (8 inch stage) | Image of Macro Results | 7210
1230 | HP Stage Repeatability (4/5/6 inch stage) | Macro Result 1 (Max 0.5um) | 7211
1230 | HP Stage Repeatability (4/5/6 inch stage) | Macro Result 2 (Max 0.5um) | 7212
1230 | HP Stage Repeatability (4/5/6 inch stage) | Macro Result 3 (Max 0.5um) | 7213
1230 | HP Stage Repeatability (4/5/6 inch stage) | Image of Macro Results | 7214
1497 | HP Stage Repeatability (8 inch stage) | Full Travel Macro Result 1 (Max 0.5um) | 7215
1497 | HP Stage Repeatability (8 inch stage) | Macro Result 2 (Max 0.5um) | 7216
1497 | HP Stage Repeatability (8 inch stage) | Macro Result 3 (Max 0.5um) | 7217
1497 | HP Stage Repeatability (8 inch stage) | Macro Result 1 (Max 0.5um) | 7218
1497 | HP Stage Repeatability (8 inch stage) | Macro Result 2 (Max 0.5um) | 7219
1497 | HP Stage Repeatability (8 inch stage) | Macro Result 3 (Max 0.5um) | 7220
1497 | HP Stage Repeatability (8 inch stage) | Image of Macro Results | 7221
1497 | HP Stage Repeatability (8 inch stage - Full Travel) | Full Travel Macro Result 1 (Max 0.5um) | 7222
1497 | HP Stage Repeatability (8 inch stage - Full Travel) | Macro Result 2 (Max 0.5um) | 7223
1497 | HP Stage Repeatability (8 inch stage - Full Travel) | Macro Result 3 (Max 0.5um) | 7224
1497 | HP Stage Repeatability (8 inch stage - Full Travel) | Macro Result 1 (Max 0.5um) | 7225
1497 | HP Stage Repeatability (8 inch stage - Full Travel) | Macro Result 2 (Max 0.5um) | 7226
1497 | HP Stage Repeatability (8 inch stage - Full Travel) | Macro Result 3 (Max 0.5um) | 7227
1497 | HP Stage Repeatability (8 inch stage - Full Travel) | Image of Macro Results | 7228
1230 | HP Stage Repeatability (4/5/6 inch stage) | Macro Result 1 (Max 0.5um) | 7229
1230 | HP Stage Repeatability (4/5/6 inch stage) | Macro Result 2 (Max 0.5um) | 7230
1230 | HP Stage Repeatability (4/5/6 inch stage) | Macro Result 3 (Max 0.5um) | 7231
1230 | HP Stage Repeatability (4/5/6 inch stage) | Confirm Ripple X (Max 0.8nm) | 7232
1230 | HP Stage Repeatability (4/5/6 inch stage) | Confirm Ripple Y (Max 0.8nm) | 7233
1230 | HP Stage Repeatability (4/5/6 inch stage) | Image of Macro Results | 7234
1502 | Standby to Operate | Rear Plinth Fan Running in Standby (yellow button) and Vacuum PCB Only Powered Up | 7235
1502 | Standby to Operate | EO/Stage Fans Not Running In Standby (Yellow button) | 7236
1502 | Standby to Operate | Rear Plinth Fan and EO/Stage Fans Running in Operate (green button) & All PCB's Powered UP. | 7237
1502 | Standby to Operate | PC Runs Up | 7238
1499 | HP Stage Mapping (8 inch stage) | Run Stage Mapping Procedure | 7239
1499 | HP Stage Mapping (8 inch stage) | Record Accuracy X (Max 10um) | 7240
1499 | HP Stage Mapping (8 inch stage) | Record Accuracy Y (Max 10um) | 7241
1500 | HP Stage Mapping (4/5/6 inch stage) | Run Stage Mapping Procedure | 7242
1500 | HP Stage Mapping (4/5/6 inch stage) | Record Accuracy X (Max 5um) | 7243
1500 | HP Stage Mapping (4/5/6 inch stage) | Record Accuracy Y (Max 5um) | 7244
1500 | HP Stage Mapping (4/5/6 inch stage) | Run Stage Mapping Procedure | 7245
1500 | HP Stage Mapping (4/5/6 inch stage) | Record Accuracy X (Max 5um) | 7246
1500 | HP Stage Mapping (4/5/6 inch stage) | Record Accuracy Y (Max 5um) | 7247
1503 | Joystick 6 Axis | All Axes Funtion | 7248
1503 | Joystick 6 Axis | M Axes Buttons Funtion | 7249
1503 | Joystick 6 Axis | Break Button Function | 7250
1495 | Stage Drift Program - 6 & 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7251
1495 | Stage Drift Program - 6 & 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7252
1495 | Stage Drift Program - 6 & 8 Inch Stages | Store stage drift macro results image | 7253
303 | Stage Drift Program - 4 & 5 Inch Stage | Actual (Spec: 40nm /6 min) | 7254
303 | Stage Drift Program - 4 & 5 Inch Stage | Actual (Spec: 40nm /6 min) | 7255
303 | Stage Drift Program - 4 & 5 Inch Stage | Store stage drift macro results image | 7256
1499 | HP Stage Mapping (8 inch stage) | Run Stage Mapping Procedure | 7257
1499 | HP Stage Mapping (8 inch stage) | Record Accuracy X (Max 10um) | 7258
1499 | HP Stage Mapping (8 inch stage) | Record Accuracy Y (Max 10um) | 7259
1500 | HP Stage Mapping (4/5/6 inch stage) | Run Stage Mapping Procedure | 7260
1500 | HP Stage Mapping (4/5/6 inch stage) | Record Accuracy X (Max 5um) | 7261
1500 | HP Stage Mapping (4/5/6 inch stage) | Record Accuracy Y (Max 5um) | 7262
1495 | Stage Drift Program - 6 & 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7263
1495 | Stage Drift Program - 6 & 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7264
1495 | Stage Drift Program - 6 & 8 Inch Stages | Store stage drift macro results image | 7265
1495 | Stage Drift Program - 6 & 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7266
1495 | Stage Drift Program - 6 & 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7267
1495 | Stage Drift Program - 6 & 8 Inch Stages | Store stage drift macro results image | 7268
1495 | Stage Drift Program - 6 + 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7269
1495 | Stage Drift Program - 6 + 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7270
1495 | Stage Drift Program - 6 + 8 Inch Stages | Store stage drift macro results image | 7271
303 | Stage Drift Program - 4 + 5 Inch Stage | Actual (Spec: 40nm /6 min) | 7272
303 | Stage Drift Program - 4 + 5 Inch Stage | Actual (Spec: 40nm /6 min) | 7273
303 | Stage Drift Program - 4 + 5 Inch Stage | Store stage drift macro results image | 7274
303 | Stage Drift Program - 4 + 5 Inch Stage | Actual (Spec: 40nm /6 min) | 7275
303 | Stage Drift Program - 4 + 5 Inch Stage | Actual (Spec: 40nm /6 min) | 7276
303 | Stage Drift Program - 4 + 5 Inch Stage | Store stage drift macro results image | 7277
1501 | Stage Drift Program - 6 + 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7278
1501 | Stage Drift Program - 6 + 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7279
1501 | Stage Drift Program - 6 + 8 Inch Stages | Store stage drift macro results image | 7280
1501 | Stage Drift Program - 6 + 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7281
1501 | Stage Drift Program - 6 + 8 Inch Stages | Actual (Spec: 100nm /8 min) | 7282
1501 | Stage Drift Program - 6 + 8 Inch Stages | Store stage drift macro results image | 7283
916 | Volutome | 1. Unpack Volutome and check all parts on the packing list are present (packing list link can be found in attachments) | 7284
916 | Volutome | 1. Upload data from USB stick to the O drive product folder | 7285
916 | Volutome | 1. Remove transit clamps from Volutome stage | 7286
916 | Volutome | 1. Ensure both safety switches are fitted to the Volutome door | 7287
916 | Volutome | 1. Check Engineers toolbox for completeness. | 7288
916 | Volutome | 1. Check Excentric pin fitment within holder. | 7289
916 | Volutome | 2. Ensure hardware has been configured according to work instructions | 7290
916 | Volutome | 2. Fit Volutome stage | 7291
916 | Volutome | 2. Check protection circuit for Piezo is fitted (349567-9013-000) | 7292
916 | Volutome | 2. Connect Volutome stage and controller to the SEM | 7293
916 | Volutome | 2. Fit Stemi holder bracket to Volutome stage door (STEMI adjustment bracket needs to fit / disconnect without force from door) | 7294
916 | Volutome | 2. Check all lights and fan come on when the electonics rack is powered on. | 7295
916 | Volutome | 3. Ethernet Connection Setup | 7296
916 | Volutome | 4. Install the latest version of Zen - Ensure to configure it for Volutome (Refer to Zen installation WI) | 7297
916 | Volutome | 4. Ensure Athena inspector manager has all green lights, Physik microtome and Physik stage are present - save a screenshot in the Volutome folder | 7298
916 | Volutome | 4. Install PI Micromove | 7299
916 | Volutome | 4. Run getversions.exe and save text file in the Volutome folder | 7300
916 | Volutome | 4. Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator | 7301
916 | Volutome | 4. Set up the Volutome workbench in Zen | 7302
916 | Volutome | 5. PID parameters configured | 7303
916 | Volutome | 5. Retractor piezo auto zero completed (after at least 24 hours at vacuum and electronics rack powered). | 7304
916 | Volutome | 6. Record VolumeBSD serial number | 7305
916 | Volutome | 6. Record Volutome stage serial number | 7306
916 | Volutome | 6. Record VolumeBSD amplifier serial number | 7307
916 | Volutome | 6. Record VolumeBSD mechanics serial number | 7308
916 | Volutome | 6. Record C-885 PI controller serial number | 7309
916 | Volutome | 6. Record E-727 PI controller serial number | 7310
916 | Volutome | 7. Volutome 'internal LED light' function check | 7311
916 | Volutome | 8. Volutome 'External cold light' function check | 7312
916 | Volutome | 9. Adjust Volutome stage centre (Crosshairs should be within 4 central squares) | 7313
916 | Volutome | 10. Install Nav cam in-line connections (If applicable) | 7314
916 | Volutome | 11. Perform the microtome and stage functionality checks | 7315
916 | Volutome | 12. Fit and align the VolumeBSD diode | 7316
916 | Volutome | 12. Check VolumeBSD alignment repeatability | 7317
916 | Volutome | 12. Check all gain ranges of the VolumeBSD (normal/inverted) | 7318
916 | Volutome | 12. Perform Signal/noise ratio check | 7319
916 | Volutome | 12. Ensure BSD diode protector is fitted and adjusted in chamber | 7320
916 | Volutome | 12. Install cable tie to loop BSD signal cable | 7321
916 | Volutome | 12. Test BSD functionality in Zen - Insert, retract and gain ranges all functional | 7322
916 | Volutome | 13. Check SCM with Volutome stage | 7323
916 | Volutome | 14. Noise check - Check there is no visible oscillation when imaging with a pixel size of 5nm | 7324
916 | Volutome | 15. Oscillator continuity check | 7325
916 | Volutome | 16. Perform X-ray check | 7326
916 | Volutome | 17. Pump Down Time | 7327
916 | Volutome | 17. Ultimate Vacuum | 7328
916 | Volutome | 18. Focal CC fitted to Volutome stage | 7329
916 | Volutome | 18. Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar | 7330
916 | Volutome | 18. Remove carbon needle from focal CC after testing and place in the engineers toolbox container | 7331
916 | Volutome | 19. Generate live image from Axiocam | 7332
916 | Volutome | 20. Dummy experiment completed with no erros reported in Zen - Results saved in Volutome folder | 7333
916 | Volutome | 21. Decouple motors | 7334
916 | Volutome | 21. Re-package Volutome and check it's completeness | 7335
916 | Volutome | 22. Check functionality of all additional software installed e.g Atlas | 7336
916 | Volutome | 22. Check four standard Volutome ZEN licenses are active (ZEN core, GeminiSEM, Volutome, Automated Imaging) | 7337
916 | Volutome | 23. Ensure Arivis Software Licence is available (1 year or permanent) | 7338
3321 | Install Customer Licences | Delete the Production Test Licences | 7339
3321 | Install Customer Licences | Copy EMServer.db to the C: Support/Data/Log Files folder | 7340
3321 | Install Customer Licences | Obtain the customer Entitlement Certificate | 7341
3321 | Install Customer Licences | Download into the C:Support/Data/Licence folder | 7342
3321 | Install Customer Licences | Print out the Customer Entitlement Certificate - To be sent to the customer | 7343
3321 | Install Customer Licences | Load customer licences, SmartSEM & ZEN, Including ZEN Promo and check against order they are correct. Please Leave Installed for Shipment. | 7344
3321 | Install Customer Licences | Test customer licenses | 7345
354 | Install Customer Licences | Delete the Production Test Licences | 7346
354 | Install Customer Licences | Obtain the customer Entitlement Certificate | 7347
354 | Install Customer Licences | Download into the Engineer folder | 7348
354 | Install Customer Licences | Install the Customer Licences, SmartSEM & ZEN, Including ZEN Promo. Please Leave Installed for Shipment. | 7349
354 | Install Customer Licences | Copy Contents of Image Folder to the Product Folder | 7350
354 | Install Customer Licences | Copy EMServer.db to the Product Folder | 7351
354 | Install Customer Licences | Print out the Customer Entitlement Certicate - To be sent to the customer | 7352
10 | Mount AC unit | Record AC unit S/N | 7353
10 | Mount AC unit | Qualification Marks Checked | 7354
10 | Mount AC unit | Fit mounting bracket to AC unit 'Angle Bracket' - 349500-8040-000 | 7355
10 | Mount AC unit | Set up mounting bracket height using Jig 356580-0174-000 | 7356
12 | Fit EHT Assembly | Serial No. | 7357
12 | Fit EHT Assembly | Version | 7358
12 | Fit EHT Assembly | Qualification Marks Checked | 7359
19 | Fit DC Unit and Level control module | Fit DC unit. - 2 Person lift or lifting equipment must be used! | 7360
19 | Fit DC Unit and Level control module | Fit level control PCB | 7361
743 | Material Availability | Are there any shortages? | 7362
9 | Assemble Mount and Setup water flow switch | Set water flow pressure to 0.9L / min | 7363
1247 | Install Ion Pump PSU 1 | Fit IPPSU 1, Record S/N | 7364
1247 | Install Ion Pump PSU 1 | Set DIP switches | 7365
1247 | Install Ion Pump PSU 1 | Qualification Marks Checked | 7366
13 | Fit Dynamic EO board - Single Scan (346001-9110-315) | Confirm EO Dynamic PCB PN is '346001-9110-315' | 7367
13 | Fit Dynamic EO board - Single Scan (346001-9110-315) | Fit EO Dynamic PCB (Record SN) | 7368
1218 | Assemble and fit ASAP / SVSEM | Check if ASAP / H-ASAP / SVSEM | 7369
1218 | Assemble and fit ASAP / SVSEM | Record Serial No | 7370
1218 | Assemble and fit ASAP / SVSEM | Fit additional ASAP cards if provided (DFE / VBE Cards) | 7371
17 | Fit Earth Bonding Straps | Fit X4 braided earth straps to rear PCB's. (Vac, EOS, EOD, ASAP /SVSEM) | 7372
17 | Fit Earth Bonding Straps | Fit earth cable to DC unit. | 7373
11253 | SCS Check | Check SCS Folder/Sales Order | 7374
11253 | SCS Check | SCS Number (If applicable) | 7375
11059 | Raith OEM Customer Data Files | Copy- 1KV.Tif (Tiff customer image) | 7376
11059 | Raith OEM Customer Data Files | Copy- 15KV.Tif (Tiff customer image) | 7377
11059 | Raith OEM Customer Data Files | Copy- 30KV.Tif (Tiff customer image) | 7378
11059 | Raith OEM Customer Data Files | Copy- 200V.Tif (Tiff customer image) | 7379
11059 | Raith OEM Customer Data Files | Copy- All Ripple images (Tiff Files) | 7380
11059 | Raith OEM Customer Data Files | Copy- All Ripple images (JPG Files) | 7381
11059 | Raith OEM Customer Data Files | Copy- 9876543.xml (XML Smartsem Licence) | 7382
11059 | Raith OEM Customer Data Files | Copy- Merlin Compact XX-XX S/O (Licence Rep File) | 7383
11059 | Raith OEM Customer Data Files | Copy- EmServer.accdb (Database) | 7384
11059 | Raith OEM Customer Data Files | Copy- EmServer.txt (Server Log File) | 7385
11059 | Raith OEM Customer Data Files | Copy- L500.ini (SmartSEM Config) | 7386
11059 | Raith OEM Customer Data Files | Copy- xx-xx.reg (Export of Registry File) | 7387
11059 | Raith OEM Customer Data Files | Copy- Mybackup.smartbackup (SmartSEM Backup) | 7388
11059 | Raith OEM Customer Data Files | Copy- Logbook (scanned)-PDF | 7389
11059 | Raith OEM Customer Data Files | Copy- DynamicEHT (mn2 logfile) | 7390
11059 | Raith OEM Customer Data Files | Copy- Probe Conditioning Stability (mn2 logfile) | 7391
11059 | Raith OEM Customer Data Files | Copy- 1KV Res.jpg (PATE Result) | 7392
11059 | Raith OEM Customer Data Files | Copy- 1KV Res.Tif (Tiff Res Image) | 7393
11059 | Raith OEM Customer Data Files | Copy- 15KV.jpg (PATE Results) | 7394
11059 | Raith OEM Customer Data Files | Copy- 15KV.Tif (Tiff Res Image) | 7395
11059 | Raith OEM Customer Data Files | Emission-Condenser OFF image from QA (Tiff Image) | 7396
11059 | Raith OEM Customer Data Files | Emission-Condenser ON image from QA (Tiff Image) | 7397
11059 | Raith OEM Customer Data Files | Merlin Compact User Manual (English) (PDF Manual) | 7398
11059 | Raith OEM Customer Data Files | Merlin Compact User Manual (German) (PDF Manual) | 7399
11059 | Raith OEM Customer Data Files | Merlin Compact Assembly (PDF Manual) | 7400
11059 | Raith OEM2 Customer Data Files | Copy- 1KV.Tif (Tiff customer image) | 7401
11059 | Raith OEM2 Customer Data Files | Copy- 15KV.Tif (Tiff customer image) | 7402
11059 | Raith OEM2 Customer Data Files | Copy- 30KV.Tif (Tiff customer image) | 7403
11059 | Raith OEM2 Customer Data Files | Copy- 200V.Tif (Tiff customer image) | 7404
11059 | Raith OEM2 Customer Data Files | Copy- All Ripple images (Tiff Files) | 7405
11059 | Raith OEM2 Customer Data Files | Copy- All Ripple images (JPG Files) | 7406
11059 | Raith OEM2 Customer Data Files | Copy- 9876543.xml (XML Smartsem Licence) | 7407
11059 | Raith OEM2 Customer Data Files | Copy- Merlin Compact XX-XX S/O (Licence Rep File) | 7408
11059 | Raith OEM2 Customer Data Files | Copy- EmServer.accdb (Database) | 7409
11059 | Raith OEM2 Customer Data Files | Copy- EmServer.txt (Server Log File) | 7410
11059 | Raith OEM2 Customer Data Files | Copy- L500.ini (SmartSEM Config) | 7411
11059 | Raith OEM2 Customer Data Files | Copy- xx-xx.reg (Export of Registry File) | 7412
11059 | Raith OEM2 Customer Data Files | Copy- Mybackup.smartbackup (SmartSEM Backup) | 7413
11059 | Raith OEM2 Customer Data Files | Copy- Logbook (scanned)-PDF | 7414
11059 | Raith OEM2 Customer Data Files | Copy- DynamicEHT (mn2 logfile) | 7415
11059 | Raith OEM2 Customer Data Files | Copy- Probe Conditioning Stability (mn2 logfile) | 7416
11059 | Raith OEM2 Customer Data Files | Copy- 1KV Res.jpg (PATE Result) | 7417
11059 | Raith OEM2 Customer Data Files | Copy- 1KV Res.Tif (Tiff Res Image) | 7418
11059 | Raith OEM2 Customer Data Files | Copy- 15KV.jpg (PATE Results) | 7419
11059 | Raith OEM2 Customer Data Files | Copy- 15KV.Tif (Tiff Res Image) | 7420
11059 | Raith OEM2 Customer Data Files | Emission-Condenser OFF image from QA (Tiff Image) | 7421
11059 | Raith OEM2 Customer Data Files | Emission-Condenser ON image from QA (Tiff Image) | 7422
11059 | Raith OEM2 Customer Data Files | Merlin Compact User Manual (English) (PDF Manual) | 7423
11059 | Raith OEM2 Customer Data Files | Merlin Compact User Manual (German) (PDF Manual) | 7424
11059 | Raith OEM2 Customer Data Files | Merlin Compact Assembly (PDF Manual) | 7425
1510 | Raith OEM7 Customer Data Files | Copy- 1KV.Tif (Tiff customer image) | 7426
1510 | Raith OEM7 Customer Data Files | Copy- 15KV.Tif (Tiff customer image) | 7427
1510 | Raith OEM7 Customer Data Files | Copy- 30KV.Tif (Tiff customer image) | 7428
1510 | Raith OEM7 Customer Data Files | Copy- 200V.Tif (Tiff customer image) | 7429
1510 | Raith OEM7 Customer Data Files | Copy- All Ripple images (Tiff Files) | 7430
1510 | Raith OEM7 Customer Data Files | Copy- All Ripple images (JPG Files) | 7431
1510 | Raith OEM7 Customer Data Files | Copy- 9876543.xml (XML Smartsem Licence) | 7432
1510 | Raith OEM7 Customer Data Files | Copy- Merlin Compact XX-XX S/O (Licence Rep File) | 7433
1510 | Raith OEM7 Customer Data Files | Copy- EmServer.accdb (Database) | 7434
1510 | Raith OEM7 Customer Data Files | Copy- EmServer.txt (Server Log File) | 7435
1510 | Raith OEM7 Customer Data Files | Copy- L500.ini (SmartSEM Config) | 7436
1510 | Raith OEM7 Customer Data Files | Copy- xx-xx.reg (Export of Registry File) | 7437
1510 | Raith OEM7 Customer Data Files | Copy- Mybackup.smartbackup (SmartSEM Backup) | 7438
1510 | Raith OEM7 Customer Data Files | Copy- Logbook (scanned)-PDF | 7439
1510 | Raith OEM7 Customer Data Files | Copy- DynamicEHT (mn2 logfile) | 7440
1510 | Raith OEM7 Customer Data Files | Copy- Probe Conditioning Stability (mn2 logfile) | 7441
1510 | Raith OEM7 Customer Data Files | Copy- 1KV Res.jpg (PATE Result) | 7442
1510 | Raith OEM7 Customer Data Files | Copy- 1KV Res.Tif (Tiff Res Image) | 7443
1510 | Raith OEM7 Customer Data Files | Copy- 15KV.jpg (PATE Results) | 7444
1510 | Raith OEM7 Customer Data Files | Copy- 15KV.Tif (Tiff Res Image) | 7445
1510 | Raith OEM7 Customer Data Files | Emission-Condenser OFF image from QA (Tiff Image) | 7446
1510 | Raith OEM7 Customer Data Files | Emission-Condenser ON image from QA (Tiff Image) | 7447
1510 | Raith OEM7 Customer Data Files | Merlin Compact User Manual (English) (PDF Manual) | 7448
1510 | Raith OEM7 Customer Data Files | Merlin Compact User Manual (German) (PDF Manual) | 7449
1510 | Raith OEM7 Customer Data Files | Merlin Compact Assembly (PDF Manual) | 7450
1510 | Raith OEM7 Customer Data Files | Copy- 1KV.Tif (Tiff customer image) | 7451
1510 | Raith OEM7 Customer Data Files | Copy- 15KV.Tif (Tiff customer image) | 7452
1510 | Raith OEM7 Customer Data Files | Copy- 30KV.Tif (Tiff customer image) | 7453
1510 | Raith OEM7 Customer Data Files | Copy- 200V.Tif (Tiff customer image) | 7454
1510 | Raith OEM7 Customer Data Files | Copy- All Ripple images (Tiff Files) | 7455
1510 | Raith OEM7 Customer Data Files | Copy- All Ripple images (JPG Files) | 7456
1510 | Raith OEM7 Customer Data Files | Copy- 9876543.xml (XML Smartsem Licence) | 7457
1510 | Raith OEM7 Customer Data Files | Copy- Merlin Compact XX-XX S/O (Licence Rep File) | 7458
1510 | Raith OEM7 Customer Data Files | Copy- EmServer.accdb (Database) | 7459
1510 | Raith OEM7 Customer Data Files | Copy- EmServer.txt (Server Log File) | 7460
1510 | Raith OEM7 Customer Data Files | Copy- L500.ini (SmartSEM Config) | 7461
1510 | Raith OEM7 Customer Data Files | Copy- xx-xx.reg (Export of Registry File) | 7462
1510 | Raith OEM7 Customer Data Files | Copy- Mybackup.smartbackup (SmartSEM Backup) | 7463
1510 | Raith OEM7 Customer Data Files | Copy- Logbook (scanned)-PDF | 7464
1510 | Raith OEM7 Customer Data Files | Copy- DynamicEHT (mn2 logfile) | 7465
1510 | Raith OEM7 Customer Data Files | Copy- Probe Conditioning Stability (mn2 logfile) | 7466
1510 | Raith OEM7 Customer Data Files | Copy- 1KV Res.jpg (PATE Result) | 7467
1510 | Raith OEM7 Customer Data Files | Copy- 1KV Res.Tif (Tiff Res Image) | 7468
1510 | Raith OEM7 Customer Data Files | Copy- 15KV.jpg (PATE Results) | 7469
1510 | Raith OEM7 Customer Data Files | Copy- 15KV.Tif (Tiff Res Image) | 7470
1510 | Raith OEM7 Customer Data Files | Emission-Condenser OFF image from QA (Tiff Image) | 7471
1510 | Raith OEM7 Customer Data Files | Emission-Condenser ON image from QA (Tiff Image) | 7472
1510 | Raith OEM7 Customer Data Files | Merlin Compact User Manual (English) (PDF Manual) | 7473
1510 | Raith OEM7 Customer Data Files | Merlin Compact User Manual (German) (PDF Manual) | 7474
1510 | Raith OEM7 Customer Data Files | Merlin Compact Assembly (PDF Manual) | 7475
1510 | Raith OEM7 Customer Data Files | Copy- 1KV.Tif (Tiff customer image) | 7476
1510 | Raith OEM7 Customer Data Files | Copy- 15KV.Tif (Tiff customer image) | 7477
1510 | Raith OEM7 Customer Data Files | Copy- 30KV.Tif (Tiff customer image) | 7478
1510 | Raith OEM7 Customer Data Files | Copy- 200V.Tif (Tiff customer image) | 7479
1510 | Raith OEM7 Customer Data Files | Copy- All Ripple images (Tiff Files) | 7480
1510 | Raith OEM7 Customer Data Files | Copy- All Ripple images (JPG Files) | 7481
1510 | Raith OEM7 Customer Data Files | Copy- 9876543.xml (XML Smartsem Licence) | 7482
1510 | Raith OEM7 Customer Data Files | Copy- Merlin Compact XX-XX S/O (Licence Rep File) | 7483
1510 | Raith OEM7 Customer Data Files | Copy- EmServer.accdb (Database) | 7484
1510 | Raith OEM7 Customer Data Files | Copy- EmServer.txt (Server Log File) | 7485
1510 | Raith OEM7 Customer Data Files | Copy- L500.ini (SmartSEM Config) | 7486
1510 | Raith OEM7 Customer Data Files | Copy- xx-xx.reg (Export of Registry File) | 7487
1510 | Raith OEM7 Customer Data Files | Copy- Mybackup.smartbackup (SmartSEM Backup) | 7488
1510 | Raith OEM7 Customer Data Files | Copy- Logbook (scanned)-PDF | 7489
1510 | Raith OEM7 Customer Data Files | Copy- DynamicEHT (mn2 logfile) | 7490
1510 | Raith OEM7 Customer Data Files | Copy- Probe Conditioning Stability (mn2 logfile) | 7491
1510 | Raith OEM7 Customer Data Files | Copy- 1KV Res.jpg (PATE Result) | 7492
1510 | Raith OEM7 Customer Data Files | Copy- 1KV Res.Tif (Tiff Res Image) | 7493
1510 | Raith OEM7 Customer Data Files | Copy- 15KV.jpg (PATE Results) | 7494
1510 | Raith OEM7 Customer Data Files | Copy- 15KV.Tif (Tiff Res Image) | 7495
1510 | Raith OEM7 Customer Data Files | Emission-Condenser OFF image from QA (Tiff Image) | 7496
1510 | Raith OEM7 Customer Data Files | Emission-Condenser ON image from QA (Tiff Image) | 7497
1510 | Raith OEM7 Customer Data Files | Merlin Compact User Manual (English) (PDF Manual) | 7498
1510 | Raith OEM7 Customer Data Files | Merlin Compact User Manual (German) (PDF Manual) | 7499
1510 | Raith OEM7 Customer Data Files | Merlin Compact Assembly (PDF Manual) | 7500
1204 | Heater HiPot Tests HV | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.5Ω) | 7501
1204 | Heater HiPot Tests HV | IGP Heaters (IGP#1) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 7502
1204 | Heater HiPot Tests HV | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 7503
1204 | Heater HiPot Tests HV | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 7504
11384 | Leak check - Vacuum Test | Warning -ONLY VENT using the electronic vent valve in the chamber | 7505
11384 | Leak check - Vacuum Test | No Leaks Detected (Prior to running up Turbo) | 7506
11384 | Leak check - Vacuum Test | No Leaks Detected (With Turbo running) when in high sensitivity | 7507
1443 | PC Power Up and Enter Bios | Check That AC Loss Policy Depend On, is Set To Power On | 7508
1443 | PC Power Up and Enter Bios | Disable Windows Fast Start UP | 7509
1443 | PC Power Up and Enter Bios | Disable USB Power Save | 7510
1443 | PC Power Up and Enter Bios | Check PC Part Number | 7511
1443 | PC Power Up and Enter Bios | Check Windows 11 Certification label is fitted | 7512
252 | Configure PC ASAP | PC S/N | 7513
252 | Configure PC ASAP | Hard disk Image (Version) | 7514
252 | Configure PC ASAP | Record Operating System | 7515
252 | Configure PC ASAP | Set the Time and Date on the PC to current UK | 7516
252 | Configure PC ASAP | Check the Entitlement certificate to determine which version of Smartsem to install | 7517
252 | Configure PC ASAP | Install Smart SEM (See WI 9693) Record Version | 7518
252 | Configure PC ASAP | Check for available SmartSEMService Packs and install them. (N/A if none) | 7519
252 | Configure PC ASAP | Check for available SmartSEM hotfixes and install them. (N/A if none) | 7520
252 | Configure PC ASAP | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 7521
252 | Configure PC ASAP | Install CZ Support Tools. Ver | 7522
252 | Configure PC ASAP | Disable Automatic Updates | 7523
252 | Configure PC ASAP | Run ASAP Screening Program | 7524
252 | Configure PC ASAP | System Specific Datazone Loaded | 7525
252 | Configure PC ASAP | In SEM status set 'Mains Sync' to off | 7526
1476 | Configure PC ASAP (Crossbeam 750) | PC S/N | 7527
1476 | Configure PC ASAP (Crossbeam 750) | Hard disk Image (Version) | 7528
1476 | Configure PC ASAP (Crossbeam 750) | Record Operating System | 7529
1476 | Configure PC ASAP (Crossbeam 750) | Set the Time and Date on the PC to current UK | 7530
1476 | Configure PC ASAP (Crossbeam 750) | Check the Entitlement certificate to determine which version of Smartsem to install | 7531
1476 | Configure PC ASAP (Crossbeam 750) | Install Smart SEM - This system has a Gemini 4 column and must be installled as a Crossbeam 750 | 7532
1476 | Configure PC ASAP (Crossbeam 750) | Confirm SmartSEM is configured as Crossbeam 750 | 7533
1476 | Configure PC ASAP (Crossbeam 750) | Check for available SmartSEMService Packs and install them. (N/A if none) | 7534
1476 | Configure PC ASAP (Crossbeam 750) | Check for available SmartSEM hotfixes and install them. (N/A if none) | 7535
1476 | Configure PC ASAP (Crossbeam 750) | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 7536
1476 | Configure PC ASAP (Crossbeam 750) | Install CZ Support Tools. Ver | 7537
1476 | Configure PC ASAP (Crossbeam 750) | Disable Automatic Updates | 7538
1476 | Configure PC ASAP (Crossbeam 750) | Run ASAP Screening Program | 7539
1476 | Configure PC ASAP (Crossbeam 750) | System Specific Datazone Loaded | 7540
1476 | Configure PC ASAP (Crossbeam 750) | In SEM status set 'Mains Sync' to off | 7541
1483 | Configure PC ASAP (Crossbeam 330) | PC S/N | 7542
1483 | Configure PC ASAP (Crossbeam 330) | Hard disk Image (Version) | 7543
1483 | Configure PC ASAP (Crossbeam 330) | Record Operating System | 7544
1483 | Configure PC ASAP (Crossbeam 330) | Set the Time and Date on the PC to current UK | 7545
1483 | Configure PC ASAP (Crossbeam 330) | Check the Entitlement certificate to determine which version of Smartsem to install | 7546
1483 | Configure PC ASAP (Crossbeam 330) | Install Smart SEM - This system should be configured as a Crossbeam 330 (Early systems will use Beta version of Smartsem V8.04) | 7547
1483 | Configure PC ASAP (Crossbeam 330) | Confirm SmartSEM is configured as Crossbeam 330 | 7548
1483 | Configure PC ASAP (Crossbeam 330) | Check for available SmartSEMService Packs and install them. (N/A if none) | 7549
1483 | Configure PC ASAP (Crossbeam 330) | Check for available SmartSEM hotfixes and install them. (N/A if none) | 7550
1483 | Configure PC ASAP (Crossbeam 330) | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 7551
1483 | Configure PC ASAP (Crossbeam 330) | Install CZ Support Tools. Ver | 7552
1483 | Configure PC ASAP (Crossbeam 330) | Disable Automatic Updates | 7553
1483 | Configure PC ASAP (Crossbeam 330) | Run ASAP Screening Program | 7554
1483 | Configure PC ASAP (Crossbeam 330) | System Specific Datazone Loaded | 7555
1483 | Configure PC ASAP (Crossbeam 330) | In SEM status set 'Mains Sync' to off | 7556
1443 | PC Power Up and Enter Bios | Set Dash Client service to Manual | 7557
1443 | PC Power Up and Enter Bios | Disable USB Power Save | 7558
1443 | PC Power Up and Enter Bios | Check PC Part Number | 7559
1443 | PC Power Up and Enter Bios | Check Windows 11 Certification label is fitted | 7560
1443 | PC Power Up and Enter Bios | Set Dash Client service to Manual | 7561
1443 | PC Power Up and Enter Bios | Disable USB Power Save | 7562
1443 | PC Power Up and Enter Bios | Check PC Part Number | 7563
1443 | PC Power Up and Enter Bios | Check Windows 11 Certification label is fitted | 7564
1443 | PC Power Up and Enter Bios | Set Dash Client service to Manual | 7565
1443 | PC Power Up and Enter Bios | Disable USB Power Save | 7566
1443 | PC Power Up and Enter Bios | Check PC Part Number | 7567
1443 | PC Power Up and Enter Bios | Check Windows 11 Certification label is fitted | 7568
1255 | Label up cables | In-Lens And ESB cable | 7569
1255 | Label up cables | IGP cables (IGP1 & IGP2) | 7570
1255 | Label up cables | SCM cable | 7571
1255 | Label up cables | CCD camera cables | 7572
58 | Accesories Configuration | Port configuration Check | 7573
1259 | Route all Chamber / Stage / Detector Cables and pipes | Configure ASAP / SVSEM detectors ports as specified in procedure. | 7574
1161 | SCS Sign off | Does system have an SCS? please select | 7575
1161 | SCS Sign off | If yes, Product Manager to enter initials. If no please type N/A: | 7576
11274 | Fit Column Cables + Pipes | Fit and route all column cables | 7577
11274 | Fit Column Cables + Pipes | Confirm IGP cable locking screws are secured and grub screw tightened | 7578
11274 | Fit Column Cables + Pipes | Fit and route objective water pipes | 7579
11274 | Fit Column Cables + Pipes | Safety – Torque the earth stud to 15 N.m | 7580
11274 | Fit Column Cables + Pipes | Torque both earth zinc plated M8 nuts to 15 N.m | 7581
11274 | Fit Column Cables + Pipes | Confirm HV cable is routed though plinth, and secured on top of plinth | 7582
1252 | Chamber checks | Check interlock keys fitted to door | 7583
1252 | Chamber checks | Check vented screws fitted to inside of chamber ceiling | 7584
1252 | Chamber checks | Check magnet clamp set correctly | 7585
1252 | Chamber checks | Check door alignment | 7586
1236 | Print and Fit SN Label (NO CSA, NO SEMI) | ✘ 𝗡𝗢 CSA marking on SN label.. | 7587
1236 | Print and Fit SN Label (NO CSA, NO SEMI) | ✘ 𝗡𝗢 SEMI marking on SN label. | 7588
11166 | Anode_TS_55um 40nA, 55um Centre Hole | Measure aperture size, 55um (+/- 2um) | 7589
11166 | Anode_TS_55um 40nA, 55um Centre Hole | Anode_TS_55um - Passed inspection criteria, Image Saved | 7590
11319 | Anode_TS_HRQA 40nA, Surface Overview and Concentricity | Anode_TS_HRQA - Passed inspection criteria, Image Saved | 7591
11334 | Anode_US_HRQA 40nA, Surface Overview | Anode_US_HRQA - Passed inspection criteria, Image Saved | 7592
11174 | Plat_TS_750um_HRQA, Surface Overview and 750um Centre Hole | Measure aperture size, 750um (+/- 15um) | 7593
11174 | Plat_TS_750um_HRQA, Surface Overview and 750um Centre Hole | Plat_TS_750um_HRQA - Passed inspection criteria, Image Saved | 7594
11173 | SF_TS_35um – SINGLE FILM APERTURE (SFA) Topside | Measure aperture size, 35um (+/- 2um) | 7595
758 | SF_TS_HRQA, Surface Overview and Concentricity | SF_TS_HRQA - Passed inspection criteria, Image Saved | 7596
11144 | Fit Water Pipes | Vacuum Check (-15PSI @ 5 Mins) | 7597
11328 | Anode_TS_Bore1 40nA, First Bore Shot | Anode_TS_Bore1 - Passed inspection criteria, Image Saved | 7598
11331 | Anode_TS_Bore2 40nA, Second Bore Shot | Anode_TS_Bore2 - Passed inspection criteria, Image Saved | 7599
11163 | Anode_TS_Bore1 100nA, First Bore Shot | Anode_TS_Bore1 - Passed inspection criteria, Image Saved | 7600
11164 | Anode_TS_Bore2 100nA, Second Bore Shot | Anode_TS_Bore2 - Passed inspection criteria, Image Saved | 7601
11167 | Anode_TS_90um 100nA, 90um Centre Hole | Measure aperture size, 90um (+/- 5um) | 7602
11167 | Anode_TS_90um 100nA, 90um Centre Hole | Anode_TS_90um - Passed inspection criteria, Image Saved | 7603
11169 | Anode_US_HRQA 100nA, Surface Overview | Anode_US_HRQA - Passed inspection criteria, Image Saved | 7604
11354 | Thermocouple Test | UHV Thermocouple resistance is 110Ὡ (+/-10Ὡ) | 7605
11266 | Bakeout cable routing | Flange Heater Earths routed correctly (where applicable) | 7606
11266 | Bakeout cable routing | Thermocouple cables routed correctly | 7607
1443 | PC Power Up and Enter Bios | Set Dash Client service to Manual | 7608
1443 | PC Power Up and Enter Bios | Disable USB Power Save | 7609
1443 | PC Power Up and Enter Bios | Check PC Part Number | 7610
1443 | PC Power Up and Enter Bios | Check Windows 11 Certification label is fitted | 7611
1443 | PC Power Up and Enter Bios | Set Dash Client service to Manual | 7612
1443 | PC Power Up and Enter Bios | Disable USB Power Save | 7613
1443 | PC Power Up and Enter Bios | Check PC Part Number | 7614
1443 | PC Power Up and Enter Bios | Check Windows 11 Certification label is fitted | 7615
1443 | PC Power Up and Enter Bios | Set Dash Client service to Manual | 7616
1443 | PC Power Up and Enter Bios | Disable USB Power Save | 7617
1443 | PC Power Up and Enter Bios | Check PC Part Number | 7618
1443 | PC Power Up and Enter Bios | Check Windows 11 Certification label is fitted | 7619
1476 | Configure PC ASAP (Crossbeam 750) | PC S/N | 7620
1476 | Configure PC ASAP (Crossbeam 750) | Hard disk Image (Version) | 7621
1476 | Configure PC ASAP (Crossbeam 750) | Record Operating System | 7622
1476 | Configure PC ASAP (Crossbeam 750) | Set the Time and Date on the PC to current UK | 7623
1476 | Configure PC ASAP (Crossbeam 750) | Check the Entitlement certificate to determine which version of Smartsem to install | 7624
1476 | Configure PC ASAP (Crossbeam 750) | Install Smart SEM - This system has a Gemini 4 column and must be installled as a Crossbeam 750 | 7625
1476 | Configure PC ASAP (Crossbeam 750) | Confirm SmartSEM is configured as Crossbeam 750 | 7626
1476 | Configure PC ASAP (Crossbeam 750) | Check for available SmartSEMService Packs and install them. (N/A if none) | 7627
1476 | Configure PC ASAP (Crossbeam 750) | Check for available SmartSEM hotfixes and install them. (N/A if none) | 7628
1476 | Configure PC ASAP (Crossbeam 750) | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 7629
1476 | Configure PC ASAP (Crossbeam 750) | Install CZ Support Tools. Ver | 7630
1476 | Configure PC ASAP (Crossbeam 750) | Run ASAP Screening Program | 7631
1476 | Configure PC ASAP (Crossbeam 750) | System Specific Datazone Loaded | 7632
1476 | Configure PC ASAP (Crossbeam 750) | In SEM status set 'Mains Sync' to off | 7633
1483 | Configure PC ASAP (Crossbeam 330) | PC S/N | 7634
1483 | Configure PC ASAP (Crossbeam 330) | Hard disk Image (Version) | 7635
1483 | Configure PC ASAP (Crossbeam 330) | Record Operating System | 7636
1483 | Configure PC ASAP (Crossbeam 330) | Set the Time and Date on the PC to current UK | 7637
1483 | Configure PC ASAP (Crossbeam 330) | Check the Entitlement certificate to determine which version of Smartsem to install | 7638
1483 | Configure PC ASAP (Crossbeam 330) | Install Smart SEM - This system should be configured as a Crossbeam 330 (Early systems will use Beta version of Smartsem V8.04) | 7639
1483 | Configure PC ASAP (Crossbeam 330) | Confirm SmartSEM is configured as Crossbeam 330 | 7640
1483 | Configure PC ASAP (Crossbeam 330) | Check for available SmartSEMService Packs and install them. (N/A if none) | 7641
1483 | Configure PC ASAP (Crossbeam 330) | Check for available SmartSEM hotfixes and install them. (N/A if none) | 7642
1483 | Configure PC ASAP (Crossbeam 330) | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 7643
1483 | Configure PC ASAP (Crossbeam 330) | Install CZ Support Tools. Ver | 7644
1483 | Configure PC ASAP (Crossbeam 330) | Run ASAP Screening Program | 7645
1483 | Configure PC ASAP (Crossbeam 330) | System Specific Datazone Loaded | 7646
1483 | Configure PC ASAP (Crossbeam 330) | In SEM status set 'Mains Sync' to off | 7647
252 | Configure PC ASAP | PC S/N | 7648
252 | Configure PC ASAP | Hard disk Image (Version) | 7649
252 | Configure PC ASAP | Record Operating System | 7650
252 | Configure PC ASAP | Set the Time and Date on the PC to current UK | 7651
252 | Configure PC ASAP | Check the Entitlement certificate to determine which version of Smartsem to install | 7652
252 | Configure PC ASAP | Install Smart SEM (See WI 9693) Record Version | 7653
252 | Configure PC ASAP | Check for available SmartSEMService Packs and install them. (N/A if none) | 7654
252 | Configure PC ASAP | Check for available SmartSEM hotfixes and install them. (N/A if none) | 7655
252 | Configure PC ASAP | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 7656
252 | Configure PC ASAP | Install CZ Support Tools. Ver | 7657
252 | Configure PC ASAP | Run ASAP Screening Program | 7658
252 | Configure PC ASAP | System Specific Datazone Loaded | 7659
252 | Configure PC ASAP | In SEM status set 'Mains Sync' to off | 7660
1443 | PC Power Up and Enter Bios | Set Dash Client service to Manual | 7661
1443 | PC Power Up and Enter Bios | Disable USB Power Save | 7662
1443 | PC Power Up and Enter Bios | Check PC Part Number | 7663
1443 | PC Power Up and Enter Bios | Check Windows 11 Certification label is fitted | 7664
1443 | PC Power Up and Enter Bios | Set Dash Client service to Manual | 7665
1443 | PC Power Up and Enter Bios | Disable USB Power Save | 7666
1443 | PC Power Up and Enter Bios | Check PC Part Number | 7667
1443 | PC Power Up and Enter Bios | Check Windows 11 Certification label is fitted | 7668
1443 | PC Power Up and Enter Bios | Set Dash Client service to Manual | 7669
1443 | PC Power Up and Enter Bios | Disable USB Power Save | 7670
1443 | PC Power Up and Enter Bios | Check PC Part Number | 7671
1443 | PC Power Up and Enter Bios | Check Windows 11 Certification label is fitted | 7672
10529 | BSD/BSD1 | BSD Diode Pre-test | 7673
10529 | BSD/BSD1 | Select Type | 7674
10529 | BSD/BSD1 | Select Amplifier Type | 7675
10529 | BSD/BSD1 | ES-BSD Only - Using ASAP Test Panel Update the ID to 0x0150 | 7676
10529 | BSD/BSD1 | Black Level Calibration | 7677
10529 | BSD/BSD1 | Check Diode Segment Function | 7678
10529 | BSD/BSD1 | Check All Gain ranges | 7679
10529 | BSD/BSD1 | Check Auto range | 7680
10529 | BSD/BSD1 | Check Auto Level (Does Not Function With BSD1) | 7681
10529 | BSD/BSD1 | Setup and Check TOPO | 7682
10529 | BSD/BSD1 | Take 3 images - Atomic, 5pA, 2KV | 7683
10529 | BSD/BSD1 | Store BSD Functional Images-Atomic, 5pA, 2KV | 7684
11237 | Ensure Open Anode (Non EBBU) | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed | 7685
11236 | Ensure Open Anode (EBBU) | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed. | 7686
11264 | Ensure Closed Anode (EBBU) | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be observed. | 7687
11264 | Ensure Closed Anode (EBBU) | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be observed. | 7688
11265 | EBBU Anode | Make sure there are 3 cut-outs on the edge of the anode. No particles are to be observed. | 7689
11263 | EBBU Anode | Make sure there are 3 cut-outs on the edge of the anode. No particles are to be observed | 7690
11012 | EBBP | Fit Test tool BB Electronics | 7691
11012 | EBBP | Functional Test completed | 7692
11012 | EBBP | Remove Test tool BB Electronics | 7693
504 | Closed Multi Hole aperture holder observations | Select PC | 7694
504 | Closed Multi Hole aperture holder observations | Hold aperture holder, look for holes on the sides.No particles are to be observed | 7695
504 | Closed Multi Hole aperture holder observations | Select PC | 7696
504 | Closed Multi Hole aperture holder observations | Hold aperture holder, look for holes on the sides.No particles are to be observed | 7697
505 | Ventilated Multi Open Hole aperture holder observations | Hold aperture holder, look for holes on the sides.No particles are to be observed | 7698
11187 | Multi_TS_HRQA 100nA, Surface Overview and Concentricity | Check there is no engraved ring around the cluster of apertures, this should be on the underside. | 7699
11187 | Multi_TS_HRQA 100nA, Surface Overview and Concentricity | Multi_TS_HRQA - Passed inspection criteria, Image Saved | 7700
11277 | Multi_TS_HRQA 100nA, Surface Overview and Concentricity | Check there is no engraved ring around the cluster of apertures, this should be on the underside. | 7701
11277 | Multi_TS_HRQA 100nA, Surface Overview and Concentricity | Multi_TS_HRQA - Passed inspection criteria, Image Saved | 7702
11011 | EBBU | Fit Customer BB Electronics (Do not fit Test Tool BB Electronics to this system) | 7703
11011 | EBBU | Functional Tests complete | 7704
10894 | Beam Deceleration | Check That Beam Deceleration Can Be Switched On., | 7705
10894 | Beam Deceleration | Beam De-Cel Sample Holder | 7706
11277 | Multi_TS_HRQA 100nA, Surface Overview and Concentricity | Check there is no engraved ring around the cluster of apertures, this should be on the underside. | 7707
11277 | Multi_TS_HRQA 100nA, Surface Overview and Concentricity | Multi_TS_HRQA - Passed inspection criteria, Image Saved | 7708
11283 | Multi_TS_HRQA 20nA, Surface Overview and Concentricity | Check there is no engraved ring around the cluster of apertures, this should be on the underside. | 7709
11283 | Multi_TS_HRQA 20nA, Surface Overview and Concentricity | Multi_TS_HRQA - Passed inspection criteria, Image Saved | 7710
11284 | Multi_TS_HRQA 20nA, Surface Overview and Concentricity | Check there is no engraved ring around the cluster of apertures, this should be on the underside. | 7711
11284 | Multi_TS_HRQA 20nA, Surface Overview and Concentricity | Multi_TS_HRQA - Passed inspection criteria, Image Saved | 7712
11252 | Multi_TS_30um 100nA, 30um Centre Hole | Measure aperture size, 30um (+/- 2um) | 7713
11252 | Multi_TS_30um 100nA, 30um Centre Hole | Multi_TS_30um - Passed inspection criteria, Image Saved | 7714
11285 | Multi_TS_30um 20nA, 30um Centre Hole | Measure aperture size, 30um (+/- 2um) | 7715
11285 | Multi_TS_30um 20nA, 30um Centre Hole | Multi_TS_30um - Passed inspection criteria, Image Saved | 7716
11156 | Multi_TS_20um 100nA, 20um Centre Hole | Measure aperture size, 20um (+/- 2um) | 7717
11156 | Multi_TS_20um 100nA, 20um Centre Hole | Multi_TS_20um - Passed inspection criteria, Image Saved | 7718
11286 | Multi_TS_20um 20nA, 20um Centre Hole | Measure aperture size, 20um (+/- 2um) | 7719
11286 | Multi_TS_20um 20nA, 20um Centre Hole | Multi_TS_20um - Passed inspection criteria, Image Saved | 7720
11160 | Multi_TS_120um 20nA, 120um Large Aperture | Measure aperture size, 120um (+/- 2um) | 7721
11160 | Multi_TS_120um 20nA, 120um Large Aperture | Multi_TS_120um - Passed inspection criteria, Image Saved | 7722
11287 | Multi_TS_120um 20nA, 120um Large Aperture | Measure aperture size, 120um (+/- 2um) | 7723
11287 | Multi_TS_120um 20nA, 120um Large Aperture | Multi_TS_120um - Passed inspection criteria, Image Saved | 7724
11288 | Multi_TS_300um 100nA, 300um Large Aperture | Measure aperture size, 300um (+/- 10um) | 7725
11288 | Multi_TS_300um 100nA, 300um Large Aperture | Multi_TS_300um - Passed inspection criteria, Image Saved | 7726
11159 | Multi_TS_300um 100nA, 300um Large Aperture | Measure aperture size, 300um (+/- 10um) | 7727
11159 | Multi_TS_300um 100nA, 300um Large Aperture | Multi_TS_300um - Passed inspection criteria, Image Saved | 7728
11161 | Multi_US_HRQA 100nA, Surface Overview | Check there is an engraved ring around the cluster of apertures, this should be on the underside | 7729
11161 | Multi_US_HRQA 100nA, Surface Overview | Multi_US_HRQA - Passed inspection criteria, Image Saved | 7730
11289 | Multi_US_HRQA 100nA, Surface Overview | Check there is an engraved ring around the cluster of apertures, this should be on the underside | 7731
11289 | Multi_US_HRQA 100nA, Surface Overview | Multi_US_HRQA - Passed inspection criteria, Image Saved | 7732
11290 | Multi_US_HRQA 20nA, Surface Overview 30um Centre Hole | Check there is an engraved ring around the cluster of apertures, this should be on the underside | 7733
11290 | Multi_US_HRQA 20nA, Surface Overview 30um Centre Hole | Multi_US_HRQA - Passed inspection criteria, Image Saved | 7734
11308 | Multi_US_HRQA 20nA, Surface Overview 20um Centre Hole | Check there is an engraved ring around the cluster of apertures, this should be on the underside | 7735
11308 | Multi_US_HRQA 20nA, Surface Overview 20um Centre Hole | Multi_US_HRQA - Passed inspection criteria, Image Saved | 7736
11262 | Anode_TS_HRQA 100nA, Surface Overview and Concentricity | Anode_TS_HRQA - Passed inspection criteria, Image Saved | 7737
11317 | Anode_TS_HRQA 300nA, Surface Overview and Concentricity | Anode_TS_HRQA - Passed inspection criteria, Image Saved | 7738
11318 | Anode_TS_HRQA 20nA, Surface Overview and Concentricity | Anode_TS_HRQA - Passed inspection criteria, Image Saved | 7739
11319 | Anode_TS_HRQA 40nA, Surface Overview and Concentricity | Anode_TS_HRQA - Passed inspection criteria, Image Saved | 7740
11163 | Anode_TS_Bore1 100nA, First Bore Shot | Anode_TS_Bore1 - Passed inspection criteria, Image Saved | 7741
11326 | Anode_TS_Bore1 300nA, First Bore Shot | Anode_TS_Bore1 - Passed inspection criteria, Image Saved | 7742
11327 | Anode_TS_Bore1 20nA, First Bore Shot | Anode_TS_Bore1 - Passed inspection criteria, Image Saved | 7743
11328 | Anode_TS_Bore1 40nA, First Bore Shot | Anode_TS_Bore1 - Passed inspection criteria, Image Saved | 7744
11164 | Anode_TS_Bore2 100nA, Second Bore Shot | Anode_TS_Bore2 - Passed inspection criteria, Image Saved | 7745
11329 | Anode_TS_Bore2 300nA, Second Bore Shot | Anode_TS_Bore2 - Passed inspection criteria, Image Saved | 7746
11330 | Anode_TS_Bore2 20nA, Second Bore Shot | Anode_TS_Bore2 - Passed inspection criteria, Image Saved | 7747
11331 | Anode_TS_Bore2 40nA, Second Bore Shot | Anode_TS_Bore2 - Passed inspection criteria, Image Saved | 7748
11165 | Anode_TS_40um 20nA, 40um Centre Hole | Measure aperture size, 40um (+/- 2um) | 7749
11165 | Anode_TS_40um 20nA, 40um Centre Hole | Anode_TS_40um - Passed inspection criteria, Image Saved | 7750
11166 | Anode_TS_55um 40nA, 55um Centre Hole | Measure aperture size, 55um (+/- 2um) | 7751
11166 | Anode_TS_55um 40nA, 55um Centre Hole | Anode_TS_55um - Passed inspection criteria, Image Saved | 7752
1476 | Configure PC ASAP (Crossbeam 750) | PC S/N | 7753
1476 | Configure PC ASAP (Crossbeam 750) | Hard disk Image (Version) | 7754
1476 | Configure PC ASAP (Crossbeam 750) | Record Operating System | 7755
1476 | Configure PC ASAP (Crossbeam 750) | Set the Time and Date on the PC to current UK | 7756
1476 | Configure PC ASAP (Crossbeam 750) | Check the Entitlement certificate to determine which version of Smartsem to install | 7757
1476 | Configure PC ASAP (Crossbeam 750) | Install Smart SEM - This system has a Gemini 4 column and must be installled as a Crossbeam 750 | 7758
1476 | Configure PC ASAP (Crossbeam 750) | Confirm SmartSEM is configured as Crossbeam 750 | 7759
1476 | Configure PC ASAP (Crossbeam 750) | Check for available SmartSEMService Packs and install them. (N/A if none) | 7760
1476 | Configure PC ASAP (Crossbeam 750) | Check for available SmartSEM hotfixes and install them. (N/A if none) | 7761
1476 | Configure PC ASAP (Crossbeam 750) | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 7762
1476 | Configure PC ASAP (Crossbeam 750) | Install CZ Support Tools. Ver | 7763
1476 | Configure PC ASAP (Crossbeam 750) | Run ASAP Screening Program | 7764
1476 | Configure PC ASAP (Crossbeam 750) | System Specific Datazone Loaded | 7765
1476 | Configure PC ASAP (Crossbeam 750) | In SEM status set 'Mains Sync' to off | 7766
11167 | Anode_TS_90um 100nA, 90um Centre Hole | Measure aperture size, 90um (+/- 5um) | 7767
11167 | Anode_TS_90um 100nA, 90um Centre Hole | Anode_TS_90um - Passed inspection criteria, Image Saved | 7768
11167 | Anode_TS_90um 100nA, 90um Centre Hole | Measure aperture size, 90um (+/- 5um) | 7769
11167 | Anode_TS_90um 100nA, 90um Centre Hole | Anode_TS_90um - Passed inspection criteria, Image Saved | 7770
11168 | Anode_TS_150um 300nA, 150um Centre Hole | Measure aperture size, 150um (+/- 5um) | 7771
11168 | Anode_TS_150um 300nA, 150um Centre Hole | Anode_TS_90um - Passed inspection criteria, Image Saved | 7772
11169 | Anode_US_HRQA 100nA, Surface Overview | Anode_US_HRQA - Passed inspection criteria, Image Saved | 7773
11332 | Anode_US_HRQA 300nA, Surface Overview | Anode_US_HRQA - Passed inspection criteria, Image Saved | 7774
1512 | Objective Shift Limit (XB750) | Actual @ 1kV WD4 = | 7775
1512 | Objective Shift Limit (XB750) | Actual @ 20kV WD2 = | 7776
11334 | Anode_US_HRQA 40nA, Surface Overview | Anode_US_HRQA - Passed inspection criteria, Image Saved | 7777
11333 | Anode_US_HRQA 20nA, Surface Overview | Anode_US_HRQA - Passed inspection criteria, Image Saved | 7778
758 | SF_TS_HRQA, Surface Overview and Concentricity | SF_TS_HRQA - Passed inspection criteria, Image Saved | 7779
11173 | SF_TS_35um – SINGLE FILM APERTURE (SFA) Topside | Measure aperture size, 35um (+/- 2um) | 7780
1443 | PC Power Up and Enter Bios | Set Dash Client service to Manual | 7781
1443 | PC Power Up and Enter Bios | Disable USB Power Save | 7782
1443 | PC Power Up and Enter Bios | Check PC Part Number | 7783
1443 | PC Power Up and Enter Bios | Check Windows 11 Certification label is fitted | 7784
252 | Configure PC ASAP | PC S/N | 7785
252 | Configure PC ASAP | Hard disk Image (Version) | 7786
252 | Configure PC ASAP | Record Operating System | 7787
252 | Configure PC ASAP | Set the Time and Date on the PC to current UK | 7788
252 | Configure PC ASAP | Check the Entitlement certificate to determine which version of Smartsem to install | 7789
252 | Configure PC ASAP | Install Smart SEM (See WI 9693) Record Version | 7790
252 | Configure PC ASAP | Check for available SmartSEMService Packs and install them. (N/A if none) | 7791
252 | Configure PC ASAP | Check for available SmartSEM hotfixes and install them. (N/A if none) | 7792
252 | Configure PC ASAP | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 7793
252 | Configure PC ASAP | Install CZ Support Tools. Ver | 7794
252 | Configure PC ASAP | Run ASAP Screening Program | 7795
252 | Configure PC ASAP | System Specific Datazone Loaded | 7796
252 | Configure PC ASAP | In SEM status set 'Mains Sync' to off | 7797
444 | Objective Shift Limit | Actual @ 1kV WD5 = | 7798
444 | Objective Shift Limit | Actual @ 20kV WD2 = | 7799
1483 | Configure PC ASAP (Crossbeam 330) | PC S/N | 7800
1483 | Configure PC ASAP (Crossbeam 330) | Hard disk Image (Version) | 7801
1483 | Configure PC ASAP (Crossbeam 330) | Record Operating System | 7802
1483 | Configure PC ASAP (Crossbeam 330) | Set the Time and Date on the PC to current UK | 7803
1483 | Configure PC ASAP (Crossbeam 330) | Check the Entitlement certificate to determine which version of Smartsem to install | 7804
1483 | Configure PC ASAP (Crossbeam 330) | Install Smart SEM - This system should be configured as a Crossbeam 330 (Early systems will use Beta version of Smartsem V8.04) | 7805
1483 | Configure PC ASAP (Crossbeam 330) | Confirm SmartSEM is configured as Crossbeam 330 | 7806
1483 | Configure PC ASAP (Crossbeam 330) | Check for available SmartSEMService Packs and install them. (N/A if none) | 7807
1483 | Configure PC ASAP (Crossbeam 330) | Check for available SmartSEM hotfixes and install them. (N/A if none) | 7808
1483 | Configure PC ASAP (Crossbeam 330) | Installation of Zen Core EM - Ensure the correct conguration is selected (if Volutome is present on the sales order, configure ZEN for Volutome). | 7809
1483 | Configure PC ASAP (Crossbeam 330) | Install CZ Support Tools. Ver | 7810
1483 | Configure PC ASAP (Crossbeam 330) | Run ASAP Screening Program | 7811
1483 | Configure PC ASAP (Crossbeam 330) | System Specific Datazone Loaded | 7812
1483 | Configure PC ASAP (Crossbeam 330) | In SEM status set 'Mains Sync' to off | 7813
311 | Objective Drift test (Macro) | Limit is 600nm .......................Actual = | 7814
311 | Objective Drift test (Macro) | Store objective drift results | 7815
11080 | EBBU | Fit Customer BB Electronics (Do not use Test Tool Electronics on this system) | 7816
11080 | EBBU | Vacuum Stability Test | 7817
11080 | EBBU | Slew Rate <2.5us | 7818
11080 | EBBU | Measure Specimen Current into a FC and then blank the beam with the EBBU. Confirm that measured Specimen current is zero. | 7819
1204 | Heater HiPot Tests HV | Confirm Earth Continuity between Column and Plinth earth terminal using DMM . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . Resistance (<0.5Ω) | 7820
1204 | Heater HiPot Tests HV | IGP Heaters (IGP#1) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 7821
1204 | Heater HiPot Tests HV | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 7822
1204 | Heater HiPot Tests HV | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2320V) (<0.10 mA) | 7823
10585 | C2D Detector | C2D Is Selected When Going Into VP Mode | 7824
10585 | C2D Detector | Image Brightness Changes When Changing Bias | 7825
10585 | C2D Detector | Breakdown Of Bias Occurs at 100% (C2D Only, Not F2) | 7826
10536 | Beam Deceleration | Check 5KV is achievable using the Bias Slider | 7827
10536 | Beam Deceleration | safety Check - Remove PL1 (Vac Ready from vac switch | 7828
10536 | Beam Deceleration | Safety Check - Remove PL2 (Turbo Speed signal) | 7829
10536 | Beam Deceleration | Safety Check - Remove PL4 (CAN Bus signal) | 7830
10536 | Beam Deceleration | Functional Check and Images (Using the BSD Detector) | 7831
10536 | Beam Deceleration | Store Beam De-Cel Functional Images - OV and 5KV | 7832
10536 | Beam Deceleration | Check Stage Bias Low Voltage Slide Varies Image Contrast | 7833
10531 | C2D/C2D F2 | Select Type | 7834
10531 | C2D/C2D F2 | Black Level Cal | 7835
10531 | C2D/C2D F2 | C2D Bias Target Check | 7836
10531 | C2D/C2D F2 | Gain Range Check | 7837
10531 | C2D/C2D F2 | Auto Level Check | 7838
10531 | C2D/C2D F2 | Auto Range Check | 7839
10531 | C2D/C2D F2 | Oxford EDX Detector Insulator Removed | 7840
10531 | C2D/C2D F2 | Function Image | 7841
10531 | C2D/C2D F2 | Check Beam Is On at Max VP Pressure | 7842
10531 | C2D/C2D F2 | Store C2D Functional Image | 7843
10531 | C2D/C2D F2 | Select Type | 7844
10531 | C2D/C2D F2 | Black Level Cal | 7845
10531 | C2D/C2D F2 | C2D Bias Target Check | 7846
10531 | C2D/C2D F2 | Gain Range Check | 7847
10531 | C2D/C2D F2 | Auto Level Check | 7848
10531 | C2D/C2D F2 | Auto Range Check | 7849
10531 | C2D/C2D F2 | Oxford EDX Detector Insulator Removed | 7850
10531 | C2D/C2D F2 | Function Image | 7851
10531 | C2D/C2D F2 | Check Beam Is On at Max VP Pressure | 7852
10531 | C2D/C2D F2 | Store C2D Functional Image | 7853
10585 | C2D Detector | Select Type | 7854
10585 | C2D Detector | C2D Is Selected When Going Into VP Mode | 7855
10585 | C2D Detector | Image Brightness Changes When Changing Bias | 7856
10585 | C2D Detector | Breakdown Of Bias Occurs at 100% (C2D Only, Not F2) | 7857