## Sheet1
mapped_df1_id | id | text
681 | 11404 | mn2 and csv
14 | 1649 | IGP 1 Pressure (≤ 9.0e-10 mbar)
14 | 1649 | IGP 2 Pressure
40 | 186 | Install 3DSM Software and Record Software Version
40 | 186 | Set Up Image Using One segment of the BSD
40 | 186 | Run 3DSM and Setup
40 | 186 | Save 3DSM Project
40 | 186 | Store 3DSM Project
40 | 186 | Install 3DSM Software and Record Software Version
40 | 186 | Set Up Image Using One segment of the BSD
40 | 186 | Run 3DSM and Setup
40 | 186 | Save 3DSM Project
40 | 186 | Store 3DSM Project
14 | 184 | Start Gun Vacuum Tracing
14 | 184 | Save Gun Vacuum Trace
28 | 409 | Gun X
28 | 409 | Gun Y
28 | 409 | Aper X
28 | 409 | Aper Y
7 | 11258 | Continuity ok
195 | 11244 | Serial No. (or N/A)
195 | 11244 | No X-Ray Leakage detected.
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | is serial number before 688330
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
7 | 210 | Visual Inspection
14 | 185 | IGP Pressure (≤ 9.0e-10 mbar)
21 | 187 | Z Move on Vent Enabled
40 | 186 | Install 3DSM Software and Record Software Version
40 | 186 | Set Up Image Using One segment of the BSD
40 | 186 | Run 3DSM and Setup
40 | 186 | Save 3DSM Project
40 | 186 | Store 3DSM Project
40 | 1251 | Resolution Result
646 | 10652 | Type of Detector
646 | 10652 | Detector Serial number
646 | 10652 | DDP Module Serial Number
646 | 10652 | Fit Flange & Mount Detector
646 | 10652 | Install EDS PCB's Into PC
646 | 10652 | Installation of DPP Module & Cables
646 | 10652 | Configure PC Network & Test
646 | 10652 | Disable Windows10 Fast Startup
646 | 10652 | Install ApexKey.edx Licences
646 | 10652 | Install Genesis.edx Licences (if Applicable)
646 | 10652 | Fit Dongle
646 | 10652 | APEX Software Version
646 | 10652 | Install APEX Software
646 | 10652 | Install Remoting
646 | 10652 | Setup RConfigure
646 | 10652 | Run API Test
646 | 10652 | Configure Imaging-Run SGShell
646 | 10652 | Calibrate Detector-Run eDPP Shell
646 | 10652 | Save Calibration Report
646 | 10652 | Detector Resolution (measured using Cu) ≈ 129 eV
646 | 10652 | Store Calibration Report
646 | 10652 | Save Calibration on Exit
646 | 10652 | Run APEX
646 | 10652 | Check Spectrum
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected
646 | 10652 | Check Mapping
646 | 10652 | Check Line Scanning
646 | 10652 | Save Aquired Spectrum, Mappimg & Line Scanning to the Image Folder
646 | 10652 | Store Aquired Spectrum (STD, Quick, HD), Mappimg & Line Scanning Images
646 | 10652 | Save Calibration Data to the Licence Folder on the SmartSEM USB Stick
646 | 10652 | Store Calibration Data File
646 | 10652 | Save Calibration Data to the Image Folder
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | is serial number before 688330
403 | 10463 | If yes reload PC image on hardisk
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
7 | 210 | Pole piece is has no damage / dents
7 | 210 | Sealing surfaces free from scratches/ imperfections
26 | 1549
14 | 181 | Bake Time =
40 | 186 | Install 3DSM Software and Record Software Version
40 | 186 | Set Up Image Using One segment of the BSD
40 | 186 | Run 3DSM and Setup
40 | 186 | Save 3DSM Project
40 | 186 | Store 3DSM Project
382 | 6928 | Macro Result ($[StageRepeatability])
285 | 3242 | Select PC type
285 | 3242 | PC S/No.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Record MAC address
285 | 3242 | Network access requested
12 | 11145 | Number of UHV Vent cycles during leak checking
12 | 11145 | Number of UHV Vent cycles during leak checking
661 | 10719 | Select Vacuum Type
661 | 10719 | Luftfilter Fitted
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
6 | 65 | Set Stage links
6 | 65 | Fill in Jumper settings label
6 | 65 | Initialisation Check
6 | 65 | Door alignment Check
6 | 65 | Magnet Setup Check
6 | 65 | Check all Interlock Keys are fitted
674 | 113 | Torque done up to 25Nm 1st nut and 15Nm Earth Nut
14 | 11122 | UHV Vent cycles
14 | 11122 | UHV Vent cycles
8 | 91 | Housing Type
8 | 91 | Inspect insulation disc using white light for cracks
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Install New Hard Disk Image (V1) - See WI
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
12 | 11145 | Number of UHV Vent cycles during leak checking
8 | 91 | Housing Type
8 | 91 | Inspect insulation disc using white light for cracks
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, is serial number brfore 688330
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
28 | 409 | Gun X
28 | 409 | Gun Y
28 | 409 | Aper X
28 | 409 | Aper Y
25 | 199 | No cut off on image at minimum Mag.
25 | 199 | Wobble removed
7 | 1650 | S/N =
7 | 1650 | Objective Coil S/N :
7 | 1650 | Scan Coil S/N :
7 | 1650 | Pole Piece S/N :
14 | 185 | IGP Pressure (≤ 9.0e-10 mbar)
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
285 | 3242 | Select PC type
285 | 3242 | PC S/No.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Record MAC address
285 | 3242 | Network access requested
20 | 252 | PC S/N
20 | 252 | Hard disk Image (Version)
20 | 252 | PC/Model (SAP No.)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Register COM API in ASAP engine
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
285 | 3242 | Select PC type
285 | 3242 | PC S/No.
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Record MAC address
285 | 3242 | Network access requested
14 | 1649 | IGP 1 Pressure (≤ 9.0e-10 mbar)
14 | 1649 | IGP 2 Pressure
285 | 3242 | Select PC type
285 | 3242 | PC S/No.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Record MAC address
285 | 3242 | Network access requested
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | Hard disk Image (Version)
20 | 252 | PC/Model (SAP No.)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Register COM API in ASAP engine
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
25 | 197 | Gun X
25 | 197 | Gun Y
25 | 197 | Aper X
25 | 197 | Aper Y
7 | 11340 | Build quality of Objective coil is satisfactory
7 | 210 | Pole piece has no damage / dents
7 | 210 | Sealing surfaces free from scratches/ imperfections
14 | 11122 | UHV Vent cycles - Gun conditioning
14 | 11122 | UHV Vent cycles - Vacuum leak
285 | 3242 | Select PC type
285 | 3242 | PC S/No.
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Record MAC address
285 | 3242 | Network access requested
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Register COM API in ASAP engine
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
653 | 193 | Cables Supplied
653 | 193 | Lindy Cables Supplied
7 | 1650 | S/N =
30 | 11069 | Dynamic focus function check - Whole image in focus
30 | 11069 | Y - Axis measurement (400μm +2%)
30 | 11069 | Orthogonality
7 | 74 | Visual Inspection
7 | 74 | Liner Fitment
7 | 74 | End Cap Fitment
20 | 10930 | Run the ‘Disable FIB–all GIS.reg’ file
14 | 11122 | UHV Vent cycles - Gun conditioning
14 | 11122 | UHV Vent cycles - Vacuum leak
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Register COM API in ASAP engine
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
653 | 193 | Lindy Cables Supplied
25 | 198 | Gun X
25 | 198 | Gun Y
25 | 198 | Aper X
25 | 198 | Aper Y
7 | 75 | Serial No.
7 | 11258 | Continuity ok
7 | 1638 | Serial No.
7 | 82 | Pressure 15 PSI for 5 minutes
29 | 11381 | Run overnight Photograph Macro
29 | 11381 | Analyse images for drift, breakdown, ESB issues
653 | 193 | Lindy Cables Supplied
401 | 194 | Lindy Cables Supplied
7 | 11340 | Build quality of Objective coil is satisfactory
7 | 74 | Visual Inspection
7 | 74 | Liner Fitment
7 | 74 | End Cap Fitment
25 | 198 | Gun X
25 | 198 | Gun Y
25 | 198 | Aper X
25 | 198 | Aper Y
7 | 1639 | Position A:
7 | 1639 | Position B:
7 | 1639 | Position C:
7 | 1639 | Position D:
7 | 1639 | Max Difference:
7 | 1641 | Stand Off 1:
7 | 1641 | Stand Off 2:
7 | 1641 | Stand Off 3:
7 | 1641 | Stand Off 4:
7 | 1641 | Max Difference:
21 | 256 | Store System checker full report
21 | 256 | Store System checker full report
45 | 205 | Run the ‘Enable FIB_MultiGIS_ UniGIS’ file
7 | 75 | Serial No.
25 | 197 | Gun X
25 | 197 | Gun Y
25 | 197 | Aper X
25 | 197 | Aper Y
7 | 1640 | Position A:
7 | 1640 | Position B:
7 | 1640 | Position C:
7 | 1640 | Position D:
7 | 1640 | Max Difference:
45 | 205 | Run the ‘Enable FIB_MultiGIS_ UniGIS’ file
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Register COM API in ASAP engine
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
29 | 11381 | Run overnight Photograph Macro
29 | 11381 | Analyse images for drift, breakdown, ESB issues
7 | 1650 | Objective Coil S/N :
7 | 1650 | Scan Coil S/N :
7 | 1650 | Pole Piece S/N :
7 | 210 | Visually inspected for damage/dents
7 | 210 | Sealing surfaces free from scratches/ imperfections
106 | 195 | Lindy Cables Supplied
7 | 11258 | Continuity ok
7 | 1650 | S/N =
45 | 205 | Run the ‘Enable FIB_MultiGIS_ UniGIS’ file
45 | 205 | Run the ‘Enable FIB_MultiGIS_ UniGIS’ file
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Register COM API in ASAP engine
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Register COM API in ASAP engine
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
8 | 91 | Housing Type
8 | 91 | Inspect insulation disc using white light for cracks
8 | 91 | Fit contact spring
7 | 1638 | Serial No.
7 | 11340 | Build quality of Objective coil is satisfactory
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Register COM API in ASAP engine
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
20 | 11038 | PC S/N
20 | 11038 | Hard disk Image (Version)
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
8 | 92 | Cooling Isolation Check (Open Circuit Coil – GND)
8 | 92 | Fit contact spring
663 | 10730 | Valves Cables
663 | 10730 | Heater Cables
663 | 10730 | Ion Pump Cable
663 | 10730 | Ion Pump PSU Serial Number
663 | 10730 | Fit Ion Pump PSU Bracket
663 | 10730 | Fit Ion Pump PSU
26 | 1549 | Use Insight Column Tool to check for cutoff.
7 | 82 | Pressure 15 PSI for 5 minutes
7 | 1639 | Position A:
7 | 1639 | Position B:
7 | 1639 | Position C:
7 | 1639 | Position D:
7 | 1639 | Max Difference:
7 | 74 | Visual Inspection
7 | 74 | Liner Fitment
7 | 74 | End Cap Fitment
7 | 75 | Serial No.
7 | 11258 | Continuity ok
7 | 1638 | Serial No.
7 | 82 | Pressure 15 PSI for 5 minutes
7 | 1639 | Position A:
7 | 1639 | Position B:
7 | 1639 | Position C:
7 | 1639 | Position D:
7 | 1639 | Max Difference:
20 | 11038 | PC S/N
20 | 11038 | Hard disk Image (Version)
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
40 | 1562 | Run Clean Cycle
7 | 1640 | Position A:
7 | 1640 | Position B:
7 | 1640 | Position C:
7 | 1640 | Position D:
7 | 1640 | Max Difference:
7 | 1640 | Position A:
7 | 1640 | Position B:
7 | 1640 | Position C:
7 | 1640 | Position D:
7 | 1640 | Max Difference:
25 | 198 | Gun X
25 | 198 | Gun Y
25 | 198 | Aper X
25 | 198 | Aper Y
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Register COM API in ASAP engine
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
647 | 10552 | Fit Test Tool SCM
647 | 10552 | Conditions set - 20KV, 1nA, Long Filament Life
647 | 10552 | Gun Monitor Setup and Running
647 | 10552 | Stability Run Passed
647 | 10552 | Gun Monitor Saved to Product Folder
647 | 10552 | Store STAB Gun Monitor Files (Parram. Mon file & Mon Export File)
7 | 1641 | Stand Off 1:
7 | 1641 | Stand Off 2:
7 | 1641 | Stand Off 3:
7 | 1641 | Stand Off 4:
7 | 1641 | Max Difference:
7 | 1641 | Stand Off 1:
7 | 1641 | Stand Off 2:
7 | 1641 | Stand Off 3:
7 | 1641 | Stand Off 4:
7 | 1641 | Max Difference:
25 | 197 | Gun X
25 | 197 | Gun Y
25 | 197 | Aper X
25 | 197 | Aper Y
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Register COM API in ASAP engine
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
647 | 206 | Dummy SCM Fitted and Aligned Correctly
7 | 11340 | Build quality of Objective coil is satisfactory
25 | 198 | Gun X
25 | 198 | Gun Y
25 | 198 | Aper X
25 | 198 | Aper Y
25 | 199 | No cut off on image at minimum Mag.
25 | 199 | Wobble removed
25 | 200 | Were you able to Align column at 99v without column intervention
25 | 200 | If Pressure Stage is fitted Tick Box In Administrator
649 | 203 | BSD Resolution 3.8nm
649 | 203 | BSD Resolution Result - If Not Achieved Contact Armin Hayn or Richard Sell
649 | 203 | Store BSD Resolution Images and Pate images for 5 x consecutive images
649 | 10556 | 30KV Resolution - Easy VP- ≤3.4nm
649 | 10556 | Store VPSE Resolution Image
649 | 202 | BSD Resolution 3.8nm
649 | 202 | BSD Resolution Result - If Not Achieved Contact Armin Hayn or Richard Sell
649 | 202 | Store BSD Resolution Images and Pate images for 5 x consecutive images
20 | 11038 | PC S/N
20 | 11038 | Hard disk Image (Version)
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
25 | 200 | Were you able to Align column at 99v without column intervention
649 | 10556 | 30KV Resolution - Easy VP- ≤3.4nm
649 | 10556 | Store VPSE Resolution Image
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Register COM API in ASAP engine
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Register COM API in ASAP engine
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Register COM API in ASAP engine
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
647 | 206 | Dummy SCM Fitted and Aligned Correctly
25 | 200 | Were you able to Align column at 99v without column intervention
649 | 204 | 30KV Resolution - Easy VP- ≤3.4nm
649 | 204 | Store VPSE Resolution Image
20 | 11038 | PC S/N
20 | 11038 | Hard disk Image (Version)
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
663 | 10730 | Valves Cables
663 | 10730 | Heater Cables
663 | 10730 | Ion Pump Cable
663 | 10730 | Ion Pump PSU Serial Number
663 | 10730 | Fit Ion Pump PSU Bracket
663 | 10730 | Fit Ion Pump PSU
647 | 206 | Dummy SCM Fitted and Aligned Correctly
647 | 10552 | Conditions set - 20KV, 1nA, Long Filament Life
647 | 10552 | Gun Monitor Setup and Running
647 | 10552 | Stability Run Passed
647 | 10552 | Gun Monitor Saved to Product Folder
647 | 10552 | Store STAB Gun Monitor Files (Parram. Mon file & Mon Export File)
7 | 11340 | Resin is uniform and not proud
7 | 11340 | No exposed wires
25 | 200 | Were you able to Align column at 99v without column intervention
648 | 201 | 15kv 0.9nm
648 | 201 | 15KV result
648 | 201 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
649 | 204 | 30KV Resolution - Easy VP- ≤3.4nm
649 | 204 | Store VPSE Resolution Image
20 | 11038 | PC S/N
20 | 11038 | Hard disk Image (Version)
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
20 | 11038 | PC S/N
20 | 11038 | Hard disk Image (Version)
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Register COM API in ASAP engine
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
25 | 200 | Were you able to Align column at 99v without column intervention
648 | 201 | 1kv 15nm
648 | 201 | 1KV result
648 | 201 | Store 1KV Resolution Images and Pate image for 5 x consecutive images
649 | 202 | 1kv 15nm
649 | 202 | 1KV result - If Not Achieved Contact Armin Hayn or Richard Sell
649 | 202 | Store 1KV Resolution Images and Pate images for 5 x consecutive images
649 | 204 | 30KV Resolution - Easy VP- ≤3.4nm
649 | 204 | Store VPSE Resolution Image
5 | 98 | Torque done up to 25Nm 1st nut and 15Nm Earth Nut
14 | 184 | Start Gun Vacuum Tracing
14 | 184 | Save Gun Vacuum Trace
665 | 10741 | EP/Water Kit Fitted
665 | 10741 | MKS Driver Box Link Set to 0-10V
665 | 10741 | TTL Kit Fitted
26 | 1549 | Objective Cap - No particles visible
26 | 1549 | Lower Aperture - No large particles visible (See WI for reference.)
26 | 1549 | Upper Aperture - No particles visible
43 | 11341 | 1kV ESB
43 | 11341 | Store 1kV ESB Customer Image
26 | 281 | Type in measured value (Limit <= 30um)
26 | 281 | Store ESB Concentricity Image
648 | 201 | 1kv 15nm
648 | 201 | 1KV result
648 | 201 | Store 1KV Resolution Images and Pate image for 5 x consecutive images
649 | 202 | BSD 15nm
649 | 202 | 1KV result - If Not Achieved Contact Armin Hayn or Richard Sell
649 | 202 | Store 1KV Resolution Images and Pate images for 5 x consecutive images
649 | 202 | BSD Resolution 4nm
649 | 202 | 1KV result - If Not Achieved Contact Armin Hayn or Richard Sell
649 | 202 | Store 1KV Resolution Images and Pate images for 5 x consecutive images
649 | 204 | 30KV Resolution - Easy VP- ≤3.4nm
649 | 204 | Store VPSE Resolution Image
285 | 3242 | Select PC type
285 | 3242 | PC S/No.
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Record MAC address
285 | 3242 | Network access requested
7 | 211 | Ensure objective coil is concentric to pole piece using alignment jig
389 | 212 | Language Set to English (United States
389 | 212 | Ethernet Connection Setup
389 | 212 | Install Aztec Software
389 | 212 | AZtec Software Version
389 | 212 | Install Licences
389 | 212 | Run AZtec Assist+
389 | 212 | Connected to Test Tool Module
25 | 219 | Ensure no cut off is visible @ 10kV on central emission image.
648 | 201 | 1kv 15nm
648 | 201 | 1KV result
648 | 201 | Store 1KV Resolution Images and Pate image for 5 x consecutive images
649 | 202 | BSD Resolution 4nm
649 | 202 | BSD Resolution Result - If Not Achieved Contact Armin Hayn or Richard Sell
649 | 202 | Store BSD Resolution Images and Pate images for 5 x consecutive images
649 | 204 | VPSE Resolution 3.0nm
649 | 204 | VPSE Resolution Result - If Not Achieved Contact Armin Hayn or Richard Sell
649 | 204 | Store BSD Resolution Images and Pate images for 5 x consecutive images
647 | 208 | Remove the Test Tool SCM
647 | 208 | Re-Connect Specimen Touch Cable
647 | 208 | Remove the Test Tool SCM
647 | 208 | Re-Connect Specimen Touch Cable to X13 on the Vac PCB
647 | 208 | Remove the Test Tool SCM
647 | 208 | Re-Connect Specimen Touch Cable to X13 on the Vac PCB
647 | 208 | Remove the Test Tool SCM
647 | 208 | Re-Connect Specimen Touch Cable to X13 on the Vac PCB
663 | 215 | Fit Valves and Pipework
663 | 215 | Connect Valve Cable to PL3 on Vac PCB
43 | 347 | 1kV Inlens
43 | 347 | Store 1kV Inlens Customer Image
43 | 347 | 15kV Inlens
43 | 347 | Store 15kV Inlens Customer Image
25 | 11401 | 20um aperture entered on central aperture
648 | 201 | 1kv 15nm
648 | 201 | 1KV result - If Not Achieved Contact Armin Hayn or Richard Sell
648 | 201 | Store 1KV Resolution Images and Pate image for 5 x consecutive images
649 | 10555 | BSD Resolution ≤4nm. On completion Proceed to Task BSD Res - New Spec
649 | 10555 | Store BSDResolution Image
647 | 207 | Test Tool SCM Fitted
647 | 208 | Remove the Test Tool SCM
647 | 208 | Re-Connect Specimen Touch Cable
647 | 208 | Remove the Test Tool SCM
647 | 208 | Re-Connect Specimen Touch Cable to X13 on the Vac PCB
647 | 208 | Remove the Test Tool SCM
647 | 208 | Re-Connect Specimen Touch Cable to X13 on the Vac PCB
652 | 213 | Take Photo of Label
43 | 347 | 1kV Inlens
43 | 347 | Store 1kV Inlens Customer Image
43 | 347 | 15kV Inlens
43 | 347 | Store 15kV Inlens Customer Image
43 | 347 | 1kV Inlens
43 | 347 | Store 1kV Inlens Customer Image
43 | 347 | 15kV Inlens
43 | 347 | Store 15kV Inlens Customer Image
25 | 216 | 20um aperture entered on central aperture
46 | 1567 | Remove the sample holder from the system.
46 | 1567 | Remove Docking Brackets and fit Transit clamps.
46 | 1567 | Airfeet deflated and free to move.
46 | 1567 | Backing lines re-configured so that PVSV valve is attached to flexi hose. Secure valve to frame.
46 | 1567 | Confirm that the 3 'o' rings are fitted to the flexi hose backing line
46 | 1567 | If applicable, ensure that the A/L push rod is removed and blank fitted in its place. Ensure A/L is pumped down before powering down system.
46 | 1567 | If applicable, ensure that the Beamsleeve funnel has been removed from the detector
46 | 1567 | Select 'Partial vent on standby' in Smartsem.
648 | 201 | 1kv 15nm
648 | 201 | 1KV result - If Not Achieved Contact Armin Hayn or Richard Sell
648 | 201 | Store 1KV Resolution Images and Pate images for 5 x consecutive images
649 | 203 | BSD Resolution 4nm
649 | 203 | BSD Resolution Result - If Not Achieved Contact Armin Hayn or Richard Sell
649 | 203 | Store BSD Resolution Images and Pate images for 5 x consecutive images
647 | 208 | Remove the Test Tool SCM
647 | 208 | Re-Connect Specimen Touch Cable to X13 on the Vac PCB
647 | 208 | Remove the Test Tool SCM
647 | 208 | Re-Connect Specimen Touch Cable to X13 on the Vac PCB
652 | 213 | Take Photo of Label
652 | 213 | Oxford Sales Order
652 | 213
43 | 347 | 1kV Inlens
43 | 347 | Store 1kV Inlens Customer Image
43 | 347 | 15kV Inlens
43 | 347 | Store 15kV Inlens Customer Image
43 | 11341 | 1kV ESB
43 | 11341 | Store 1kV ESB Customer Image
46 | 11408 | Remove the sample holder from the system.
46 | 11408 | Remove Test Tool SE2 and fit to warehouse stand. Fit SE2 blank plate to chamber and SHV connector to Spellman (00000-0302-367).
46 | 11408 | Confirm that Raith top mounted SE2 detector is mounted to a warehouse stand and that a protective cap is fitted over the light guide.
46 | 11408 | Set Stage position to Z=41mm & X & Y =65mm.
46 | 11408 | Remove SCM and pack for shipment (include screws)
46 | 11408 | Remove Docking Brackets and fit Transit clamps.
46 | 11408 | Airfeet deflated and free to move.
46 | 11408 | Backing lines re-configured so that PVSV valve is attached to flexi hose. Secure valve to frame.
46 | 11408 | Confirm that the 3 'o' rings are fitted to the flexi hose backing line
46 | 11408 | Select 'Partial vent on standby' in Smartsem.
25 | 11401 | 30um aperture entered on central aperture
674 | 113 | Torque the M8 Zinc plated earth nuts to 15 N.m
648 | 10553 | 30KV Res- ≤2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 1KV Res- ≤20nm - On completion Proceed to Task 1KV W - New Spec.
648 | 10553 | 200V Res- Not measured
648 | 10553 | Store Standard Resolution Images
649 | 203 | BSD Resolution 4nm
649 | 203 | BSD Resolution Result - If Not Achieved Contact Armin Hayn or Richard Sell
649 | 203 | Store BSD Resolution Images and Pate images for 5 x consecutive images
647 | 208 | Remove the Test Tool SCM
647 | 208 | Re-Connect Specimen Touch Cable to X13 on the Vac PCB
652 | 213 | Take Photo of Label
652 | 213 | Oxford Sales Order
44 | 349 | Carry out safety tests using the Seaward Prime 100 Tester.
44 | 349 | Fill out completely the Product Safety Tests form in Section 13.
389 | 3300 | Function Test
648 | 10553 | 30KV Res- ≤2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 1KV Res- ≤15nm
648 | 10553 | 200V Res- Not measured
648 | 10553 | Store Standard Resolution Images
661 | 10719 | Select Vacuum Type
661 | 10719 | Luftfilter Fitted
385 | 3283 | EXT Voltage
25 | 219 | Ensure no cut off is visible @ 10kV, SS9 on central aperture of emission image.
12 | 224 | Confirm correct fitting of Raith spray aperture
648 | 201 | 1kv 15nm
648 | 201 | 1KV result - If Not Achieved Contact Armin Hayn or Richard Sell
648 | 201 | Store 1KV Resolution Images and Pate images for 5 x consecutive images
400 | 214 | Take Photo of Label
400 | 214 | Oxford Sales Order
661 | 10719 | Select Vacuum Type
661 | 10719 | Luftfilter Fitted
646 | 10960 | Install 3DSM Software and Record Software Version
646 | 10960 | Set Up Image Using One segment of the BSD
646 | 10960 | Run 3DSM and Setup
646 | 10960 | Save 3DSM Project
646 | 10960 | Store 3DSM Project
40 | 1555 | Leak check
40 | 1555 | Pneumatics OK
40 | 1555 | Diode Alignment OK
40 | 1555 | Good Repeatability
40 | 1555 | Bias Voltage Working
40 | 1555 | Vac Degradation < 3x10^-5 mbar
40 | 1555 | NanoVP Operation Working over full range
40 | 1555 | XVP Operation working over full range
25 | 219 | No cut off is visible @ 10kV, SS9 on central aperture of emission image.
646 | 10547 | ZEN Version
646 | 10547 | Correct Version of Windows 10 Installed
646 | 10547 | Install ZEN
646 | 10547 | Install Licence
646 | 10547 | Install Remote API Client
646 | 10547 | Functional Test
646 | 10547 | Save Snapped Image
646 | 10547 | Store Snapped Image
646 | 10547 | Store Entitlement Certificate
389 | 10883 | ZEN Version
389 | 10883 | Correct Version of Windows 10 Installed
389 | 10883 | Install ZEN
389 | 10883 | Install Licence
389 | 10883 | Install Remote API Client
389 | 10883 | Functional Test
389 | 10883 | Save Snapped Image
389 | 10883 | Store Snapped Image
389 | 10883 | Store Entitlement Certificate
647 | 208 | Remove the Test Tool SCM
647 | 208 | Connect Specimen Touch Cable to X13 on Vac PCB
389 | 217 | Function Test
27 | 406 | Alignment done and confirmed.
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, Install New Hard Disk Image (V1) - See WI
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
646 | 84 | Language Set to English (United States
646 | 84 | Ethernet Connection Setup
646 | 84 | Install Aztec Software
646 | 84 | AZtec Software Version
646 | 84 | Install Licences
646 | 84 | Run AZtec Assist+
646 | 84 | Connected to Test Tool Module
646 | 84 | Language Set to English (United States
646 | 84 | Ethernet Connection Setup
646 | 84 | Install Aztec Software
646 | 84 | AZtec Software Version
646 | 84 | Install Licences
646 | 84 | Run AZtec Assist+
646 | 84 | Connected to Test Tool Module
26 | 1549 | Objective Cap - No particles visible
26 | 1549 | Lower Aperture - No large particles visible (See WI for details)
26 | 1549 | Upper Aperture - No cut-off on image and no particles visible.
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, then the PC needs reimaging – follow the attached WI. The correct image version is V1.
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
647 | 206 | Dummy SCM Fitted and Aligned Correctly
647 | 206 | Specimen Touch Cable Connected to Dummy SCM
389 | 3300 | Function Test
389 | 3300 | Check Break Button
646 | 84 | Language Set to English (United States
646 | 84 | Ethernet Connection Setup
646 | 84 | Install Aztec Software
646 | 84 | AZtec Software Version
646 | 84 | Install Licences
646 | 84 | Run AZtec Assist+
646 | 84 | Connected to Test Tool Module
285 | 3242 | Select PC type
285 | 3242 | PC S/No.
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, then the PC needs reimaging – follow the attached WI. The correct image version is V1.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Record MAC address
285 | 3242 | Network access requested
12 | 11266 | Flange Heater Earths routed correctly (where applicable)
12 | 11266 | Thermocouple cables routed correctly
389 | 10883 | ZEN Version
389 | 10883 | Correct Version of Windows 10 Installed
389 | 10883 | Install ZEN
389 | 10883 | Install Licence
389 | 10883 | Install Remote API Client
389 | 10883 | Functional Test
389 | 10883 | Store Entitlement Certificate
40 | 11023 | ZEN Version
40 | 11023 | Correct Version of Windows 10 Installed
40 | 11023 | Install ZEN
40 | 11023 | Install Licence
40 | 11023 | Install Remote API Client
40 | 11023 | Functional Test
40 | 11023 | Save Snapped Image
40 | 11023 | Store Snapped Image
40 | 11023 | Store Entitlement Certificate
30 | 11069 | Dynamic focus function check - Whole image in focus
30 | 11069 | Y - Axis measurement (400μm +/- 2%)
30 | 11069 | Orthogonality
5 | 10989 | Align Needle
5 | 10989 | Pneumatics ok
5 | 10989 | Pressure Test (CC Valve fully open)
5 | 10989 | Take image of paper
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, then the PC needs reimaging – follow the attached WI. The correct image version is V1.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
25 | 11036 | Large Aperture in Top Left Quadrant
25 | 11036 | Save Emission Image with Condenser On
25 | 11036 | Save Emission Image with Condenser Off
646 | 10960 | Install 3DSM Software and Record Software Version
646 | 10960 | Set Up Image Using One segment of the BSD
646 | 10960 | Run 3DSM and Setup
646 | 10960 | Save 3DSM Project
646 | 10960 | Store 3DSM Project
4 | 11389 | Fit chamber pumping pipe
4 | 11389 | Fit chamber to damper
4 | 11389 | Fit chamber bolts
4 | 11389 | Fit earthquake bolts
5 | 98 | Torque both earth zinc plated M8 nuts to 15 N.m
38 | 337 | Blanking Drift Distance Median (Limit <= 300nm) =
38 | 337 | Blanking Drift Time Median (Limit <= 3.5 Sec) =
38 | 337 | Corner Test Max Drift (Limit <= 300nm) =
38 | 337 | Store Autobeamshift results PDF
648 | 10658 | 2KV - 100um VP Aperture- ≤12nm
648 | 10658 | 2KV - 140um VP Aperture- ≤12nm
648 | 10658 | Store 2KV Aperture Resolution Images
12 | 225 | Confirm correct fitting of Raith spray aperture
5 | 98 | Torque both earth zinc plated M8 nuts to 15 N.m
40 | 1561 | EHT Cal
40 | 1561 | Spark Alarm
40 | 1561 | Resolution Preperation
40 | 1561 | Acquire a customer proof image @ 1kV landing energy 5kV bias on gold on carbon (Resolution not required)
650 | 10567 | Install Customer Licences
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer to the Product Folder
650 | 10567 | Store EM Server
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
650 | 10567 | Add Licence to the SmartSEM USB Stick
674 | 113 | Safety – Torque the earth stud to 15 N.m
674 | 113 | Torque the M8 Zinc plated earth nuts to 15 N.m
12 | 225 | Confirm correct fitting of Raith spray aperture
652 | 213 | Take Photo of Label
652 | 213 | Oxford Sales Order
652 | 213 | Store Oxford Sales Order Label Images
652 | 213 | Take Photo of Label
652 | 213 | Oxford Sales Order
652 | 213 | Store Oxford Sales Order Label Images
652 | 213 | Take Photo of Label
652 | 213 | Oxford Sales Order
652 | 213 | Store Oxford Sales Order Label Images
5 | 87 | Verify the SVSEM cables are connected to the customer port configuration
38 | 11315 | Confirm that when the beam is blanked the measured I Probe <5pA
38 | 11315 | Difference between EHT on and off < 1pA.
708 | 222 | Check SCS Folder/Sales Order
708 | 222 | SCS Number (If applicable)
5 | 98 | Safety – Torque the earth stud to 20 N.m
5 | 98 | Torque both earth zinc plated M8 nuts to 15 N.m
5 | 98 | Safety – Torque the earth stud to 15 N.m
5 | 98 | Torque both earth zinc plated M8 nuts to 15 N.m
12 | 225 | Confirm that blanker cable routing is correct
652 | 213 | Take Photo of Label
652 | 213 | Oxford Sales Order
652 | 213 | Store Oxford Sales Order Label Images
102 | 1626 | Function of Internal Specimen Current Monitor checked.
102 | 1626 | Confirm when Beam Blanked measured Specimen Current <5pa
652 | 213 | Take Photo of Label
652 | 213 | Oxford Sales Order
652 | 213 | Store Oxford Sales Order Label Images
5 | 11275 | Verify the ASAP cables are connected to the customer port configuration
652 | 213 | Take Photo of Label
652 | 213 | Oxford Sales Order
652 | 213 | Store Oxford Sales Order Label Images
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
36 | 324 | Max Dev= Measured value +/-15% of demanded current
36 | 321 | Max dev = Measured value +3pa of slider
674 | 113 | Safety – Torque the earth stud to 20 N.m
674 | 113 | Torque the M8 Zinc plated earth nuts to 15 N.m
652 | 213 | Take Photo of Label
652 | 213 | Oxford Sales Order
652 | 213 | Store Oxford Sales Order Label Images
649 | 203 | BSD Resolution 3.8nm
649 | 203 | BSD Resolution Result - If Not Achieved Contact Armin Hayn or Richard Sell
649 | 203 | Store BSD Resolution Images and Pate images for 5 x consecutive images
36 | 322 | Max dev = Measured value +/-15% of slider (Use Gun Aligns to maximise Probe I)
18 | 395 | Objective Coil S/N :
18 | 395 | Scan Coil S/N :
18 | 395 | Pole Piece S/N :
18 | 382 | Ensure objective coil is concentric to pole piece, this should be done as best as possible by eye as no jig yet available.
18 | 382 | Ensure objective coil is concentric to pole piece ( this should be done as best as possible by eye as no jig yet available)
27 | 434 | Image Shift (um) =
39 | 339 | After conditioning confirm there are no EMServer error messages
39 | 339 | Store Gun monitor trace as 30kV Cycle Conditioning.mn2
39 | 339 | After conditioning confirm there are no EMServer error messages
39 | 339 | Store Gun monitor trace as 30kV Cycle Conditioning.mn2
7 | 223 | Check SCS Folder/Sales Order
7 | 223 | SCS Number (If applicable)
652 | 213 | Take Photo of Label
652 | 213 | Oxford Sales Order
652 | 213 | Store Oxford Sales Order Label Images
649 | 203 | BSD Resolution 3.8nm
649 | 203 | BSD Resolution Result - If Not Achieved Contact Armin Hayn or Richard Sell
649 | 203 | Store BSD Resolution Images and Pate images for 5 x consecutive images
649 | 10555 | BSD Resolution ≤4nm. On completion Proceed to Task BSD Res - New Spec
649 | 10555 | Store BSDResolution Image
36 | 323 | Max dev= Measured value +/-15% of slider.
7 | 211 | Ensure objective coil is concentric to pole piece ( this should be done as best as possible by eye as no jig yet available)
17 | 251 | Ensure objective coil is concentric to pole piece ( this should be done as best as possible by eye as no jig yet available)
39 | 339 | After conditioning confirm there are no EMServer error messages
39 | 339 | Store Gun monitor trace as 30kV Cycle Conditioning.mn2
663 | 414 | No Leaks Detected
649 | 204 | VPSE Resolution 3.0nm
649 | 204 | VPSE Resolution Result - If Not Achieved Contact Armin Hayn or Richard Sell
649 | 204 | Store BSD Resolution Images and Pate images for 5 x consecutive images
649 | 204 | VPSE Resolution 3.0nm
649 | 204 | Store BSD Resolution Image
649 | 10555 | BSD Resolution ≤4nm. On completion Proceed to Task BSD Res - New Spec
649 | 10555 | Store BSDResolution Image
649 | 10555 | BSD Resolution ≤4nm. On completion Proceed to Task BSD Res - New Spec
649 | 10555 | Store BSDResolution Image
649 | 10555 | BSD Resolution ≤4nm. On completion Proceed to Task BSD Res - New Spec
649 | 10555 | Store BSDResolution Image
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
25 | 279 | Fil I
25 | 279 | Ext V
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
15 | 228 | Check SCS Folder/Sales Order
15 | 228 | SCS Number (If applicable)
649 | 204 | VPSE Resolution 3.0nm
649 | 204 | Store VPSE Resolution Image
40 | 1557 | CCD
40 | 1557 | Focus
18 | 393 | Visual Inspection
18 | 393 | Liner Fitment
18 | 393 | End Cap Fitment
19 | 408 | After conditioning confirm there are no EMServer error messages
19 | 408 | Store Gun monitor trace as 30kV Cycle Conditioning.mn2
4 | 11384 | No Leaks Detected
44 | 358 | All verification & Safety tests done and within threshold limits defined
44 | 358 | Sales Order is latest version
44 | 358 | All items on Sales order accounted for
44 | 358 | ESB SN matches
15 | 229 | Objective Coil S/N :
15 | 229 | Scan Coil S/N :
15 | 229 | Pole Piece S/N :
649 | 10659 | 30KV Resolution- ≤5.5nm
649 | 10659 | Store VPSE Resolution Image
649 | 10659 | 30KV Resolution- ≤5.5nm
649 | 10659 | Store VPSE Resolution Image
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
649 | 10555 | BSD Resolution ≤4nm. On completion Proceed to Task BSD Res - New Spec
649 | 10555 | Store BSDResolution Image
649 | 230 | BSD Resolution ≤4nm.
649 | 230 | Store BSD Resolution Image
649 | 230 | BSD Resolution ≤4nm.
649 | 230 | Store BSD Resolution Image
649 | 230 | BSD Resolution ≤4nm.
649 | 230 | Store BSD Resolution Image
649 | 227 | 30KV Resolution - Easy VP- ≤5.5nm
649 | 227 | Store VPSE Resolution Image
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
649 | 10555 | BSD Resolution ≤4nm. On completion Proceed to Task BSD Res - New Spec
649 | 10555 | Store BSDResolution Image
649 | 230 | BSD Resolution ≤4nm.
649 | 230 | Store BSD Resolution Image
649 | 230 | BSD Resolution ≤4nm.
649 | 230 | Store BSD Resolution Image
40 | 1558 | CCD
40 | 1558 | Focus
10 | 109 | Continuity (Close Circuit)
10 | 109 | Confirm that the coil is NOT short circuit to the condenser body. Do not use test box, check manually.
26 | 1549 | Objective Cap - No particles visible
26 | 1549 | Lower Aperture - No large particles visible (See WI for details)
26 | 1549 | Upper Aperture - No cut-off on image and no particles visible.
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
649 | 230 | BSD Resolution ≤4nm.
649 | 230 | Store BSD Resolution Image
649 | 230 | BSD Resolution ≤4nm.
649 | 230 | Store BSD Resolution Image
649 | 230 | BSD Resolution ≤4nm.
649 | 230 | Store BSD Resolution Image
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
44 | 348 | Chamber Door Interlock
44 | 348 | Service Switch
44 | 348 | Detector de-activation
44 | 348 | Detector De-activation
44 | 348 | EHT + IPPSU De-activation
44 | 348 | Chamber Door Interlock
44 | 348 | Service Switch
44 | 348 | Detector De-activation
44 | 348 | EHT + IPPSU De-activation
16 | 231 | Check SCS Folder/Sales Order
16 | 231 | SCS Number (If applicable)
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
44 | 348 | Chamber Door Interlock
44 | 348 | Service Switch
44 | 348 | Detector De-activation
44 | 348 | EHT + IPPSU De-activation
16 | 231 | Check SCS Folder/Sales Order
16 | 231 | SCS Number (If applicable)
16 | 232 | Objective Coil S/N :
16 | 232 | Scan Coil S/N :
16 | 232 | Pole Piece S/N :
16 | 231 | Check SCS Folder/Sales Order
16 | 231 | SCS Number (If applicable)
40 | 1557 | CCD
40 | 1557 | Focus
43 | 346 | 1kV =
43 | 346 | Store 1KV Resolution Images and Pate image
43 | 346 | 15kV =
43 | 346 | Store 1KV Resolution Images and Pate image
18 | 392 | Continuity ok
649 | 10555 | BSD Resolution ≤4nm. On completion Proceed to Task BSD Res - New Spec
649 | 10555 | Store BSDResolution Image
649 | 204 | VPSE Resolution 3.0nm
649 | 204 | Store VPSE Resolution Image
16 | 232 | Objective Coil S/N :
16 | 232 | Scan Coil S/N :
16 | 232 | Pole Piece S/N :
40 | 1558 | CCD
40 | 1558 | Focus
43 | 346 | 1kV =
43 | 346 | Store 1KV Resolution Images and Pate image
43 | 346 | 15kV =
43 | 346 | Store 1KV Resolution Images and Pate image
18 | 391 | Visually inspected for damage/dents
18 | 391 | Sealing surfaces free from scratches/ imperfections
19 | 408 | After stability confirm there are no EMServer error messages
19 | 408 | Store Probe Stability trace as 30kV Cycle Conditioning.mn2
674 | 416 | No Leaks Detected
285 | 3242 | Select PC type
285 | 3242 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
285 | 3242 | PC S/No.
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, then the PC needs reimaging – follow the attached WI. The correct image version is V1.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Record MAC address
285 | 3242 | Network access requested
649 | 10555 | BSD Resolution ≤4nm. On completion Proceed to Task BSD Res - New Spec
649 | 10555 | Store BSDResolution Image
17 | 247 | Confirm there is no Electical breakdown between Scan coils and Objective body.
17 | 248 | Objective Coil
17 | 248 | Scan Coil
18 | 378 | Confirm there is no Electical breakdown between Scan coils and Objective body.
25 | 398 | Are Gun Head clamping screws secure?
25 | 398 | Are Condensor clamping screws secure?
25 | 398 | Store 1kV emission image (Condensor off)
25 | 398 | Check emission image against 'emissions distortion criteria'
25 | 398 | Store 10kV emission image (Condensor off)
25 | 400 | Stig X
25 | 400 | Stig Y
18 | 390 | Position A:
18 | 390 | Position B:
18 | 390 | Position C:
18 | 390 | Position D:
18 | 390 | Max Difference:
18 | 386 | Resin is uniform and not proud
18 | 386 | No exposed wires
649 | 10555 | BSD Resolution ≤4nm. On completion Proceed to Task BSD Res - New Spec
649 | 10555 | Store BSDResolution Image
16 | 234 | Visual Inspection
16 | 234 | Liner Fitment
16 | 234 | End Cap Fitment
16 | 235 | Continuity ok
16 | 234 | Visual Inspection
16 | 234 | Liner Fitment
16 | 234 | End Cap Fitment
25 | 219 | No cut off is visible @ 10kV, SS9 on central aperture of emission image.
25 | 400 | Stig X
25 | 400 | Stig Y
18 | 384 | Pressure 15 PSI for 5 minutes
18 | 383 | Stand Off 1:
18 | 383 | Stand Off 2:
18 | 383 | Stand Off 3:
18 | 383 | Stand Off 4:
18 | 383 | Max Difference:
18 | 382 | Ensure objective coil is concentric to pole piece using alignment jig
19 | 408 | After stability confirm there are no EMServer error messages
19 | 408 | Store Probe Stability trace as Probe I Stability.mn2
649 | 10555 | BSD Resolution ≤4nm. On completion Proceed to Task BSD Res - New Spec
649 | 10555 | Store BSDResolution Image
16 | 236 | Visually inspected for damage/dents
16 | 236 | Sealing surfaces free from scratches/ imperfections
16 | 235 | Continuity ok
17 | 249 | Open Circuit (Coil – GND).
25 | 398 | Are Gun Head clamping screws secure?
25 | 398 | Are Condensor clamping screws secure?
25 | 398 | Store 1kV emission image (Condensor off)
25 | 398 | Check emission image against 'emissions distortion criteria'
25 | 398 | Store 10kV emission image (Condensor off)
25 | 398 | Are Gun Head clamping screws secure?
25 | 398 | Are Condensor clamping screws secure?
25 | 398 | Store 1kV emission image (Condensor off)
25 | 398 | Check emission image against 'emissions distortion criteria'
25 | 398 | Store 10kV emission image (Condensor off)
25 | 399 | Aperture Finder run succesfully
649 | 10555 | BSD Resolution ≤4nm. On completion Proceed to Task BSD Res - New Spec
649 | 10555 | Store BSDResolution Image
16 | 237 | Position A:
16 | 237 | Position B:
16 | 237 | Position C:
16 | 237 | Position D:
16 | 237 | Max Difference:
16 | 238 | Position A:
16 | 238 | Position B:
16 | 238 | Position C:
16 | 238 | Position D:
16 | 238 | Max Difference:
16 | 236 | Visually inspected for damage/dents
16 | 236 | Sealing surfaces free from scratches/ imperfections
17 | 251 | Ensure objective coil is concentric to pole piece using alignment jig
18 | 382 | Ensure objective coil is concentric to pole piece using alignment jig
25 | 275 | Are Gun Head clamping screws secure?
25 | 275 | Are Condensor clamping screws secure?
25 | 275 | Store 1kV emission image (Condensor off)
25 | 275 | Check emission image against 'emissions distortion criteria'
25 | 275 | Store 10kV emission image (Condensor off)
18 | 380 | Open Circuit (Coil – GND).
18 | 379 | Objective Coil
18 | 379 | Scan Coil
16 | 239 | Resin is uniform and not proud
16 | 239 | No exposed wires
16 | 237 | Position A:
16 | 237 | Position B:
16 | 237 | Position C:
16 | 237 | Position D:
16 | 237 | Max Difference:
17 | 263 | Stand Off 1:
17 | 263 | Stand Off 2:
17 | 263 | Stand Off 3:
17 | 263 | Stand Off 4:
17 | 263 | Max Difference:
18 | 379 | Objective Coil
18 | 379 | Scan Coil
25 | 275 | Are Gun Head clamping screws secure?
25 | 275 | Are Condensor clamping screws secure?
25 | 275 | Store 1kV emission image (Condensor off)
25 | 275 | Check emission image against 'emissions distortion criteria'
25 | 275 | Store 10kV emission image (Condensor off)
18 | 378 | Confirm there is no Electical breakdown between Scan coils and Objective body.
19 | 408 | After stability confirm there are no EMServer error messages
19 | 408 | Store Probe Stability trace as Probe I Stability.mn2
16 | 240 | Pressure 15 PSI for 5 minutes
16 | 239 | Resin is uniform and not proud
16 | 239 | No exposed wires
18 | 380 | Open Circuit (Coil – GND).
16 | 231 | Check SCS Folder/Sales Order
16 | 231 | SCS Number (If applicable)
19 | 408 | Actual =
19 | 408 | After conditioning confirm there are no EMServer error messages
19 | 408 | Store Gun monitor trace as Probestability.mn2
646 | 84 | Language Set to English (United States
646 | 84 | Ethernet Connection Setup
646 | 84 | Install Aztec Software
646 | 84 | AZtec Software Version
646 | 84 | Install Licences
646 | 84 | Run AZtec Assist+
646 | 84 | Connected to Test Tool Module
389 | 212 | Language Set to English (United States
389 | 212 | Ethernet Connection Setup
389 | 212 | Install Aztec Software
389 | 212 | AZtec Software Version
389 | 212 | Install Licences
389 | 212 | Run AZtec Assist+
389 | 212 | Connected to Test Tool Module
389 | 212 | Language Set to English (United States
389 | 212 | Ethernet Connection Setup
389 | 212 | Install Aztec Software
389 | 212 | AZtec Software Version
389 | 212 | Install Licences
389 | 212 | Run AZtec Assist+
389 | 212 | Connected to Test Tool Module
| 439 | Ion Pump Current draw < 1.0e-07 A
50 | 480 | DC Tablefiller successfully completed.
649 | 10555 | BSD Resolution ≤4nm.
649 | 10555 | Store BSDResolution Image
16 | 240 | Pressure 15 PSI for 5 minutes
16 | 241 | Stand Off 1:
16 | 241 | Stand Off 2:
16 | 241 | Stand Off 3:
16 | 241 | Stand Off 4:
16 | 241 | Max Difference:
17 | 271 | Pressure 15 PSI for 5 minutes
16 | 232 | Objective Coil S/N :
16 | 232 | Scan Coil S/N :
16 | 232 | Pole Piece S/N :
19 | 408 | Actual =
19 | 408 | After conditioning confirm there are no EMServer error messages
19 | 408 | Store Gun monitor trace as Probestability.mn2
646 | 84 | Language Set to English (United States
646 | 84 | Ethernet Connection Setup
646 | 84 | Install Aztec Software
646 | 84 | AZtec Software Version
646 | 84 | Install Licences
646 | 84 | Run AZtec Assist+
646 | 84 | Connected to Test Tool Module
25 | 11410 | Align offset angle set to -110 degrees.
12 | 145 | Checked By Assembly Engineer:
649 | 230 | BSD Resolution ≤4nm.
649 | 230 | Store BSDResolution Image
16 | 241 | Stand Off 1:
16 | 241 | Stand Off 2:
16 | 241 | Stand Off 3:
16 | 241 | Stand Off 4:
16 | 241 | Max Difference:
16 | 242 | Ensure objective coil is concentric to pole piece using alignment jig
17 | 274 | Resin is uniform and not proud
17 | 274 | No exposed wires
16 | 234 | Visual Inspection
16 | 234 | Liner Fitment
16 | 234 | End Cap Fitment
16 | 235 | Continuity ok
16 | 236 | Visually inspected for damage/dents
16 | 236 | Sealing surfaces free from scratches/ imperfections
16 | 237 | Position A:
16 | 237 | Position B:
16 | 237 | Position C:
16 | 237 | Position D:
16 | 237 | Max Difference:
12 | 1660 | Initials of 2nd Engineer confirming.
649 | 230 | BSD Resolution ≤4nm.
649 | 230 | Store BSD Resolution Image
649 | 10555 | BSD Resolution ≤4nm.
649 | 10555 | Store BSD Resolution Image
16 | 242 | Ensure objective coil is concentric to pole piece using alignment jig
16 | 244 | Open Circuit (Coil – GND).
16 | 245 | Objective Coil
16 | 245 | Scan Coil
16 | 246 | Confirm there is no Electical breakdown between Scan coils and Objective body.
18 | 383 | Stand Off 1:
18 | 383 | Stand Off 2:
18 | 383 | Stand Off 3:
18 | 383 | Stand Off 4:
18 | 383 | Max Difference:
16 | 239 | Resin is uniform and not proud
16 | 239 | No exposed wires
16 | 240 | Pressure 15 PSI for 5 minutes
646 | 84 | Language Set to English (United States
646 | 84 | Ethernet Connection Setup
646 | 84 | Install Aztec Software
646 | 84 | AZtec Software Version
646 | 84 | Install Licences
646 | 84 | Run AZtec Assist+
646 | 84 | Connected to Test Tool Module
23 | 269 | Run macro 'ESB Runup 3h pumpen & 9h hochfahren.MLF'
649 | 230 | BSD Resolution ≤4nm.
649 | 230 | Store BSD Resolution Image
649 | 230 | BSD Resolution ≤4nm.
649 | 230 | Store BSD Resolution Image
17 | 288 | Position A:
17 | 288 | Position B:
17 | 288 | Position C:
17 | 288 | Position D:
17 | 288 | Max Difference:
18 | 384 | Pressure 15 PSI for 5 minutes
18 | 386 | Resin is uniform and not proud
18 | 386 | No exposed wires
18 | 388 | Position A:
18 | 388 | Position B:
18 | 388 | Position C:
18 | 388 | Position D:
18 | 388 | Max Difference:
18 | 390 | Position A:
18 | 390 | Position B:
18 | 390 | Position C:
18 | 390 | Position D:
18 | 390 | Max Difference:
18 | 391 | Visually inspected for damage/dents
18 | 391 | Sealing surfaces free from scratches/ imperfections
16 | 241 | Stand Off 1:
16 | 241 | Stand Off 2:
16 | 241 | Stand Off 3:
16 | 241 | Stand Off 4:
16 | 241 | Max Difference:
16 | 242 | Ensure objective coil is concentric to pole piece using alignment jig
17 | 301 | Position A:
17 | 301 | Position B:
17 | 301 | Position C:
17 | 301 | Position D:
17 | 301 | Max Difference:
17 | 304 | Visually inspected for damage/dents
17 | 304 | Sealing surfaces free from scratches/ imperfections
17 | 313 | Continuity ok
17 | 332 | Visual Inspection
17 | 332 | Liner Fitment
17 | 332 | End Cap Fitment
17 | 338 | Resistance:
17 | 361 | 12 hours drying time
16 | 244 | Open Circuit (Coil – GND).
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, then the PC needs reimaging – follow the attached WI. The correct image version is V1.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
18 | 392 | Continuity ok
18 | 393 | Visual Inspection
18 | 393 | Liner Fitment
18 | 393 | End Cap Fitment
18 | 395 | Objective Coil S/N :
18 | 395 | Scan Coil S/N :
18 | 395 | Pole Piece S/N :
18 | 397 | Check SCS Folder/Sales Order
18 | 397 | SCS Number (If applicable)
16 | 245 | Objective Coil
16 | 245 | Scan Coil
16 | 246 | Confirm there is no Electical breakdown between Scan coils and Objective body.
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, then the PC needs reimaging – follow the attached WI. The correct image version is V1.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
23 | 269 | Run relevant ESB conditioning MACRO
| 441 | S/N =
16 | 244 | Open Circuit (Coil – GND).
17 | 376 | Objective Coil S/N :
17 | 376 | Scan Coil S/N :
17 | 376 | Pole Piece S/N :
7 | 223 | Check SCS Folder/Sales Order
7 | 223 | SCS Number (If applicable)
17 | 313 | Continuity ok
20 | 252 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, then the PC needs reimaging – follow the attached WI. The correct image version is V1.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
30 | 303 | Actual
30 | 303 | Store stage drift macro results image
285 | 3242 | Select PC type
285 | 3242 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
285 | 3242 | PC S/No.
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, then the PC needs reimaging – follow the attached WI. The correct image version is V1.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Record MAC address
285 | 3242 | Network access requested
46 | 1567 | Remove the sample holder from the system.
46 | 1567 | Remove Docking Brackets and fit Transit clamps.
46 | 1567 | Airfeet deflated and free to move.
46 | 1567 | Backing lines re-configured so that PVSV valve is attached to flexi hose. Secure valve to frame.
46 | 1567 | Confirm that the 3 'o' rings are fitted to the flexi hose backing line
46 | 1567 | If applicable, ensure that the A/L push rod is removed and blank fitted in its place. Ensure A/L is pumped down before powering down system.
46 | 1567 | If applicable, ensure that the Beamsleeve funnel has been removed from the detector
46 | 1567 | Select 'Partial vent on standby' in Smartsem.
17 | 377 | Check SCS Folder/Sales Order
17 | 377 | SCS Number (If applicable)
17 | 377 | Check SCS Folder/Sales Order
17 | 377 | SCS Number (If applicable)
17 | 376 | Objective Coil S/N :
17 | 376 | Scan Coil S/N :
17 | 376 | Pole Piece S/N :
17 | 332 | Visual Inspection
17 | 332 | Liner Fitment
17 | 332 | End Cap Fitment
7 | 1650 | Objective Coil S/N :
7 | 1650 | Scan Coil S/N :
7 | 1650 | Pole Piece S/N :
17 | 313 | Continuity ok
29 | 11381 | Run overnight ESB Run Up & Photograph Macro
29 | 11381 | Analyse images for drift, breakdown, ESB issues
703 | 637 | Run 3 hour plasma clean
16 | 245 | Objective Coil
16 | 245 | Scan Coil
16 | 246 | Confirm there is no Electical breakdown between Scan coils and Objective body.
17 | 304 | Visually inspected for damage/dents
17 | 304 | Sealing surfaces free from scratches/ imperfections
17 | 301 | Position A:
17 | 301 | Position B:
17 | 301 | Position C:
17 | 301 | Position D:
17 | 301 | Max Difference:
17 | 274 | Resin is uniform and not proud
17 | 274 | No exposed wires
17 | 271 | Pressure 15 PSI for 5 minutes
7 | 74 | Visual Inspection
7 | 74 | Liner Fitment
7 | 74 | End Cap Fitment
25 | 398 | Are Gun Head clamping screws secure?
25 | 398 | Are Condensor clamping screws secure?
25 | 398 | Store 1kV emission image (Condensor off)
25 | 398 | Ensure 1kV On-Axis Stig Values
25 | 398 | Check emission image against 'emissions distortion criteria'
25 | 398 | Store 10kV emission image (Condensor off)
25 | 398 | Ensure 10kV On-Axis Stig Values
44 | 348 | Chamber Door Interlock
44 | 348 | Service Switch
44 | 348 | Detector De-activation
44 | 348 | EHT + IPPSU De-activation
44 | 454 | Interlock tested and working
44 | 354 | Check Installed Licences Against Sales Order
17 | 263 | Stand Off 1:
17 | 263 | Stand Off 2:
17 | 263 | Stand Off 3:
17 | 263 | Stand Off 4:
17 | 263 | Max Difference:
7 | 11258 | Continuity ok
7 | 210 | Visually inspected for damage/dents
7 | 210 | Sealing surfaces free from scratches/ imperfections
25 | 398 | Are Gun Head clamping screws secure?
25 | 398 | Are Condensor clamping screws secure?
25 | 398 | Store 1kV emission image (Condensor off)
25 | 398 | Actual 1kV X Stigs =
25 | 398 | Actual 1kV Y Stigs (Max 20%) =
25 | 398 | Check emission image against 'emissions distortion criteria'
25 | 398 | Store 10kV emission image (Condensor off)
25 | 398 | Ensure 10kV On-Axis Stig Values
25 | 398 | Are Gun Head clamping screws secure?
25 | 398 | Are Condensor clamping screws secure?
25 | 398 | Store 1kV emission image (Condensor off)
25 | 398 | Actual 1kV X Stigs (Max 20%) =
25 | 398 | Actual 1kV Y Stigs (Max 20%) =
25 | 398 | Check emission image against 'emissions distortion criteria'
25 | 398 | Store 10kV emission image (Condensor off)
25 | 398 | Ensure 10kV On-Axis Stig Values
443
443
26 | 10965 | ESB Serial Number.
17 | 251 | Ensure objective coil is concentric to pole piece using alignment jig
7 | 1640 | Position A:
7 | 1640 | Position B:
7 | 1640 | Position C:
7 | 1640 | Position D:
7 | 1640 | Max Difference:
25 | 398 | Are Gun Head clamping screws secure?
25 | 398 | Are Condensor clamping screws secure?
25 | 398 | Store 1kV emission image (Condensor off)
25 | 398 | Actual 1kV X Stigs (Max 20%) =
25 | 398 | Actual 1kV Y Stigs (Max 20%) =
25 | 398 | Check emission image against 'emissions distortion criteria'
25 | 398 | Store 10kV emission image (Condensor off)
25 | 398 | Actual 10kV X Stigs (Max 20%) =
25 | 398 | Actual 10kV Y Stigs (Max 20%) =
25 | 401 | Stig: 30μm, All EHT’s is ≤ 20%
25 | 401 | Stig: Off Axis Aper, All EHT’s is ≤ 30%
25 | 401 | Aper/Gun 30μm, All EHT’s is ≤ 60%
20 | 252 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, then the PC needs reimaging – follow the attached WI. The correct image version is V1.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
17 | 248 | Objective Coil
17 | 248 | Scan Coil
17 | 247 | Confirm there is no Electical breakdown between Scan coils and Objective body.
7 | 11340 | Resin is uniform and not proud
7 | 11340 | No exposed wires
7 | 82 | Pressure 15 PSI for 5 minutes
25 | 399 | Aperture Finder run succesfully
25 | 399 | Stig Check 1kV X (Max 20%) =
25 | 399 | Stig Check 1kV Y (Max 20%) =
25 | 401 | Stig: 30μm, 99V, 10kV, 20kV is ≤ 20%
25 | 401 | Stig: Off Axis Aper, All EHT’s is ≤ 30%
25 | 401 | Aper/Gun 30μm, All EHT’s is ≤ 60%
25 | 401 | Stig: 30μm, 99V, 10kV, 20kV is ≤ 20%
25 | 401 | Stig: Off Axis Aper, 10kV is ≤ 30%
25 | 401 | Aper/Gun 30μm, All EHT’s is ≤ 60%
25 | 401 | Stig: 30μm, 99V, 10kV, 20kV is ≤ 20%
25 | 401 | Stig: Off Axis Aper, 10kV is ≤ 30%
25 | 401 | Aper/Gun 30μm, 10kV, 500V, 20kV is ≤ 60%
25 | 401 | Stig: 30μm, 99V, 10kV, 20kV is ≤ 20%
25 | 401 | Stig: Off Axis Aper, 10kV is ≤ 30%
25 | 401 | Aper/Gun 30μm, 500V, 10kV, 20kV is ≤ 60%
403 | 402 | Fit labels to PC serial cable com port 4 to 7
285 | 3242 | Select PC type
285 | 3242 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
285 | 3242 | PC S/No.
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, then the PC needs reimaging – follow the attached WI. The correct image version is V1.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Record MAC address
285 | 3242 | Network access requested
50 | 481 | VP Tablefiller Entries successfully completed.
703 | 637 | Run 3 hour plasma clean
17 | 249 | Open Circuit (Coil – GND).
7 | 1641 | Stand Off 1:
7 | 1641 | Stand Off 2:
7 | 1641 | Stand Off 3:
7 | 1641 | Stand Off 4:
7 | 1641 | Max Difference:
7 | 211 | Ensure objective coil is concentric to pole piece using alignment jig
7 | 83 | Open Circuit (Coil – GND).
25 | 399 | Aperture Finder run succesfully
25 | 399 | Stig Check 1kV X (Max 20%) =
25 | 399 | Stig Check 1kV Y (Max 20%) =
25 | 399 | Actual 10kV X Stigs (Max 20%) =
25 | 399 | Actual 10kV Y Stigs (Max 20%) =
25 | 398 | Align offset angle set to -110 degrees.
27 | 289 | C2 Calibration Value
27 | 289 | C3 Calibration Value
27 | 289 | Are Calibration factors typical? (see table of typical values in WI)
27 | 430 | C1 Cal =
27 | 430 | C2 Cal =
27 | 430 | C3 Cal =
27 | 430 | Are Calibration factors typical? (see table of typical values in WI)
403 | 6935 | Record MAC address
403 | 6935 | Network access requested
403 | 6935 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
23 | 269 | Run ESB conditioning MACRO
13 | 11191 | Driver A In – Anode (>1000MΩ)
13 | 11191 | Driver A In – GND (>1000MΩ)
13 | 11191 | Anode - GND (>1000MΩ)
50 | 480 | DC Tablefiller successfully completed.
392 | 3342 | No pipes are touching cladding.
392 | 3342 | Cladding is secured correctly and not able to move.
392 | 3342 | Safety stickers are applied correctly and are the correct language.
17 | 313 | Continuity ok
7 | 1642 | Objective Coil
7 | 1642 | Scan Coil
7 | 11393 | Confirm there is no Electical breakdown between Scan coils and Objective body.
25 | 399 | Aperture Finder run succesfully
25 | 399 | Stig Check 1kV X (Max 20%) =
25 | 399 | Stig Check 1kV Y (Max 20%) =
25 | 399 | Stig Check 10kV X (Max 20%) =
25 | 399 | Stig Check 10kV Y (Max 20%) =
25 | 398 | Align offset angle set to -110 degrees.
403 | 10465 | Windows 10 Version
403 | 6935 | Record MAC address
403 | 6935 | Network access requested
403 | 6935 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
5 | 423 | Verify the SVSEM cables are connected to the customer port configuration
46 | 1568 | Power system down.
46 | 1568 | Blow water out of the cooling system.
46 | 1568 | Chamber secured to vibration plate
46 | 1568 | Chamber Sercuring Screws Tightened to 25NM
46 | 1568 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 1568 | Secure Airfeet with foam blocks.
46 | 1568 | Fit foam blocks around column cladding.
46 | 1568 | Secure Turbo pump with foam blocks.
46 | 1568 | Fit Covers to all unused connectors.
46 | 1568 | Remove Cekon Plug From Mains Cable
46 | 1568 | Fit Live and Neutral Labels to Mains Cable
46 | 1568 | Mains and Earth cable stowed on Mains unit.
46 | 1568 | Fit shipping Eyes to transit blocks.
46 | 1568 | Confirm all column cladding is secure and screws Loctited in place.
46 | 1568 | Put thread lock on pendulum screws and transit block Screws.
46 | 1568 | All doors closed and secured.
46 | 1568 | All plinth panels removed and stored in transit trolley.
46 | 1568 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 1568 | Print shipping stickers.
46 | 1568 | Fit M8 screw to secure stage door to chamber.
23 | 269 | Run relevant ESB conditioning MACRO
50 | 481 | VP Tablefiller Entries successfully completed.
18 | 397 | Check SCS Folder/Sales Order
18 | 397 | SCS Number (If applicable)
25 | 399 | Aperture Finder run succesfully
25 | 275 | Are Gun Head clamping screws secure?
25 | 275 | Are Condensor clamping screws secure?
25 | 275 | Store 1kV emission image (Condensor off)
25 | 275 | Check emission image against 'emissions distortion criteria'
25 | 275 | Store 10kV emission image (Condensor off)
5 | 423 | Verify the SVSEM cables are connected to the customer port configuration
5 | 423 | Verify the SVSEM cables are connected to the customer port configuration
46 | 1568 | Power system down.
46 | 1568 | Blow water out of the cooling system.
46 | 1568 | Chamber secured to vibration plate
46 | 1568 | Chamber Sercuring Screws Tightened to 25NM
46 | 1568 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 1568 | Secure Airfeet with foam blocks.
46 | 1568 | Fit foam blocks around column cladding.
46 | 1568 | Secure Turbo pump with foam blocks.
46 | 1568 | Fit Covers to all unused connectors.
46 | 1568 | Remove Cekon Plug From Mains Cable
46 | 1568 | Fit Live and Neutral Labels to Mains Cable
46 | 1568 | Mains and Earth cable stowed on Mains unit.
46 | 1568 | Fit shipping Eyes to transit blocks.
46 | 1568 | Confirm all column cladding is secure and screws Loctited in place.
46 | 1568 | Put thread lock on pendulum screws and transit block Screws.
46 | 1568 | All doors closed and secured.
46 | 1568 | All plinth panels removed and stored in transit trolley.
46 | 1568 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 1568 | Print shipping stickers.
46 | 1568 | Fit M8 screw to secure stage door to chamber.
23 | 269 | Run relevant ESB conditioning MACRO
49 | 482 | DC Tablefiller successfully completed.
717 | 638 | Run 3 hour plasma clean
18 | 397 | Check SCS Folder/Sales Order
18 | 397 | SCS Number (If applicable)
18 | 395 | Objective Coil S/N :
18 | 395 | Scan Coil S/N :
18 | 395 | Pole Piece S/N :
18 | 393 | Visual Inspection
18 | 393 | Liner Fitment
18 | 393 | End Cap Fitment
49 | 482 | DC Tablefiller successfully completed.
37 | 11233 | 10kv VP
38 | 335 | Measured Specimen Current @ 20kV =
38 | 335 | Store Optiprobe log
5 | 423 | Confirm Liner cable routed as specified
30 | 1230 | Stage mapping completed successfuly
667 | 11414 | Confirm label has CSA mark and SN and PN are correct.
37 | 328 | Tablefiller successfully completed.
6 | 1252 | Check interlock keys fitted to door
5 | 1256 | Fit SCM
18 | 397 | Check SCS Folder/Sales Order
18 | 397 | SCS Number (If applicable)
18 | 395 | Objective Coil S/N :
18 | 395 | Scan Coil S/N :
18 | 395 | Pole Piece S/N :
18 | 393 | Visual Inspection
18 | 393 | Liner Fitment
18 | 393 | End Cap Fitment
| 596 | Confirm there is no Electical breakdown between Scan coils and Objective body.
37 | 11314 | Tablefiller successfully completed.
46 | 1568 | Power system down.
46 | 1568 | Blow water out of the cooling system.
46 | 1568 | Put thread lock on pendulum screws and transit block Screws.
46 | 1568 | Chamber secured to vibration plate
46 | 1568 | Chamber Sercuring Screws Tightened to 25NM
46 | 1568 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 1568 | Secure Airfeet with foam blocks.
46 | 1568 | Fit foam blocks around column cladding.
46 | 1568 | Secure Turbo pump with foam blocks.
46 | 1568 | Fit Covers to all unused connectors.
46 | 1568 | Remove Cekon Plug From Mains Cable
46 | 1568 | Fit Live and Neutral Labels to Mains Cable
46 | 1568 | Mains and Earth cable stowed on Mains unit.
46 | 1568 | Fit shipping Eyes to transit blocks.
46 | 1568 | Confirm all column cladding is secure and screws Loctited in place.
46 | 1568 | All doors closed and secured.
46 | 1568 | All plinth panels removed and stored in transit trolley.
46 | 1568 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 1568 | Print shipping stickers.
46 | 1568 | Fit M8 screw to secure stage door to chamber.
19 | 408 | Actual =
19 | 408 | After conditioning confirm there are no EMServer error messages
19 | 408 | Store Gun monitor trace as Probestability.mn2
37 | 328 | Tablefiller successfully completed.
18 | 392 | Continuity ok
25 | 398 | Are Gun Head clamping screws secure?
25 | 398 | Are Condensor clamping screws secure?
25 | 398 | Store 1kV emission image (Condensor off)
25 | 398 | Actual 1kV X Stigs (Max 20%) =
25 | 398 | Actual 1kV Y Stigs (Max 20%) =
25 | 398 | Check emission image against 'emissions distortion criteria'
25 | 398 | Store 10kV emission image (Condensor off)
25 | 398 | Actual 10kV X Stigs (Max 20%) =
25 | 398 | Actual 10kV Y Stigs (Max 20%) =
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
648 | 10553 | 30KV Res- ≤2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 1KV Res- ≤15nm
648 | 10553 | 200V Res- Not measured
648 | 10553 | Store Standard Resolution Images
6 | 921 | Record Water Flow Temperature (20ºC +/- 4ºC)
389 | 924 | Fit to Stage Port
18 | 391 | Visually inspected for damage/dents
18 | 391 | Sealing surfaces free from scratches/ imperfections
648 | 10611 | 30KV Res- ≤2nm
648 | 10611 | 3KV Res- ≤6nm
648 | 10611 | 1KV Res- ≤9nm
648 | 10611 | 200V Res- Not Measured
648 | 10611 | Store Standard LaB6 Resolution Images
| 922 | Remove heat shrink from objective coil cooling pipes
37 | 11110 | 99v
37 | 11110 | 22kv
37 | 11110 | 30kv
707 | 11187 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11187 | Inspection Passed and Images Saved
707 | 11187 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11187 | Inspection Passed and Images Saved
49 | 482 | DC Tablefiller successfully completed .
20 | 407 | Fit labels to PC serial cable com port 4 to 7
20 | 407 | Fit labels to PC serial cable com port 4 to 7
18 | 390 | Position A:
18 | 390 | Position B:
18 | 390 | Position C:
18 | 390 | Position D:
18 | 390 | Max Difference:
388 | 11425 | 15kv 0.9nm
388 | 11425 | 15KV result
388 | 11425 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 1kv 1.3nm
388 | 11425 | 1KV result
388 | 11425 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 500V 2nm
388 | 11425 | 500V result
388 | 11425 | Store 500V Resolution Images and Pate image for 2 x consecutive images
44 | 348 | Chamber Door Interlock
44 | 348 | Service Switch
44 | 348 | Lens De-activation
44 | 348 | Detector De-activation
44 | 348 | EHT + IPPSU De-activation
25 | 197 | Gun X
25 | 197 | Gun Y
25 | 197 | Aper X
25 | 197 | Aper Y
389 | 11074 | Check each individual segment in normal and inverse modes
389 | 11074 | Analyze Segments
389 | 11074 | Gain Ranges
389 | 11074 | Pneumatics
389 | 11074 | Alignment
389 | 11074 | Vac Degradation < 3x10-5 mbar
389 | 11074 | CZ Detector Cal
389 | 11074 | Z Adjustment
707 | 11277 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11277 | Use HR conditions (6144x4608, N=5)
707 | 11277 | Save to TIFF as Multi_TS_HRQA
707 | 11277 | Inspection Passed and Images Saved
32 | 444 | Actual @ 1kV WD5 =
32 | 444 | Actual @ 20kV WD2 =
18 | 386 | Resin is uniform and not proud
18 | 386 | No exposed wires
18 | 384 | Pressure 15 PSI for 5 minutes
18 | 383 | Stand Off 1:
18 | 383 | Stand Off 2:
18 | 383 | Stand Off 3:
18 | 383 | Stand Off 4:
18 | 383 | Max Difference:
18 | 382 | Ensure objective coil is concentric to pole piece using alignment jig
18 | 380 | Open Circuit (Coil – GND).
18 | 379 | Objective Coil
18 | 379 | Scan Coil
18 | 378 | Confirm there is no Electical breakdown between Scan coils and Objective body.
388 | 11426 | 15kv 0.7nm
388 | 11426 | 15KV result
388 | 11426 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 1kv 1.2nm
388 | 11426 | 1KV result
388 | 11426 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 500V 1.9nm
388 | 11426 | 500V result
388 | 11426 | Store 500V Resolution Images and Pate image for 2 x consecutive images
37 | 328 | Tablefiller successfully completed.
40 | 1178 | SE Detector
40 | 1178 | SE Detector has no light leak.
285 | 11101 | Check SCS Folder/Sales Order
285 | 11101 | SCS Number (If applicable)
7 | 1650 | Objective Coil S/N :
7 | 1650 | Scan Coil S/N :
7 | 1650 | Pole Piece S/N :
648 | 10554 | 2KV Res - 100um VP Aperture- ≤16nm
648 | 10554 | 2KV Res - 400um Easy VP Aperture- ≤16nm
648 | 10554 | Store VP & Easy VP Aperture Resolution Images
25 | 197 | Gun X
25 | 197 | Gun Y
25 | 197 | Aper X
25 | 197 | Aper Y
37 | 328 | Tablefiller successfully completed.
389 | 11074 | Check each individual segment in normal and inverse modes
389 | 11074 | Analyze Segments
389 | 11074 | Gain Ranges
389 | 11074 | Pneumatics
389 | 11074 | Alignment
389 | 11074 | Vac Degradation < 3x10-5 mbar
389 | 11074 | CZ Detector Cal
389 | 11074 | Z Adjustment
663 | 414 | No Leaks Detected with GROSS TEST MODE off
31 | 483 | Aperture Finder run succesfully
707 | 11277 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11277 | Inspection Passed and Images Saved
707 | 11283 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11283 | Inspection Passed and Images Saved
7 | 74 | Visual Inspection
7 | 74 | Liner Fitment
7 | 74 | End Cap Fitment
648 | 10612 | 2KV Res - 100um VP Apertutre- ≤12nm
648 | 10612 | Store VP Aperture (100u) Resolution Images
646 | 10545 | Correct PC - 347823-9024-660 Installed
646 | 10545 | Print Out SmartPI Check List (Link in Documents)
646 | 10545 | Install SmartSEM Licences
646 | 10545 | Install MIL Redistribution
646 | 10545 | Fit MIL Dongle
646 | 10545 | SmartPI Software Version
646 | 10545 | Install SmartPI Software
646 | 10545 | I/O-SCAN Card Fitted (Bruker Only)
646 | 10545 | I/O-SCAN Card Calibrated
646 | 10545 | Setup & Run Application
646 | 10545 | Check Results & Image Montage
646 | 10545 | Store Compressed Database,Montage Image, Log files etc
37 | 328 | Tablefiller successfully completed.
667 | 10773 | No Leaks Detected with GROSS TEST MODE off
52 | 485 | Stig: 30μm, 99V, 10kV, 20kV is ≤ 20%
52 | 485 | Stig: Off Axis Aper, 10kV is ≤ 30%
52 | 485 | Aper/Gun 30μm, 500V, 10kV, 20kV is ≤ 60%
| 487 | Objective Coil S/N :
| 487 | Scan Coil S/N :
| 487 | Pole Piece S/N :
7 | 11258 | Continuity ok
648 | 10881 | 2KV Res - 750um Easy VP Aperture- ≤12nm
648 | 10881 | Store Easy VP Aperture (750u) Resolution Images
648 | 10614 | 2KV Res - 750um Easy Align Aperture- ≤12nm
648 | 10614 | Store Easy Align Aperture (750u) Resolution Images
648 | 10657 | 2KV - 100um VP Aperture- ≤16nm
648 | 10657 | 2KV - 140um VP Aperture- ≤16nm
648 | 10657 | Store 2KV Aperture Resolution Images
648 | 10658 | 2KV - 100um VP Aperture- ≤12nm
648 | 10658 | 2KV - 140um VP Aperture- ≤12nm
648 | 10658 | Store 2KV Aperture Resolution Images
37 | 11110 | 99v
37 | 11110 | 22kv
37 | 11110 | 30kv
663 | 414 | No Leaks Detected (If using Agilent leaker tester GROSS TEST MODE must be off)
52 | 484 | Aperture Finder run succesfully
7 | 210 | Visually inspected for damage/dents
7 | 210 | Sealing surfaces free from scratches/ imperfections
7 | 1640 | Position A:
7 | 1640 | Position B:
7 | 1640 | Position C:
7 | 1640 | Position D:
7 | 1640 | Max Difference:
7 | 11340 | Resin is uniform and not proud
7 | 11340 | No exposed wires
7 | 82 | Pressure 15 PSI for 5 minutes
649 | 230 | BSD Resolution ≤4nm.
649 | 230 | Store BSD Resolution Image
30 | 303 | Actual
30 | 303 | Store stage drift macro results image
667 | 10773 | No Leaks Detected (If using Agilent leaker tester GROSS TEST MODE must be off)
| 11379 | Continuity (Close Circuit)
| 11379 | Using Megger tester check there is not short circuit between lens and the condenser casing
7 | 1641 | Stand Off 1:
7 | 1641 | Stand Off 2:
7 | 1641 | Stand Off 3:
7 | 1641 | Stand Off 4:
7 | 1641 | Max Difference:
7 | 211 | Ensure objective coil is concentric to pole piece using alignment jig
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
649 | 10556 | 30KV Resolution - Easy VP- ≤3.4nm
649 | 10556 | Store VPSE Resolution Image
649 | 204 | VPSE Resolution 3.0nm
649 | 204 | Store VPSE Resolution Image
649 | 11206 | 30KV Res- ≤2nm
649 | 11206 | Store 30KV ( Large ZTR)Resolution Image
649 | 10659 | 30KV Resolution- ≤5.5nm
649 | 10659 | Store VPSE Resolution Image
649 | 10557 | Resolution @ 7KV- ≤9nm
649 | 10557 | Store Beam Deceleration Resolution Image
37 | 461 | 50v
37 | 461 | 21kv
37 | 461 | 30kv
37 | 461 | 50v - All Aperture at Normal and High current WD3
37 | 461 | 21k - All Apertures at Normal and High current WD5
37 | 461 | 30kv - Aperture 1 and 7 at Normal and High current at WD5
| 463 | S/N =
| 464 | Objective Coil S/N :
| 464 | Scan Coil S/N :
| 464 | Pole Piece S/N :
4 | 11384 | No Leaks Detected (If using Agilent leaker tester GROSS TEST MODE must be off)
707 | 11284 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11284 | Inspection Passed and Images Saved
707 | 11277 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11277 | Inspection Passed and Images Saved
10 | 676 | S/N =
7 | 83 | Open Circuit (Coil – GND).
7 | 11393 | Confirm there is no Electical breakdown between Scan coils and Objective body.
| 463 | S/N =
35 | 319 | Aperture Finder successfully run
6 | 68 | No Leaks Detected (If using Agilent leaker tester GROSS TEST MODE must be off)
| 490 | Visual Inspection
| 490 | Liner Fitment
| 490 | End Cap Fitment
707 | 11283 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11283 | Inspection Passed and Images Saved
707 | 11284 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11284 | Inspection Passed and Images Saved
707 | 11187 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11187 | Inspection Passed and Images Saved
7 | 1642 | Objective Coil
7 | 1642 | Scan Coil
| 466 | Visual Inspection
| 466 | Liner Fitment
| 466 | End Cap Fitment
| 467 | Continuity ok
6 | 68 | No Leaks Detected (If using Agilent leaker tester GROSS TEST MODE must be off)
| 491 | Continuity ok
707 | 11283 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11283 | Inspection Passed and Images Saved
20 | 252 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, then the PC needs reimaging – follow the attached WI. The correct image version is V1.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Run ASAP Screening Program
20 | 252 | System Specific Datazone Loaded
678 | 621 | Confirm PAT test passed and results stored
18 | 395 | Objective Coil S/N :
18 | 395 | Scan Coil S/N :
18 | 395 | Pole Piece S/N :
| 492 | Visually inspected for damage/dents
| 492 | Sealing surfaces free from scratches/ imperfections
707 | 11284 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11284 | Inspection Passed and Images Saved
389 | 115 | Setup Ethernet Connection
389 | 115 | Install Aztec Software
389 | 115 | Aztec Software Version
389 | 115 | Licence Installation
389 | 115 | Run Aztec Assist+
389 | 115 | Install Detector
389 | 115 | Imaging Setup
389 | 115 | Stage Calibration Setup
389 | 115 | Calibrate Detector
389 | 115 | Save Aztec Sign Off PDF and Calibration Folder
23 | 269 | Run ESB conditioning MACRO
23 | 269 | Run ESB conditioning MACRO prior to switching on EHT.
23 | 269 | Run ESB conditioning MACRO prior to switching on EHT.
707 | 11170 | VP_TS_HRQA - Passed inspection criteria, Image Saved
18 | 395 | Objective Coil S/N :
18 | 395 | Scan Coil S/N :
18 | 395 | Pole Piece S/N :
18 | 393 | Visual Inspection
18 | 393 | Liner Fitment
18 | 393 | End Cap Fitment
18 | 392 | Continuity ok
18 | 391 | Visually inspected for damage/dents
18 | 391 | Sealing surfaces free from scratches/ imperfections
| 468 | Visually inspected for damage/dents
| 468 | Sealing surfaces free from scratches/ imperfections
| 494 | Resin is uniform and not proud
| 494 | No exposed wires
707 | 11283 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11283 | Inspection Passed and Images Saved
707 | 11286 | Ensure size of hole is 20um (+/- 2um)
707 | 11286 | Inspection Passed and Images Saved
18 | 390 | Position A:
18 | 390 | Position B:
18 | 390 | Position C:
18 | 390 | Position D:
18 | 390 | Max Difference:
18 | 388 | Position A:
18 | 388 | Position B:
18 | 388 | Position C:
18 | 388 | Position D:
18 | 388 | Max Difference:
| 469 | Position A:
| 469 | Position B:
| 469 | Position C:
| 469 | Position D:
| 469 | Max Difference:
| 493 | Position A:
| 493 | Position B:
| 493 | Position C:
| 493 | Position D:
| 493 | Max Difference:
707 | 11252 | Ensure size of hole is 30um (+/- 2um)
707 | 11252 | Inspection Passed and Images Saved
707 | 11286 | Ensure size of hole is 20um (+/- 2um)
707 | 11286 | Inspection Passed and Images Saved
707 | 11252 | Ensure size of hole is 30um (+/- 2um)
707 | 11252
707 | 11252 | Inspection Passed and Images Saved
18 | 386 | Resin is uniform and not proud
18 | 386 | No exposed wires
18 | 384 | Pressure 15 PSI for 5 minutes
18 | 383 | Stand Off 1:
18 | 383 | Stand Off 2:
18 | 383 | Stand Off 3:
18 | 383 | Stand Off 4:
18 | 383 | Max Difference:
18 | 382 | Ensure objective coil is concentric to pole piece using alignment jig
| 495 | Pressure 15 PSI for 5 minutes
| 496 | Stand Off 1:
| 496 | Stand Off 2:
| 496 | Stand Off 3:
| 496 | Stand Off 4:
| 496 | Max Difference:
707 | 11285 | Ensure size of hole is 30um (+/- 2um)
707 | 11285 | Inspection Passed and Images Saved
707 | 11156 | Ensure size of hole is 30um (+/- 2um)
707 | 11156 | Inspection Passed and Images Saved
18 | 380 | Open Circuit (Coil – GND).
18 | 379 | Objective Coil
18 | 379 | Scan Coil
18 | 378 | Confirm there is no Electical breakdown between Scan coils and Objective body.
| 470 | Resin is uniform and not proud
| 470 | No exposed wires
| 497 | Ensure objective coil is concentric to pole piece ( this should be done as best as possible by eye as no jig yet available)
707 | 11160 | Ensure size of hole is 30um (+/- 2um)
707 | 11160 | Inspection Passed and Images Saved
707 | 11287 | Ensure size of hole is 30um (+/- 2um)
707 | 11287 | Inspection Passed and Images Saved
707 | 11252 | Ensure size of hole is 30um (+/- 2um)
707 | 11252 | Inspection Passed and Images Saved
17 | 376 | Objective Coil S/N :
17 | 376 | Scan Coil S/N :
17 | 376 | Pole Piece S/N :
| 499 | Open Circuit (Coil – GND).
707 | 11156 | Ensure size of hole is 20um (+/- 2um)
707 | 11156 | Inspection Passed and Images Saved
707 | 11252 | Measure aperture size, 30um (+/- 2um)
707 | 11252 | Inspection Passed and Images Saved
707 | 11283 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11283 | Multi_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11284 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11284 | Multi_TS_HRQA - Passed inspection criteria, Image Saved
38 | 335 | Measured Specimen Current @ 20kV =
38 | 335 | Store Optiprobe log
| 540 | S/N =
17 | 361 | 12 hours drying time
17 | 338 | Resistance:
17 | 332 | Visual Inspection
17 | 332 | Liner Fitment
17 | 332 | End Cap Fitment
| 471 | Pressure 15 PSI for 5 minutes
| 500 | Objective Coil
| 500 | Scan Coil
| 501 | Confirm there is no Electical breakdown between Scan coils and Objective body.
707 | 11286 | Ensure size of hole is 20um (+/- 2um)
707 | 11286 | Inspection Passed and Images Saved
707 | 11160 | Ensure size of hole is 120um (+/- 2um)
707 | 11160 | Inspection Passed and Images Saved
707 | 11287 | Ensure size of hole is 120um (+/- 2um)
707 | 11287 | Inspection Passed and Images Saved
17 | 304 | Visually inspected for damage/dents
17 | 304 | Sealing surfaces free from scratches/ imperfections
| 472 | Stand Off 1:
| 472 | Stand Off 2:
| 472 | Stand Off 3:
| 472 | Stand Off 4:
| 472 | Max Difference:
707 | 11288 | Ensure size of hole is 300um (+/- 10um)
707 | 11288 | Inspection Passed and Images Saved
707 | 11285 | Measure aperture size, 30um (+/- 2um)
707 | 11285 | Inspection Passed and Images Saved
707 | 11286 | Measure aperture size, 20um (+/- 2um)
707 | 11286 | Multi_TS_20um - Passed inspection criteria, Image Saved
707 | 11286 | Measure aperture size, 20um (+/- 2um)
707 | 11286 | Multi_TS_20um - Passed inspection criteria, Image Saved
| 541 | Housing Type
| 541 | Inspect insulation disc using white light for cracks
17 | 301 | Position A:
17 | 301 | Position B:
17 | 301 | Position C:
17 | 301 | Position D:
17 | 301 | Max Difference:
17 | 288 | Position A:
17 | 288 | Position B:
17 | 288 | Position C:
17 | 288 | Position D:
17 | 288 | Max Difference:
17 | 274 | Resin is uniform and not proud
17 | 274 | No exposed wires
17 | 263 | Stand Off 1:
17 | 263 | Stand Off 2:
17 | 263 | Stand Off 3:
17 | 263 | Stand Off 4:
17 | 263 | Max Difference:
17 | 251 | Ensure objective coil is concentric to pole piece using alignment jig
| 502 | S/N =
707 | 11159 | Ensure size of hole is 300um (+/- 10um)
707 | 11159 | Inspection Passed and Images Saved
707 | 11156 | Measure aperture size, 20um (+/- 2um)
707 | 11156 | Inspection Passed and Images Saved
707 | 11286 | Measure aperture size, 30um (+/- 2um)
707 | 11286 | Inspection Passed and Images Saved
17 | 271 | Pressure 15 PSI for 5 minutes
| 473 | Ensure objective coil is concentric to pole piece using alignment jig
33 | 316 | Actual for 1kV =
33 | 316 | Actual for 15kV =
33 | 316 | Actual for 1kV =
33 | 316 | Actual for 15kV =
33 | 316 | Actual for 1kV =
33 | 316 | Actual for 15kV =
33 | 316 | Actual for 1kV =
33 | 316 | Actual for 15kV =
33 | 316 | Actual for 1kV =
33 | 316 | Actual for 15kV =
33 | 316 | Actual for 1kV =
33 | 316 | Store 1KV Resolution Images and Pate image
33 | 316 | Actual for 15kV =
33 | 316 | Store 15KV Resolution Images and Pate image
32 | 444 | Actual @ 1kV WD5 =
32 | 444 | Actual @ 20kV WD2 =
17 | 249 | Open Circuit (Coil – GND).
17 | 248 | Objective Coil
17 | 248 | Scan Coil
17 | 247 | Confirm there is no Electical breakdown between Scan coils and Objective body.
707 | 11252 | Ensure size of hole is 30um (+/- 2um)
707 | 11252 | Inspection Passed and Images Saved
707 | 11160 | Measure aperture size, 30um (+/- 2um)
707 | 11160 | Inspection Passed and Images Saved
| 544 | Operational repeatability test x 3
20 | 252 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, then the PC needs reimaging – follow the attached WI. The correct image version is V1.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Run ASAP Screening Program
20 | 252 | System Specific Datazone Loaded
28 | 700 | Confirm EM Server screeening programme is installed and running
28 | 700 | Confirm no EM Server screening errors have been flagged
| 475 | Open Circuit (Coil – GND).
707 | 11285 | Ensure size of hole is 30um (+/- 2um)
707 | 11285 | Inspection Passed and Images Saved
707 | 11160 | Measure aperture size, 120um (+/- 2um)
707 | 11160 | Inspection Passed and Images Saved
| 546 | Pneumatic Operation
| 546 | Pin Touching when Retracted
19 | 408 | Actual =
19 | 408 | After conditioning confirm there are no EMServer error messages
19 | 408 | Store Gun monitor trace as Probestability.mn2
| 464 | Objective Coil S/N :
| 464 | Scan Coil S/N :
| 464 | Pole Piece S/N :
| 469 | Position A:
| 469 | Position B:
| 469 | Position C:
| 469 | Position D:
| 469 | Max Difference:
| 470 | Resin is uniform and not proud
| 470 | No exposed wires
| 473 | Ensure objective coil is concentric to pole piece using alignment jig
| 476 | Objective Coil
| 476 | Scan Coil
707 | 11156 | Ensure size of hole is 20um (+/- 2um)
707 | 11156 | Inspection Passed and Images Saved
707 | 11287 | Measure aperture size, 120um (+/- 2um)
707 | 11287 | Inspection Passed and Images Saved
| 468 | Visually inspected for damage/dents
| 468 | Sealing surfaces free from scratches/ imperfections
| 471 | Pressure 15 PSI for 5 minutes
| 475 | Open Circuit (Coil – GND).
| 476 | Objective Coil
| 476 | Scan Coil
| 477 | Confirm there is no Electical breakdown between Scan coils and Objective body.
32 | 622 | Actual @ 1kV WD5 =
32 | 622 | Actual @ 20kV WD2 =
32 | 622 | Actual @ 1kV WD5 =
32 | 622 | Actual @ 20kV WD2 =
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | Confirm the following tests were succesfully run as part of the Auto Routine:................................................ PE Test, Insulation Test (520V @ 5sec)
| 477 | Confirm there is no Electical breakdown between Scan coils and Objective body.
707 | 11252 | Ensure size of hole is 30um (+/- 2um)
707 | 11252 | Inspection Passed and Images Saved
707 | 11285 | Ensure size of hole is 30um (+/- 2um)
707 | 11285 | Inspection Passed and Images Saved
707 | 11287 | Measure aperture size, 120um (+/- 2um)
707 | 11287 | Inspection Passed and Images Saved
707 | 11287 | Measure aperture size, 120um (+/- 2um)
707 | 11287 | Inspection Passed and Images Saved
| 560 | S/N =
| 11373 | S/N =
| 11373 | S/N =
| 467 | Continuity ok
| 472 | Stand Off 1:
| 472 | Stand Off 2:
| 472 | Stand Off 3:
| 472 | Stand Off 4:
| 472 | Max Difference:
707 | 11156 | Ensure size of hole is 20um (+/- 2um)
707 | 11156 | Inspection Passed and Images Saved
707 | 11160 | Measure aperture size, 120um (+/- 2um)
707 | 11160 | Inspection Passed and Images Saved
707 | 11160 | Measure aperture size, 120um (+/- 2um)
707 | 11160 | Inspection Passed and Images Saved
| 466 | Visual Inspection
| 466 | Liner Fitment
| 466 | End Cap Fitment
32 | 622 | Actual @ 1kV WD5 =
32 | 622 | Actual @ 20kV WD2 =
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | Confirm the following tests were succesfully run as part of the Auto Routine:................................................ PE Test <0.1, Insulation Test (520V @ 5sec)
667 | 10765 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
667 | 10765 | HAL SN:
667 | 10765 | HAL Calibration Date:
667 | 10765 | DC IR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
667 | 10765 | DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
667 | 10765 | Confirm System Passed Hi-Pot Tests
6 | 710 | Actual
707 | 11288 | Measure aperture size, 300um (+/- 2um)
707 | 11288 | Inspection Passed and Images Saved
707 | 11290 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11290 | Multi_US_HRQA - Passed inspection criteria, Image Saved
707 | 11308 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11308 | Multi_US_HRQA - Passed inspection criteria, Image Saved
707 | 11165 | Ensure size of hole is 40um (+/- 2um)
707 | 11165 | Use LR conditions (1024x768, N=1)
707 | 11165 | Save to TIFF as Anode_TS_40um
707 | 11165 | Inspection Passed and Images Saved
707 | 11166 | Ensure size of hole is 55um (+/- 2um
707 | 11166 | Use LR conditions (1024x768, N=1)
707 | 11166 | Save to TIFF as Anode_TS_55um
707 | 11166 | Inspection Passed and Images Saved
| 560 | S/N =
| 563 | Housing Type
| 563 | Inspect insulation disc using white light for cracks
707 | 11288 | Measure aperture size, 300um (+/- 10um)
707 | 11288 | Inspection Passed and Images Saved
707 | 11262 | (There is no specific checks not regarding the search for particles)
707 | 11262 | Use HR conditions (6144x4608, N=5)
707 | 11262 | Save to TIFF as Anode_TS_HRQA
707 | 11262 | Inspection Passed and Images Saved
707 | 11165 | Ensure size of hole is 40um (+/- 2um)
707 | 11165 | Use LR conditions (1024x768, N=1)
707 | 11165 | Save to TIFF as Anode_TS_40um
707 | 11165 | Inspection Passed and Images Saved
707 | 11166 | Ensure size of hole is 55um (+/- 2um
707 | 11166 | Inspection Passed and Images Saved
707 | 11166 | Ensure size of hole is 55um (+/- 2um
707 | 11166 | Anode_TS_40um - Passed inspection criteria, Image Saved
| 574 | Housing Type
| 574 | Inspect insulation disc using white light for cracks
707 | 11288 | Measure aperture size, 300um (+/- 10um)
707 | 11288 | Inspection Passed and Images Saved
707 | 11288 | Measure aperture size, 300um (+/- 10um)
707 | 11288 | Inspection Passed and Images Saved
707 | 11262 | (There is no specific checks not regarding the search for particles)
707 | 11262 | Use HR conditions (6144x4608, N=5)
707 | 11262 | Save to TIFF as Anode_TS_HRQA
707 | 11262 | Inspection Passed and Images Saved
707 | 11165 | Measure aperture size, 40um (+/- 2um)
707 | 11165 | Anode_TS_40um - Passed inspection criteria, Image Saved
707 | 11166 | Ensure size of hole is 55um (+/- 2um
707 | 11166 | Anode_TS_55um - Passed inspection criteria, Image Saved
| 487 | Objective Coil S/N :
| 487 | Scan Coil S/N :
| 487 | Pole Piece S/N :
707 | 11159 | Measure aperture size, 300um (+/- 10um)
707 | 11159 | Inspection Passed and Images Saved
707 | 11262 | Save to TIFF as Anode_TS_HRQA
707 | 11262 | Inspection Passed and Images Saved
707 | 11165 | Measure aperture size, 40um (+/- 2um)
707 | 11165 | Anode_TS_40um - Passed inspection criteria, Image Saved
707 | 11166 | Ensure size of hole is 55um (+/- 2um
707 | 11166 | Anode_TS_55um - Passed inspection criteria, Image Saved
32 | 622 | Actual @ 1kV WD5 =
32 | 622 | Actual @ 20kV WD2 =
678 | 11415 | Confirm test unit fails using the High voltage probe, Hi-Pot Test Set (356180-9217-000) and shorting test box connected (354933-9205-000).
522 | 479
707 | 11159 | Measure aperture size, 300um (+/- 10um)
707 | 11159 | Inspection Passed and Images Saved
707 | 11262 | Anode_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11167 | Ensure size of hole is 90um (+/- 5um)
707 | 11167 | Use LR conditions (1024x768, N=1)
707 | 11167 | Save to TIFF as Anode_TS_90um
707 | 11167 | Inspection Passed and Images Saved
| 490 | Visual Inspection
| 490 | Liner Fitment
| 490 | End Cap Fitment
| 491 | Continuity ok
| 492 | Visually inspected for damage/dents
| 492 | Sealing surfaces free from scratches/ imperfections
| 493 | Position A:
| 493 | Position B:
| 493 | Position C:
| 493 | Position D:
| 493 | Max Difference:
667 | 11411 | Connect Hal Hi-Pot Tester to orange test box and confirm test fails.
522 | 479 | Auto Inductance Test Passed
522 | 479 | Auto Flux Test Passed
707 | 11187 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11187 | Inspection Passed and Images Saved
707 | 11317 | (There is no specific checks not regarding the search for particles)
707 | 11317 | Use HR conditions (6144x4608, N=5)
707 | 11317 | Save to TIFF as Anode_TS_HRQA
707 | 11317 | Inspection Passed and Images Saved
707 | 11167 | Ensure size of hole is 90um (+/- 5um)
707 | 11167 | Use LR conditions (1024x768, N=1)
707 | 11167 | Save to TIFF as Anode_TS_90um
707 | 11167 | Inspection Passed and Images Saved
| 494 | Resin is uniform and not proud
| 494 | No exposed wires
| 495 | Pressure 15 PSI for 5 minutes
| 496 | Stand Off 1:
| 496 | Stand Off 2:
| 496 | Stand Off 3:
| 496 | Stand Off 4:
| 496 | Max Difference:
| 497 | Ensure objective coil is concentric to pole piece ( this should be done as best as possible by eye as no jig yet available)
6 | 710 | Actual
522 | 479 | Auto Inductance Test Passed
522 | 479 | Auto Flux Test Passed
707 | 11187 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11187 | Multi_TS_HRQA Images Saved
707 | 11317 | Anode_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11317 | Anode_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11167 | Ensure size of hole is 55um (+/- 2um
707 | 11167 | Anode_TS_55um - Passed inspection criteria, Image Saved
707 | 11167 | Ensure size of hole is 55um (+/- 2um
707 | 11167 | Anode_TS_90um - Passed inspection criteria, Image Saved
| 499 | Open Circuit (Coil – GND).
707 | 11187 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11187 | Multi_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11318 | (There is no specific checks not regarding the search for particles)
707 | 11318 | Use HR conditions (6144x4608, N=5)
707 | 11318 | Save to TIFF as Anode_TS_HRQA
707 | 11318 | Inspection Passed and Images Saved
707 | 11167 | Ensure size of hole is 90um (+/- 2um
707 | 11167 | Anode_TS_90um - Passed inspection criteria, Image Saved
707 | 11167 | Ensure size of hole is 90um (+/- 5um
707 | 11167 | Anode_TS_90um - Passed inspection criteria, Image Saved
| 500 | Objective Coil
| 500 | Scan Coil
| 501 | Confirm there is no Electical breakdown between Scan coils and Objective body.
667 | 11411 | Confirm test unit fails with shorting test box connected.
674 | 10819 | Ion Pump Cables
674 | 10819 | Column X1 Lens cable
674 | 10819 | Objective Lens Cable
674 | 10819 | Scan Cable
707 | 11187 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11187 | Multi_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11318 | Anode_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11166 | Measure aperture size, 55um (+/- 2um)
707 | 11166 | Anode_TS_55um - Passed inspection criteria, Image Saved
667 | 10765 | HAL Test
667 | 10765 | PAT Test
667 | 10765 | HAL Test
667 | 10765 | PAT Test
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | Confirm the following tests were succesfully run as part of the Auto Routine:................................................ PE Test (<0.1Ω), Insulation Test (520V @ 5sec)
644 | 10498 | Gun Monitor Setup
644 | 10498 | Test No. 96 Selected
644 | 10498 | Run up to 32KV and leave for 1 minute
644 | 10498 | Shutdown and run up to 32KV, Repeat 5 times
644 | 10498 | Run up to 32KV and leave for 30 minutes
644 | 10498 | 32KV Gun Monitor Traced Saved
28 | 700 | Confirm EM Server screeening programme is installed and running
28 | 700 | Confirm no repeated EM Server screening errors have been flagged
707 | 11277 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11277 | Multi_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11284 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11284 | Multi_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11319 | (There is no specific checks not regarding the search for particles)
707 | 11319 | Use HR conditions (6144x4608, N=5)
707 | 11319 | Save to TIFF as Anode_TS_HRQA
707 | 11319 | Inspection Passed and Images Saved
707 | 11167 | Measure aperture size, 90um (+/- 5um)
707 | 11167 | Anode_TS_90um - Passed inspection criteria, Image Saved
707 | 11283 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11283 | Multi_TS_HRQA - Passed inspection criteria, Image Saved
| 532 | Position A:
| 532 | Position B:
| 532 | Position C:
| 532 | Position D:
| 532 | Max. Difference
707 | 11161 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11161 | Use HR conditions (6144x4608, N=5)
707 | 11161 | Save to TIFF as Multi_US_HRQA
707 | 11161 | Inspection Passed and Images Saved
707 | 11319 | Anode_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11167 | Measure aperture size, 90um (+/- 5um)
707 | 11167 | Anode_TS_90um - Passed inspection criteria, Image Saved
| 534 | Has Transzorb been fitted
667 | 10765 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
667 | 10765 | Complete Hi Pot test with safety check rig enabled
667 | 10765 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
667 | 10765 | Complete Hi Pot test with safety check rig enabled
667 | 11411 | Confirm test unit fails with shorting test box connected (354933-9205-000).
707 | 11161 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11161 | Inspection Passed and Images Saved
707 | 11163 | Inspection Passed and Images Saved
707 | 11168 | Ensure size of hole is 150um (+/- 5um)
707 | 11168 | Use LR conditions (1024x768, N=1)
707 | 11168 | Save to TIFF as Anode_TS_150um
707 | 11168 | Inspection Passed and Images Saved
707 | 11168 | Ensure size of hole is 150um (+/- 5um)
707 | 11168 | Use LR conditions (1024x768, N=1)
707 | 11168 | Save to TIFF as Anode_TS_150um
707 | 11168 | Inspection Passed and Images Saved
707 | 11168 | Measure aperture size, 90um (+/- 5um)
707 | 11168 | Anode_TS_90um - Passed inspection criteria, Image Saved
| 535 | Continuity (Close Circuit)
| 535 | Confirm that the coil is NOT short circuit to the condenser body. Do not use test box, check manually.
| 536 | Continuity (Close Circuit)
| 536 | Coils to GND (Open Circuit)
667 | 10765 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
667 | 10765 | Enter HAL SN:
667 | 10765 | Complete Hi Pot test with safety check rig enabled
83 | 701 | Confirm EM Server screeening programme is installed and running
83 | 701 | Confirm no repeated EM Server screening errors have been flagged
707 | 11161 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11161 | Multi_US_HRQA - Passed inspection criteria, Image Saved
707 | 11326 | Inspection Passed and Images Saved
707 | 11168 | Measure aperture size, 150um (+/- 5um)
707 | 11168 | Anode_TS_90um - Passed inspection criteria, Image Saved
707 | 11169 | Anode_US_HRQA - Passed inspection criteria, Image Saved
667 | 10765 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
667 | 10765 | HAL SN:
667 | 10765 | HAL Calibration Date:
667 | 10765 | Complete Hi Pot test with safety check rig enabled
667 | 11412 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | Confirm the following tests were succesfully run as part of the Auto Routine:................................................ PE Test (<0.1Ω), Insulation Test (520V @ 5sec), Residual Current Test (2sec)
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | Confirm the following tests were succesfully run as part of the Auto Routine:................................................ PE Test (<0.1Ω), Insulation Test (520V @ 5sec), Residual Current Test (2sec)
707 | 11289 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11289 | Use HR conditions (6144x4608, N=5)
707 | 11289 | Save to TIFF as Multi_US_HRQA
707 | 11289 | Inspection Passed and Images Saved
707 | 11163 | Inspection Passed and Images Saved
707 | 11169 | (There is no specific checks not regarding the search for particles)
707 | 11169 | Use HR conditions (6144x4608, N=5)
707 | 11169 | Save to TIFF as Anode_US_HRQA
707 | 11169 | Inspection Passed and Images Saved
707 | 11169 | (There is no specific checks not regarding the search for particles)
707 | 11169 | Use HR conditions (6144x4608, N=5)
707 | 11169 | Save to TIFF as Anode_US_HRQA
707 | 11169 | Inspection Passed and Images Saved
707 | 11332 | (There is no specific checks not regarding the search for particles)
707 | 11332 | Use HR conditions (6144x4608, N=5)
707 | 11332 | Save to TIFF as Anode_US_HRQA
707 | 11332 | Inspection Passed and Images Saved
667 | 11412 | Earth Bond Test (TP1) 0.09Ω - 0.07Ω
83 | 701 | Confirm no repeated EM Server screening errors have been flagged
707 | 11289 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11289 | Multi_US_HRQA - Passed inspection criteria, Image Saved
707 | 11326 | Inspection Passed and Images Saved
707 | 11332 | Anode_US_HRQA - Passed inspection criteria, Image Saved
707 | 11333 | Anode_US_HRQA - Passed inspection criteria, Image Saved
667 | 10765 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
667 | 10765 | HAL SN:
667 | 10765 | HAL Calibration Date:
667 | 10765 | Complete Hi Pot test at the following conditions:........................................................................................... Test Voltage: 2000Vdc,Time Up: 2 seconds, Time Hold 5 seconds, Time Down: 2 seconds
667 | 11412 | Earth Bond Test TP1 (0.1Ω - 0.07Ω)
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | R INS - Confirm all test points (>0.500MΩ)
667 | 11413 | Confirm the following tests were succesfully run as part of the Auto Routine:................................................ PE Test (<0.1Ω), Insulation Test (520V @ 5sec), Residual Current Test (2sec)
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | R INS - ( > 0.500MΩ)
667 | 11413 | Confirm EN 61010 Passed
667 | 11413 | Confirm Visual Inspection Criterea Passed
707 | 11290 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11290 | Use HR conditions (6144x4608, N=5)
707 | 11290 | Save to TIFF as Multi_US_HRQA
707 | 11290 | Inspection Passed and Images Saved
707 | 11308 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11308 | Use HR conditions (6144x4608, N=5)
707 | 11308 | Save to TIFF as Multi_US_HRQA
707 | 11308 | Inspection Passed and Images Saved
707 | 11308 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11308 | Inspection Passed and Images Saved
707 | 11327 | (There is no specific checks not regarding the search for particles)
707 | 11327 | Use LR conditions (1024x768, N=1)
707 | 11327 | Save to TIFF as Anode_TS_Bore1
707 | 11327 | Inspection Passed and Images Saved
707 | 11333 | Anode_US_HRQA - Passed inspection criteria, Image Saved
707 | 11333 | Anode_US_HRQA - Passed inspection criteria, Image Saved
707 | 11290 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11290 | Inspection Passed and Images Saved
707 | 11328 | (There is no specific checks not regarding the search for particles)
707 | 11328 | Use LR conditions (1024x768, N=1)
707 | 11328 | Save to TIFF as Anode_TS_Bore1
707 | 11328 | Inspection Passed and Images Saved
707 | 11334 | (There is no specific checks not regarding the search for particles)
707 | 11334 | Use HR conditions (6144x4608, N=5)
707 | 11334 | Save to TIFF as Anode_US_HRQA
707 | 11334 | Inspection Passed and Images Saved
667 | 11412 | Earth Bond Test TP1 (0.1Ω - 0.07Ω)
667 | 11412 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
667 | 11412 | Earth Bond Test TP1 (0.1Ω - 0.07Ω)
667 | 11412 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
667 | 11412 | Insulation Test (2.0MΩ - 1.9MΩ)
667 | 11412 | Earth Bond Test TP1 (0.1Ω - 0.07Ω)
667 | 11412 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
667 | 11412 | Insulation Test (2.0MΩ - 1.9MΩ)
667 | 11412 | Earth Leakage Test (0.9mA - 0.7mA)
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
2 | 743
707 | 11252 | Measure aperture size, 30um (+/- 2um)
707 | 11252 | Inspection Passed and Images Saved
707 | 11285 | Measure aperture size, 30um (+/- 2um)
707 | 11285 | Inspection Passed and Images Saved
707 | 11156 | Measure aperture size, 20um (+/- 2um)
707 | 11156 | Inspection Passed and Images Saved
707 | 11334 | Anode_US_HRQA - Passed inspection criteria, Image Saved
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | Confirm the following tests were succesfully run as part of the Auto Routine:................................................ PE Test, Insulation Test (520V @ 5sec)
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | R INS - ( > 0.500MΩ)
667 | 11413 | Confirm the following tests were succesfully run as part of the Auto Routine:................................................ PE Test (<0.1Ω), Insulation Test (520V @ 5sec), Residual Current Test (2sec)
678 | 11415 | Confirm test unit fails with shorting test box connected (354933-9205-000).
28 | 700 | Confirm EM Server screeening programme is installed and running
28 | 700 | Confirm no repeated EM Server screening errors have been flagged
707 | 11156 | Measure aperture size, 20um (+/- 2um)
707 | 11156 | Inspection Passed and Images Saved
707 | 11170 | (Check surface finish on the inner bore edge is free from marks)
707 | 11170 | Use HR conditions (6144x4608, N=5)
707 | 11170 | Save to TIFF as VP_TS_HRQA
707 | 11170 | Inspection Passed and Images Saved
707 | 11170 | VP_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11171 | (Check surface finish on the inner bore edge is free from marks)
707 | 11171 | Use LR conditions (1024x768, N=1))
707 | 11171 | Save to TIFF as VP_TS_Bore
707 | 11171 | Inspection Passed and Images Saved
707 | 11172 | (Check there is no engraved ring around the aperture, this should be on the underside.
707 | 11172 | Use HR conditions (6144x4608, N=5)
707 | 11172 | Save to TIFF as SF_TS_HRQA
707 | 11172 | Inspection Passed and Images Saved
707 | 11163 | Anode_TS_Bore1 - Passed inspection criteria, Image Saved
707 | 11156 | Measure aperture size, 20um (+/- 2um)
707 | 11156 | Inspection Passed and Images Saved
707 | 11171 | (Check surface finish on the inner bore edge is free from marks)
707 | 11171 | Use LR conditions (1024x768, N=1))
707 | 11171 | Save to TIFF as VP_TS_Bore
707 | 11171 | Inspection Passed and Images Saved
707 | 11171 | VP_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11172 | VP_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11172 | VP_TS_HRQA - Passed inspection criteria, Image Saved
708 | 504 | Select PC
708 | 504 | Hold aperture holder, look for holes on the sides.No particles are to be observed
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | R INS - ( > 0.500MΩ)
667 | 11413 | Confirm the following tests were succesfully run as part of the Auto Routine:................................................ PE Test (<0.1Ω), Insulation Test (520V @ 5sec), Residual Current Test (2sec)
707 | 11286 | Measure aperture size, 30um (+/- 2um)
707 | 11286 | Inspection Passed and Images Saved
707 | 11172 | SF_TS_HRQA - Passed inspection criteria, Image Saved
708 | 505 | Hold aperture holder, look for holes on the sides.No particles are to be observed
708 | 504 | Select PC
708 | 504 | Hold aperture holder, look for holes on the sides.No particles are to be observed
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
667 | 11412 | PAT Checkbox Calibrati0n Date
667 | 11412 | Earth Bond Test TP1 (0.1Ω - 0.07Ω)
667 | 11412 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
667 | 11412 | Insulation Test (2.0MΩ - 1.9MΩ)
667 | 11412 | Earth Leakage Test (0.9mA - 0.7mA)
707 | 11160 | Measure aperture size, 120um (+/- 2um)
707 | 11160 | Inspection Passed and Images Saved
707 | 11173 | Ensure size of hole is 35um (+/- 2 um)
707 | 11173 | Use HR conditions (6144x4608, N=5)
707 | 11173 | Save to TIFF as SF_TS_35um
707 | 11173 | Inspection Passed and Images Saved
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
667 | 11412 | PAT Checkbox Calibrati0n Date
667 | 11412 | Earth Bond Test TP1 (0.1Ω - 0.07Ω)
667 | 11412 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
667 | 11412 | Insulation Test (2.0MΩ - 1.9MΩ)
667 | 11412 | Earth Leakage Test (0.9mA - 0.7mA)
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413
667 | 11413 | R INS - ( > 0.500MΩ)
667 | 11413 | Confirm the following tests were succesfully run as part of the Auto Routine:................................................ PE Test (<0.1Ω), Insulation Test (520V @ 5sec), Residual Current Test (2sec)
707 | 11287 | Measure aperture size, 120um (+/- 2um)
707 | 11287 | Inspection Passed and Images Saved
707 | 11173 | Measure aperture size, 35um (+/- 2um)
707 | 11173 | Anode_TS_40um - Passed inspection criteria, Image Saved
708 | 11238 | Make sure there are no cut-outs on the edge of the anode. No particles are to be observed.
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
708 | 505 | Hold aperture holder, look for holes on the sides.No particles are to be observed
667 | 11412 | PAT Checkbox Calibration Date
667 | 11412 | Earth Bond Test TP1 (0.1Ω - 0.07Ω)
667 | 11412 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
667 | 11412 | Insulation Test (2.0MΩ - 1.9MΩ)
667 | 11412 | Earth Leakage Test (0.9mA - 0.7mA)
707 | 11288 | Measure aperture size, 300um (+/- 10um)
707 | 11288 | Inspection Passed and Images Saved
707 | 11159 | Measure aperture size, 300um (+/- 10um)
707 | 11159 | Inspection Passed and Images Saved
707 | 11252 | Measure aperture size, 30um (+/- 2um)
707 | 11252 | Multi_TS_30um - Passed inspection criteria, Image Saved
707 | 11173 | Measure aperture size, 35um (+/- 2um)
707 | 11173 | Anode_TS_40um - Passed inspection criteria, Image Saved
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
708 | 11238 | Make sure there are no cut-outs on the edge of the anode.
707 | 11285 | Measure aperture size, 30um (+/- 2um)
707 | 11285 | Multi_TS_30um - Passed inspection criteria, Image Saved
707 | 11173 | Measure aperture size, 35um (+/- 2um)
707 | 11173 | Anode_TS_40um - Passed inspection criteria, Image Saved
708 | 11241 | Make sure there are no cut-outs on the edge of the anode. No particles are to be observed.
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
708 | 507 | Anode Inspected & cleaned in Ultra Sonic tank.
50 | 480 | DC Tablefiller successfully completed.
707 | 11156 | Measure aperture size, 20um (+/- 2um)
707 | 11156 | Multi_TS_20um - Passed inspection criteria, Image Saved
707 | 11174 | Ensure size of hole is 750um (+/- 15um)
707 | 11174 | Use HR conditions (6144x4608, N=5)
707 | 11174 | Save to TIFF as Plat_TS_750um_HRQA
707 | 11174 | Inspection Passed and Images Saved
708 | 11236 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed.
708 | 507 | Anode Inspected & cleaned in Ultra Sonic tank.
708 | 507 | Multi-Hole holder Inspected & cleaned in Ultra Sonic tank.
43 | 347 | 1kV Inlens
43 | 347 | Store 1kV Inlens Customer Image
43 | 347 | 15kV Inlens
43 | 347 | Store 15kV Inlens Customer Image
| 510 | Objective Coil S/N :
| 510 | Scan Coil S/N :
| 510 | Pole Piece S/N :
| 517 | Position A:
| 517 | Position B:
| 517 | Position C:
| 517 | Position D:
| 517 | Max Difference:
13 | 168 | Driver A – Liner (>1000MΩ)
13 | 168 | Driver B – Liner (>1000MΩ)
13 | 168 | Driver A – Driver B (>1000MΩ)
707 | 11286 | Measure aperture size, 30um (+/- 2um)
707 | 11286 | Multi_TS_20um - Passed inspection criteria, Image Saved
707 | 11174 | Measure aperture size, 35um (+/- 15um)
707 | 11174 | Anode_TS_40um - Passed inspection criteria, Image Saved
707 | 11174 | Measure aperture size, 750um (+/- 15um)
707 | 11174 | Anode_TS_40um - Passed inspection criteria, Image Saved
708 | 11263 | Make sure there are 3 cut-outs on the edge of the anode. No particles are to be observed
708 | 507 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
| 518 | Position A:
| 518 | Position B:
| 518 | Position C:
| 518 | Position D:
| 518 | Max Difference:
83 | 701 | Confirm no repeated EM Server screening errors have been flagged
2 | 743 | Are there any shortages?
2 | 743 | Shortage PN
707 | 11160 | Measure aperture size, 120um (+/- 2um)
707 | 11160 | Multi_TS_120um - Passed inspection criteria, Image Saved
707 | 11174 | Measure aperture size, 750um (+/- 15um)
707 | 11174 | Plat_TS_750um_HRQA - Passed inspection criteria, Image Saved
708 | 11236 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed.
707 | 506 | Measure aperture size, 300um (+/- 10um)
707 | 506 | Multi_TS_300um - Passed inspection criteria, Image Saved
708 | 507 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
708 | 507 | Anode Inspected & cleaned in Ultra Sonic tank.
708 | 507 | Multi-Hole holder Inspected & cleaned in Ultra Sonic tank.
| 512 | 12 hours drying time
| 513 | Resistance:
| 514 | Visual Inspection
| 514 | Liner Fitment
| 514 | End Cap Fitment
| 519 | Resin is uniform and not proud
| 519 | No exposed wires
45 | 367 | High Res Fil I:
45 | 367 | High Res Extractor V:
707 | 11287 | Measure aperture size, 120um (+/- 2um)
707 | 11287 | Multi_TS_120um - Passed inspection criteria, Image Saved
707 | 11185 | Hold aperture holder, look for holes on the sides.No particles are to be observed
707 | 11186 | Select PC
707 | 11186 | Hold aperture holder, look for holes on the sides.No particles are to be observed
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
707 | 506 | Measure aperture size, 7.5um (+/-2um)
707 | 506 | Multi_TS_7.5um - Passed inspection criteria, Image Saved
708 | 507 | Anode Inspected & Ultra Soniced
708 | 508 | Assemble Multi-Hole Aperture into holder. Inspect using optical microscope.
708 | 508 | Multi-Hole holder Inspected & cleaned in Ultra Sonic tank.
| 515 | Continuity ok
707 | 11288 | Measure aperture size, 300um (+/- 10um)
707 | 11288 | Multi_TS_300um - Passed inspection criteria, Image Saved
707 | 11159 | Measure aperture size, 300um (+/- 10um)
707 | 11159 | Multi_TS_300um - Passed inspection criteria, Image Saved
707 | 11185 | Hold aperture holder, look for holes on the sides.No particles are to be observed
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
707 | 506 | Measure aperture size, 7.5um (+/-2um)
707 | 506 | Multi_TS_7.5um - Passed inspection criteria, Image Saved
707 | 506 | Measure aperture size, 10um (+/-2um)
707 | 506 | Multi_TS_10um - Passed inspection criteria, Image Saved
708 | 508 | Assemble Multi-Hole Aperture into holder. Inspect using optical microscope.
708 | 508 | Assemble Anode Aperture into Anode.
| 520 | Pressure 15 PSI for 5 minutes
45 | 367 | High Res Fil I:
45 | 367 | High Res Extractor V:
707 | 11163 | Inspection Passed and Images Saved
708 | 11264 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be observed.
707 | 506 | Measure aperture size, 7.5um (+/-2um)
707 | 506 | Multi_TS_7.5um - Passed inspection criteria, Image Saved
707 | 506 | Measure aperture size, 10um (+/-2um)
707 | 506 | Multi_TS_10um - Passed inspection criteria, Image Saved
708 | 508 | Assemble Multi-Hole Aperture into holder. Inspect using optical microscope.
708 | 508 | Flame Anode aperture and assemble into Anode.
708 | 508 | Assemble Multi-Hole Aperture into holder. Inspect using optical microscope.
708 | 508 | Flame Anode aperture and assemble into Anode.
| 516 | Visually inspected for damage/dents
| 516 | Sealing surfaces free from scratches/ imperfections
| 521 | Stand Off 1:
| 521 | Stand Off 2:
| 521 | Stand Off 3:
| 521 | Stand Off 4:
| 521 | Max Difference:
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | R INS - ( > 0.500MΩ)
667 | 11413 | Confirm the following tests were succesfully run as part of the Auto Routine:................................................ PE Test (<0.1Ω), Insulation Test (520V @ 5sec), Residual Current Test (2sec)
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | R INS - ( > 0.500MΩ)
667 | 11413 | I HLNC - ( < 0.500mA)
667 | 11413 | Confirm EN 61010 Passed
667 | 11413 | Confirm Visual Inspection Criterea Passed
45 | 367 | Nominal Fil I:
45 | 367 | Nominal extractor V:
707 | 11163 | Anode_TS_Bore1 - Passed inspection criteria, Image Saved
708 | 11264 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be observed.
707 | 506 | Measure aperture size, 7.5um (+/-2um)
707 | 506 | Multi_TS_7.5um - Passed inspection criteria, Image Saved
707 | 506 | Measure aperture size, 10um (+/-2um)
707 | 506 | Multi_TS_10um - Passed inspection criteria, Image Saved
707 | 506 | Measure aperture size, 15um (+/-2um)
707 | 506 | Multi_TS_15um - Passed inspection criteria, Image Saved
25 | 11401 | 30um aperture entered on central aperture
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | R INS - ( > 0.500MΩ)
667 | 11413 | Confirm EN 61010 Passed
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | R INS - ( > 0.500MΩ)
667 | 11413 | I HLNC - ( < 0.500mA)
667 | 11413 | I HLSF - ( < 3.500mA)
667 | 11413 | Confirm EN 61010 Passed
667 | 11413 | Confirm Visual Inspection Criterea Passed
707 | 11290 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11290 | Multi_US_HRQA - Passed inspection criteria, Image Saved
9 | 1652 | S/N =
707 | 11326 | Inspection Passed and Images Saved
708 | 11265 | Make sure there are 3 cut-outs on the edge of the anode. No particles are to be observed.
707 | 506 | Measure aperture size, 7.5um (+/-2um)
707 | 506 | Multi_TS_7.5um - Passed inspection criteria, Image Saved
707 | 506 | Measure aperture size, 10um (+/-2um)
707 | 506 | Multi_TS_10um - Passed inspection criteria, Image Saved
707 | 506 | Measure aperture size, 15um (+/-2um)
707 | 506 | Multi_TS_15um - Passed inspection criteria, Image Saved
707 | 506 | Measure aperture size, 20um (+/-2um)
707 | 506 | Multi_TS_20um - Passed inspection criteria, Image Saved
707 | 506 | Measure aperture size, 7.5um (+/-2um)
707 | 506 | Multi_TS_7.5um - Passed inspection criteria, Image Saved
707 | 506 | Measure aperture size, 10um (+/-2um)
707 | 506 | Multi_TS_10um - Passed inspection criteria, Image Saved
707 | 506 | Measure aperture size, 15um (+/-2um)
707 | 506 | Multi_TS_15um - Passed inspection criteria, Image Saved
707 | 506 | Measure aperture size, 20um (+/-2um)
707 | 506 | Multi_TS_20um - Passed inspection criteria, Image Saved
707 | 506 | Measure aperture size, 60um (+/-2um)
707 | 506 | Multi_TS_60um - Passed inspection criteria, Image Saved
667 | 10765 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
667 | 10765 | HAL SN:
667 | 10765 | HAL Calibration Date:
667 | 10765 | DCIR (Direct Current Insulation Resistance @500V) (>10.0M)
667 | 10765 | Complete Hi Pot test at the following conditions:........................................................................................... Test Voltage: 2000Vdc,Time Up: 2 seconds, Time Hold 5 seconds, Time Down: 2 seconds
667 | 10765 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
667 | 10765 | HAL SN:
667 | 10765 | HAL Calibration Date:
667 | 10765 | DC IR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
667 | 10765 | DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
667 | 10765 | Confirm System Passed Hi-Pot Tests
707 | 11326 | Anode_TS_Bore1 - Passed inspection criteria, Image Saved
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
17 | 711 | Objective Coil S/N :
17 | 711 | Scan Coil S/N :
17 | 711 | Pole Piece S/N :
667 | 10765 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
667 | 10765 | HAL SN:
667 | 10765 | HAL Calibration Date:
667 | 10765 | DCIR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
667 | 10765 | Complete Hi Pot test at the following conditions:........................................................................................... Test Voltage: 2000Vdc,Time Up: 2 seconds, Time Hold 5 seconds, Time Down: 2 seconds
667 | 10765 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
667 | 10765 | HAL SN:
667 | 10765 | HAL Calibration Date:
667 | 10765 | DCIR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
667 | 10765 | Complete Hi Pot test at the following conditions:........................................................................................... Test Voltage: 2000Vdc,Time Up: 2 seconds, Time Hold 5 seconds, Time Down: 2 seconds
| 580 | S/N =
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | R INS - ( > 500µΩ)
667 | 11413 | I HLNC - ( < 500µA)
667 | 11413 | I HLSF - ( < 3.500mA)
667 | 11413 | Confirm EN 61010 Passed
667 | 11413 | Confirm Visual Inspection Criterea Passed
45 | 367 | High Res Fil I:
45 | 367 | High Res Extractor V:
707 | 11327 | (There is no specific checks not regarding the search for particles)
707 | 11327 | Use LR conditions (1024x768, N=1)
707 | 11327 | Save to TIFF as Anode_TS_Bore1
707 | 11327 | Inspection Passed and Images Saved
667 | 10765 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
667 | 10765 | HAL SN:
667 | 10765 | HAL Calibration Date:
667 | 10765 | DCIR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
667 | 10765 | Confirm Hi-Pot Test Result
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | R INS - ( > 0.5MΩ)
667 | 11413 | I HLNC - ( < 500µA)
667 | 11413 | I HLSF - ( < 3.500mA)
667 | 11413 | Confirm EN 61010 Passed
667 | 11413 | Confirm Visual Inspection Criterea Passed
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | R INS - ( > 0.5MΩ)
667 | 11413 | I HLNC - ( < 500µA)
667 | 11413 | I HLSF - ( < 3500µA)
667 | 11413 | Confirm EN 61010 Passed
667 | 11413 | Confirm Visual Inspection Criterea Passed
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | R INS - ( > 0.5MΩ)
667 | 11413 | I HLNC - ( < 500µA)
667 | 11413 | I HLSF - ( < 3500µA)
667 | 11413 | Confirm EN 61010 Passed
667 | 11413 | Confirm Visual Inspection Criterea Passed
13 | 171 | Driver A – Liner (>1000MΩ)
13 | 171 | Driver B – Liner (>1000MΩ)
13 | 171 | Driver A – Driver B (>1000MΩ)
707 | 11327 | Anode_TS_Bore1 - Passed inspection criteria, Image Saved
17 | 713 | Visual Inspection
17 | 713 | Liner Fitment
17 | 713 | End Cap Fitment
667 | 10765 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
667 | 10765 | HAL SN:
667 | 10765 | HAL Calibration Date:
667 | 10765 | DCIR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
667 | 10765 | DCIR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
667 | 10765 | Confirm Hi-Pot Test Result
| 510 | Objective Coil S/N :
| 510 | Scan Coil S/N :
| 510 | Pole Piece S/N :
| 512 | 12 hours drying time
| 513 | Resistance:
6 | 625 | HAL SN:
6 | 625 | HAL Calibration Date:
6 | 625 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | HAL SN:
6 | 1546 | HAL Calibration Date:
6 | 1546 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
45 | 368 | Analytic Fil I
45 | 368 | Analytic Extractor V:
707 | 11328 | (There is no specific checks not regarding the search for particles)
707 | 11328 | Use LR conditions (1024x768, N=1)
707 | 11328 | Save to TIFF as Anode_TS_Bore1
707 | 11328 | Inspection Passed and Images Saved
707 | 11328 | Anode_TS_Bore1 - Passed inspection criteria, Image Saved
| 527 | S/N =
50 | 480 | DC Tablefiller successfully completed.
| 541 | Drying time of 12hrs.
| 557 | A - B (< 100kΩ)
| 557 | C - D (< 10Ω)
| 557 | B - D (Open Circuit)
| 560 | S/N =
707 | 11164 | (There is no specific checks not regarding the search for particles)
707 | 11164 | Use LR conditions (1024x768, N=1)
707 | 11164 | Save to TIFF as Anode_TS_Bore2
707 | 11164 | Inspection Passed and Images Saved
| 527 | S/N =
| 530 | S/N =
| 537 | S/N =
50 | 480 | DC Tablefiller successfully completed.
| 542 | C1 Coil
| 542 | C2 Coil
| 543 | C1 Coil
| 543 | C2 Coil
| 563 | Housing Type
| 563 | Inspect insulation disc using white light for cracks
| 574 | Housing Type
| 574 | Inspect insulation disc using white light for cracks
| 582 | Objective Coil S/N :
| 582 | Scan Coil S/N :
| 582 | Pole Piece S/N :
25 | 11401 | 30um aperture entered on central aperture
707 | 11164 | (There is no specific checks not regarding the search for particles)
707 | 11164 | Use LR conditions (1024x768, N=1)
707 | 11164 | Save to TIFF as Anode_TS_Bore2
707 | 11164 | Inspection Passed and Images Saved
| 537 | S/N =
50 | 480 | DC Tablefiller successfully completed.
| 559 | S/N =
| 564 | Cooling Isolation Check (Open Circuit Coil – GND)
| 586 | Continuity ok
667 | 10765 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
667 | 10765 | HAL SN:
667 | 10765 | HAL Calibration Date:
667 | 10765 | DC IR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
667 | 10765 | DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
667 | 10765 | Confirm Hi-Pot Test Result
| 514 | Visual Inspection
| 514 | Liner Fitment
| 514 | End Cap Fitment
6 | 1547 | PAT Checkbox Calibration Date
6 | 1547 | Earth Bond Test TP1 (0.1Ω - 0.07Ω)
6 | 1547 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
6 | 1547 | Insulation Test (2.0MΩ - 1.9MΩ)
6 | 1547 | Earth Leakage Test (0.9mA - 0.7mA)
707 | 11329 | (There is no specific checks not regarding the search for particles)
707 | 11329 | Use LR conditions (1024x768, N=1)
707 | 11329 | Save to TIFF as Anode_TS_Bore2
707 | 11329 | Inspection Passed and Images Saved
707 | 11329 | (There is no specific checks not regarding the search for particles)
707 | 11329 | Use LR conditions (1024x768, N=1)
707 | 11329 | Save to TIFF as Anode_TS_Bore2
707 | 11329 | Inspection Passed and Images Saved
| 532 | Position A:
| 532 | Position B:
| 532 | Position C:
| 532 | Position D:
| 532 | Max. Difference
| 537 | S/N =
| 576 | Operational repeatability test x 3
| 587 | Visually inspected for damage/dents
| 587 | Sealing surfaces free from scratches/ imperfections
667 | 10765 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
667 | 10765 | HAL SN:
667 | 10765 | HAL Calibration Date:
667 | 10765 | DC IR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
667 | 10765 | DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
667 | 10765 | Confirm Hi-Pot Test Result
| 515 | Continuity ok
707 | 11330 | (There is no specific checks not regarding the search for particles)
707 | 11330 | Use LR conditions (1024x768, N=1)
707 | 11330 | Save to TIFF as Anode_TS_Bore2
707 | 11330 | Inspection Passed and Images Saved
707 | 11331 | (There is no specific checks not regarding the search for particles)
707 | 11331 | Use LR conditions (1024x768, N=1)
707 | 11331 | Save to TIFF as Anode_TS_Bore2
707 | 11331 | Inspection Passed and Images Saved
| 546 | Inspect
| 546 | Clean
| 560 | S/N =
| 585 | Visual Inspection
| 585 | Liner Fitment
| 585 | End Cap Fitment
| 588 | Position A:
| 588 | Position B:
| 588 | Position C:
| 588 | Position D:
| 588 | Max Difference:
| 516 | Visually inspected for damage/dents
| 516 | Sealing surfaces free from scratches/ imperfections
707 | 11164 | Anode_TS_Bore2 - Passed inspection criteria, Image Saved
707 | 11329 | Anode_TS_Bore2 - Passed inspection criteria, Image Saved
707 | 11330 | Anode_TS_Bore2 - Passed inspection criteria, Image Saved
707 | 11331 | Anode_TS_Bore2 - Passed inspection criteria, Image Saved
| 534 | Has Transzorb been fitted
| 578 | Pneumatic Operation
| 578 | Pin Touching when Retracted
| 589 | Resin is uniform and not proud
| 589 | No exposed wires
| 591 | Stand Off 1:
| 591 | Stand Off 2:
| 591 | Stand Off 3:
| 591 | Stand Off 4:
| 591 | Max Difference:
52 | 484 | Aperture Finder run succesfully
| 535 | Continuity (Close Circuit)
| 535 | Confirm that the coil is NOT short circuit to the condenser body. Do not use test box, check manually.
| 549 | Continuity (Close Circuit)
| 549 | Coils to GND (Open Circuit)
| 590 | Pressure 15 PSI for 5 minutes
| 592 | Ensure objective coil is concentric to pole piece ( this should be done as best as possible by eye as no jig yet available)
667 | 10765 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
667 | 10765 | HAL SN:
667 | 10765 | HAL Calibration Date:
667 | 10765 | DC IR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
667 | 10765 | DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
667 | 10765 | Confirm Hi-Pot Test Result
| 517 | Position A:
| 517 | Position B:
| 517 | Position C:
| 517 | Position D:
| 517 | Max Difference:
6 | 1546 | HAL SN:
6 | 1546 | HAL Calibration Date:
6 | 1546 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 628 | Secutest SN:
6 | 628 | Secutest Calibration Date:
6 | 628 | PE Test - Confirm all test points (<0.1Ω)
6 | 628 | R INS - ( > 0.5MΩ)
6 | 628 | I HLNC - ( < 500µA)
6 | 628 | I HLSF - ( < 3500µA)
6 | 628 | Confirm EN 61010 Passed
6 | 628 | Confirm Visual Inspection Criterea Passed
13 | 172 | Driver A – Driver A In
13 | 172 | Driver B – Driver B In
13 | 172 | Anode – Liner In
| 536 | Continuity (Close Circuit)
| 536 | Coils to GND (Open Circuit)
| 550 | Continuity (Close Circuit)
| 550 | Coils to GND (Open Circuit)
| 567 | Operational repeatability test x 3
| 580 | S/N =
| 594 | Open Circuit (Coil – GND).
| 518 | Position A:
| 518 | Position B:
| 518 | Position C:
| 518 | Position D:
| 518 | Max Difference:
| 519 | Resin is uniform and not proud
| 519 | No exposed wires
| 520 | Pressure 15 PSI for 5 minutes
| 521 | Stand Off 1:
| 521 | Stand Off 2:
| 521 | Stand Off 3:
| 521 | Stand Off 4:
| 521 | Max Difference:
| 522 | Ensure objective coil is concentric to pole piece ( this should be done as best as possible by eye as no jig yet available)
| 524 | Open Circuit (Coil – GND).
708 | 508 | Assemble Final Aperture (Multi / Single). Inspect using optical microscope.
708 | 508 | Flame Anode aperture and assemble into Anode.
| 551 | Continuity (Close Circuit)
| 551 | Coils to GND (Open Circuit)
| 580 | S/N =
| 595 | Objective Coil
| 595 | Scan Coil
| 600 | Objective Coil S/N :
| 600 | Scan Coil S/N :
| 600 | Pole Piece S/N :
| 607 | Pressure 15 PSI for 5 minutes
| 608 | Stand Off 1:
| 608 | Stand Off 2:
| 608 | Stand Off 3:
| 608 | Stand Off 4:
| 608 | Max Difference:
707 | 11285 | Measure aperture size, 30um (+/- 2um)
707 | 11285 | Multi_TS_30um - Passed inspection criteria, Image Saved
| 552 | Continuity (Close Circuit)
| 552 | Coils to GND (Open Circuit)
| 569 | Pneumatic Operation
| 569 | Pin Touching when Retracted
| 596 | Confirm there is no Electical breakdown between Scan coils and Objective body.
| 609 | Ensure objective coil is concentric to pole piece ( this should be done as best as possible by eye as no jig yet available)
707 | 11160 | Measure aperture size, 120um (+/- 2um)
707 | 11160 | Multi_TS_120um - Passed inspection criteria, Image Saved
707 | 11170 | VP_TS_HRQA - Passed inspection criteria, Image Saved
| 525 | Confirm there is no Electical breakdown between Scan coils and Objective body.
6 | 625 | HAL SN:
6 | 625 | HAL Calibration Date:
6 | 625 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
708 | 507 | Anode Inspected & cleaned in Ultra Sonic tank.
708 | 507 | Final Aperture Holder (Multi / Single) Inspected & cleaned in Ultra Sonic tank.
6 | 710 | Actual
6 | 710 | Portable IGP controller fitted?
| 597 | S/N =
| 603 | Visual Inspection
| 603 | Liner Fitment
| 603 | End Cap Fitment
707 | 11170 | VP_TS_HRQA - Passed inspection criteria, Image Saved
| 526 | Objective Coil
| 526 | Scan Coil
6 | 1546 | HAL SN:
6 | 1546 | HAL Calibration Date:
6 | 1546 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
646 | 10529 | BSD Diode Pre-test
646 | 10529 | Select Type
646 | 10529 | Select Amplifier Type
646 | 10529 | Black Level Calibration
646 | 10529 | Check Diode Segment Function
646 | 10529 | Check All Gain ranges
646 | 10529 | Check Auto range
646 | 10529 | Check Auto Level
646 | 10529 | Setup and Check TOPO
646 | 10529 | Take 3 images - Atomic, 5pA, 2KV
646 | 10529 | Store BSD Functional Images-Atomic, 5pA, 2KV
678 | 11417 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
678 | 11417 | HAL SN:
678 | 11417 | HAL Calibration Date:
678 | 11417 | DC IR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
678 | 11417 | DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 11417 | Confirm System Passed Hi-Pot Tests
2 | 743 | Are there any shortages?
| 571 | S/N =
| 597 | S/N =
| 602 | Continuity ok
| 604 | Visually inspected for damage/dents
| 604 | Sealing surfaces free from scratches/ imperfections
| 611 | Open Circuit (Coil – GND).
707 | 11170 | VP_TS_HRQA - Passed inspection criteria, Image Saved
6 | 625 | HAL SN:
6 | 625 | HAL Calibration Date:
6 | 625 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
646 | 10529 | BSD Diode Pre-test
646 | 10529 | Select Type
646 | 10529 | Select Amplifier Type
646 | 10529 | Black Level Calibration
646 | 10529 | Check Diode Segment Function
646 | 10529 | Check All Gain ranges
646 | 10529 | Check Auto range
646 | 10529 | Check Auto Level
646 | 10529 | Setup and Check TOPO
646 | 10529 | Take 3 images - Atomic, 5pA, 2KV
646 | 10529 | Store BSD Functional Images-Atomic, 5pA, 2KV
21 | 257 | In Factory Tab: Omit Dongle Check
21 | 257 | In Factory Tab: Enter S/N
21 | 257 | In License Tab: Set All Licenses
| 571 | S/N =
| 605 | Position A:
| 605 | Position B:
| 605 | Position C:
| 605 | Position D:
| 605 | Max Difference:
707 | 11171 | VP_TS_HRQA - Passed inspection criteria, Image Saved
| 527 | S/N =
13 | 167 | Dowell pins have been mounted correctly to the mounting plate
6 | 1545 | Complete HiPot Result Sheet
6 | 1546 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
6 | 1546 | HAL SN:
6 | 1546 | HAL Calibration Date:
6 | 1546 | DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | Confirm Hi-Pot Test Result
646 | 10531 | Select Type
646 | 10531 | Black Level Cal
646 | 10531 | C2D Bias Target Check
646 | 10531 | Gain Range Check
646 | 10531 | Auto Level Check
646 | 10531 | Auto Range Check
646 | 10531 | Function Image
646 | 10531 | Store C2D Functional Image
37 | 461 | 50v - All Aperture at Normal and High current WD3
37 | 461 | 21k - All Apertures at Normal and High current WD5
37 | 461 | 30kv - Aperture 1 and 7 at Normal and High current at WD5
3 | 745 | Are there any shortages?
| 606 | Resin is uniform and not proud
| 606 | No exposed wires
| 612 | Objective Coil
| 612 | Scan Coil
30 | 303 | Actual
30 | 303 | Store stage drift macro results image
707 | 11174 | Measure aperture size, 750um (+/- 15um)
707 | 11174 | Plat_TS_750um_HRQA - Passed inspection criteria, Image Saved
20 | 11038 | PC S/N
20 | 11038 | Hard disk Image (Version)
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
6 | 1545 | Confirm test unit fails using the high voltgae probe and shorting test box connected (354933-9205-000).
21 | 187 | Z Move on Vent Enabled
| 613 | Confirm there is no Electical breakdown between Scan coils and Objective body.
30 | 303 | Actual
30 | 303 | Store stage drift macro results image
20 | 11038 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
20 | 11038 | PC S/N
20 | 11038 | Hard disk Image (Version)
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
20 | 11038 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
20 | 11038 | PC S/N
20 | 11038 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, then the PC needs reimaging – follow the attached WI. The correct image version is V1.
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Hard disk Image (Version)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
6 | 1546 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
6 | 1546 | HAL SN:
6 | 1546 | HAL Calibration Date:
6 | 1546 | DC IR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
6 | 1546 | DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | Confirm Hi-Pot Test Result
6 | 1545 | Confirm test unit fails using the High voltage probe, Hi-Pot Test Set (356180-9217-000) and shorting test box connected (354933-9205-000).
| 595 | Objective Coil
| 595 | Scan Coil
| 614 | S/N =
20 | 11038 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
20 | 11038 | PC S/N
20 | 11038 | If SEM Enhanced PC(349535-9001-020) has Serial Number Before 688330, then the PC needs reimaging – follow the attached WI. The correct image version is V1.
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Hard disk Image (Version)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
6 | 1546 | HAL SN:
6 | 1546 | HAL Calibration Date:
6 | 1546 | DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | Confirm Hi-Pot Test Result
646 | 10531 | Select Type
646 | 10531 | Black Level Cal
646 | 10531 | C2D Bias Target Check
646 | 10531 | Gain Range Check
646 | 10531 | Auto Level Check
646 | 10531 | Auto Range Check
646 | 10531 | Function Image
646 | 10531 | Store C2D Functional Image
708 | 11236 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed.
703 | 637 | Run 3 hour plasma clean
646 | 10543 | Ferrite Fitted To Cables - PC End
646 | 10543 | $NavCam Postion Setup
646 | 10543 | Image Navigation Setup
646 | 10543 | Functional Check
20 | 11038 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
20 | 11038 | PC S/N
20 | 11038 | If SEM Enhanced PC(349535-9001-020) has Serial Number Before 688330, then the PC needs reimaging – follow the attached WI. The correct image version is V1.
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Hard disk Image (Version)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
6 | 625 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
6 | 625 | HAL SN:
6 | 625 | HAL Calibration Date:
6 | 625 | DC IR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
6 | 625 | DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | Confirm Hi-Pot Test Result
6 | 1546 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
6 | 1546 | HAL SN:
6 | 1546 | HAL Calibration Date:
6 | 1546 | DC IR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
6 | 1546 | DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | Confirm Hi-Pot Test Result
6 | 625 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
6 | 625 | HAL SN:
6 | 625 | HAL Calibration Date:
6 | 625 | DC IR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
6 | 625 | DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | Confirm Hi-Pot Test Result
6 | 1546 | HAL SN:
6 | 1546 | HAL Calibration Date:
6 | 1546 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | Confirm Hi-Pot Test Result
4 | 629 | Ensure Pirani gauge is isolated from the roughing line
6 | 1546 | HAL SN:
6 | 1546 | HAL Calibration Date:
6 | 1546 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | Cathode Head Heater (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
21 | 12567 | Enter the FIB Extension number using EEPromUpdater.exe
21 | 12567 | Record FIB Extension Serial No.
708 | 11236 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed.
703 | 637 | Run 3 hour plasma clean
9 | 11242 | O-Rings TEM oiled
4 | 11384 | No Leaks Detected (If using Agilent leaker tester GROSS TEST MODE must be off)
4 | 11384 | On completion before venting, Turn of the Turbo using the omni controller top right hand corner icon
4 | 11384 | Isolate the Leak detector from the roughing line
4 | 11384 | Isolate the system from the roughing line
4 | 11384 | Connect the air filter up to the roughing line
4 | 11384 | Wait until the Turbo freq has dropped to below 500, as displayed on the omni controller. Then open up the system to the vent line
674 | 416 | No Leaks Detected (If using Agilent leaker tester GROSS TEST MODE must be off)
674 | 416 | On completion before venting, Turn of the Turbo using the omni controller top right hand corner icon
674 | 416 | Isolate the Leak detector from the roughing line
674 | 416 | Isolate the system from the roughing line
674 | 416 | Connect the air filter up to the roughing line
674 | 416 | Wait until the Turbo freq has dropped to below 500, as displayed on the omni controller. Then open up the system to the vent line
4 | 629 | Ensure Pirani gauge is isolated from the roughing line
6 | 625 | HAL SN:
6 | 625 | HAL Calibration Date:
6 | 625 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | Cathode Head Heater (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
40 | 11084 | VP Mode Macro selects C2D when going to VP
40 | 11084 | Black Level calibration
40 | 11084 | Functional Checks ok
40 | 11084 | Take Functional Image
717 | 638 | Run 3 hour plasma clean
18 | 639 | Objective Coil S/N :
18 | 639 | Scan Coil S/N :
18 | 639 | Pole Piece S/N :
9 | 103 | Log Aperture number as shown on pot on Aperture database (see link attached in 'Documents')
31 | 306 | Actual Gun Axis X =
31 | 306 | Actual Gun Axis Y =
31 | 307 | Actual Align 1 X =
31 | 307 | Actual Align 1 Y =
663 | 414 | No Leaks Detected (If using Agilent leaker tester GROSS TEST MODE must be off)
663 | 414 | On completion before venting, Turn of the Turbo using the omni controller top right hand corner icon
663 | 414 | Isolate the Leak detector from the roughing line
663 | 414 | Isolate the system from the roughing line
663 | 414 | Connect the air filter up to the roughing line
663 | 414 | Wait until the Turbo freq has dropped to below 400, as displayed on the omni controller. Then open up the system to the vent line
22 | 11031 | Set joystick angle to 180 degrees.
22 | 421 | Select "Tools and User Preferences": Set "Expert Gun Mode = Yes"
| 597 | Enter pole shoe S/N =
646 | 10531 | Select Type
646 | 10531 | Black Level Cal
646 | 10531 | C2D Bias Target Check
646 | 10531 | Gain Range Check
646 | 10531 | Auto Level Check
646 | 10531 | Auto Range Check
646 | 10531 | Function Image
646 | 10531 | Store C2D Functional Image
708 | 11264 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be observed.
18 | 641 | Visual Inspection
18 | 641 | Liner Fitment
18 | 641 | End Cap Fitment
31 | 306 | Actual Gun Axis X =
31 | 306 | Actual Gun Axis Y =
31 | 308 | Actual 1kV X Axis =
31 | 308 | Actual 1kV Y Axis =
31 | 308 | Actual 20kV X Axis =
31 | 308 | Actual 20kV Y Axis =
4 | 629 | Ensure Pirani gauge is isolated from the roughing line
650 | 10567 | Install Customer Licences
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer.db to the Product Folder
650 | 10567 | Store EM Server
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
650 | 10567 | Add Licence to the SmartSEM USB Stick
| 600 | Objective Coil S/N :
| 600 | Scan Coil S/N :
| 600 | Pole Piece S/N :
708 | 11264 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be observed.
18 | 642 | Continuity ok
18 | 643 | Visually inspected for damage/dents
18 | 643 | Sealing surfaces free from scratches/ imperfections
4 | 668 | Fit chamber to damper
4 | 668 | Fit chamber bolts item 10 and 12 and Torque to 42Nm.
4 | 668 | Fit chamber bolts item 13 and 14 and Torque to 21Nm.
4 | 668 | Fit chamber pumping tube and Toruqe bolts item 7 to 12Nm.
4 | 11389 | Fit chamber pumping pipe
4 | 11389 | Fit chamber to damper
4 | 11389 | Fit chamber bolts
4 | 11389 | Fit earthquake bolts
675 | 10824 | All Cables Are Neatly Routed
675 | 10824 | All Plugs Are Secured Correctly
675 | 10824 | All Earth Cables Are Tightly Secured
675 | 10824 | Verify the EHT gun assembly earth is torqued p to 15Nm.
708 | 11264 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be observed.
18 | 644 | Position A:
18 | 644 | Position B:
18 | 644 | Position C:
18 | 644 | Position D:
18 | 644 | Max Difference:
4 | 668 | Fit chamber pumping pipe
4 | 668 | Fit chamber to damper
4 | 668 | Fit chamber bolts
4 | 668 | Fit earthquake bolts
650 | 10567 | Install Customer Licences
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer.db to the Product Folder
650 | 10567 | Store EM Server.db
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
650 | 10567 | Add Licence to the SmartSEM USB Stick
43 | 671 | 1kV =
43 | 671 | Store 1KV Resolution Images and Pate image
43 | 671 | 15kV =
43 | 671 | Store 1KV Resolution Images and Pate image
18 | 645 | Resin is uniform and not proud
18 | 645 | No exposed wires
18 | 648 | Ensure objective coil is concentric to pole piece ( this should be done as best as possible by eye as no jig yet available)
4 | 668 | Fit chamber to damper
4 | 668 | Fit chamber bolts item 10 and 12 and Torque to 42Nm.
4 | 668 | Fit chamber bolts item 13 and 14 and Torque to 21Nm.
4 | 668 | Fit chamber pumping tube and Toruqe bolts item 7 to 12Nm.
650 | 10567 | Install Customer Licences
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer.db to the Product Folder
650 | 10567 | Store EM Server.db
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
650 | 10567 | Add Licence to the SmartSEM USB Stick
650 | 10567 | Add EM Server.db to the SmartSEM USB Stick
651 | 669 | The Licence file has been copied to the SmartSEM USB Stick
651 | 669 | The EM Server.db has been copied to the SmartSEM USB Stick
6 | 710 | Actual
6 | 710 | Portable IGP controller fitted? (See work instruction for details)
44 | 923 | Confirm that all deviations have been cleared on OST - The Andon lamp should be Green
44 | 923 | Remove all old / invalid data and images from the PC.
44 | 923 | Ensure all image folders are labelled accordingly.
44 | 923 | Ensure all 'return to stock' or 'reject parts have been processed.
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
4 | 11384 | No Leaks Detected (If using Agilent leaker tester GROSS TEST MODE must be off)
4 | 11384 | On completion before venting, Turn of the Turbo using the omni controller top right hand corner icon
4 | 11384 | Isolate the Leak detector from the roughing line
4 | 11384 | Isolate the system from the roughing line
4 | 11384 | Connect the air filter up to the roughing line
4 | 11384 | Wait until the Turbo freq has dropped from 1000 to below 900, as displayed on the omni controller. Then open up the system to the vent line
4 | 11384 | No Leaks Detected (If using Agilent leaker tester GROSS TEST MODE must be off)
4 | 11384 | On completion before venting, Turn of the Turbo using the omni controller top right hand corner icon
4 | 11384 | Isolate the Leak detector from the roughing line
4 | 11384 | Isolate the system from the roughing line
4 | 11384 | Connect the air filter up to the roughing line
4 | 11384 | Wait until the Turbo freq has dropped to below 900, as displayed on the omni controller. Then open up the system to the vent line
674 | 416 | No Leaks Detected (If using Agilent leaker tester GROSS TEST MODE must be off)
674 | 416 | On completion before venting, Turn of the Turbo using the omni controller top right hand corner icon
674 | 416 | Isolate the Leak detector from the roughing line
674 | 416 | Isolate the system from the roughing line
674 | 416 | Connect the air filter up to the roughing line
674 | 416 | Wait until the Turbo freq has dropped to below 900, as displayed on the omni controller. Then open up the system to the vent line
642 | 10482 | Baratron Gauge Set To Zero
642 | 10482
642 | 10482 | 2nd Rotary Pump Is Not Running
642 | 10482
642 | 10482 | Baratron Gauge Set To Zero
642 | 10482 | 2nd Rotary Pump Is NOT Running
644 | 10510 | 500um Lower Aperture Fitted
644 | 10510 | 2nd Rotary Pump is ON in EP Mode
644 | 10510 | 10Pa Chamber Vacuum Achieved
644 | 10510 | 3000Pa Vacuum Achieved (Air)
644 | 10510 | Water Function Checked
644 | 10510 | Approx. 2000Pa Chamber Vacuum Achieved with Water
642 | 10483 | UnPlug PL4 on Vac PCB
642 | 10483 | Check That 2nd Rotary Pump is OFF in HV
642 | 10483 | Go To VP, 2nd Rotary Pump is ON
642 | 10483 | Set Up Driver Box
642 | 10483 | Go TO HV, 2nd Rotary Pump Switched OFF
642 | 10483 | ReFit PL4 on Vac PCB
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
663 | 414 | No Leaks Detected (If using Agilent leaker tester GROSS TEST MODE must be off)
663 | 414 | On completion before venting, Turn of the Turbo using the omni controller top right hand corner icon
663 | 414 | Isolate the Leak detector from the roughing line
663 | 414 | Isolate the system from the roughing line
663 | 414 | Connect the air filter up to the roughing line
663 | 414 | Wait until the Turbo freq has dropped to below 720, as displayed on the omni controller. Then open up the system to the vent line
644 | 10510 | 500um Lower Aperture Fitted
644 | 10510 | 2nd Rotary Pump is ON in EP Mode
644 | 10510 | 10Pa Chamber Vacuum Achieved
644 | 10510 | 3000Pa Vacuum Achieved (Air)
644 | 10510 | Water Function Checked
644 | 10510 | Approx. 2000Pa Chamber Vacuum Achieved with Water
644 | 10510 | Go To HV, 2nd Rotary Pump is OFF
18 | 646 | Pressure 15 PSI for 5 minutes
651 | 669 | The Licence file has been copied to the SmartSEM USB Stick
651 | 669 | The EM Server.db has been copied to the SmartSEM USB Stick
650 | 10567 | Install Customer Licences
650 | 10567
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer.db to the Product Folder
650 | 10567 | Store EM Server.db
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
650 | 10567 | Add Licence to the SmartSEM USB Stick
650 | 10567 | Add EM Server.db to the SmartSEM USB Stick
43 | 671 | 1kV =
43 | 671 | Store 1KV Resolution Images and Pate image
43 | 671 | 15kV: (Spec = 1.1nm)
43 | 671 | Store 1KV Resolution Images and Pate image
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
651 | 631 | VP/EP Function Checked
651 | 631 | VP/EP Function Checked
651 | 632 | 2nd Rotary Pump is OFF in HV Mode
651 | 632 | 2nd Rotary Pump is ON in VP/EP Mode
651 | 669 | The Licence file has been copied to the SmartSEM USB Stick
651 | 669 | The EM Server.db has been copied to the SmartSEM USB Stick
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
29 | 11381 | Run overnight ESB Run Up & Photograph Macro
29 | 11381 | Analyse images for drift, breakdown, ESB issues
18 | 647 | Stand Off 1:
18 | 647 | Stand Off 2:
18 | 647 | Stand Off 3:
18 | 647 | Stand Off 4:
18 | 647 | Max Difference:
43 | 671 | 1kV =
43 | 671 | Store 1KV Resolution Images and Pate image
43 | 671 | 15kV: (Spec = 1.1nm)
43 | 671 | Store 1KV Resolution Images and Pate image
43 | 346 | 1kV =
43 | 346 | Store 1KV Resolution Images and Pate image
43 | 346 | 15kV =
43 | 346 | Store 1KV Resolution Images and Pate image
| 603 | Visual Inspection
| 603 | Liner Fitment
| 603 | End Cap Fitment
| 602 | Continuity ok
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
26 | 10965 | ESB Serial Number.
23 | 269 | Run ESB conditioning MACRO prior to switching on EHT.
26 | 281 | Type in measured value (Limit <= 30um)
26 | 281 | Store ESB Concentricity Image
5 | 189 | Ensure the detector is fitted the correct way up
9 | 103 | Log Aperture number as shown on pot on Aperture database (see link attached in 'Documents'
18 | 649 | Open Circuit (Coil – GND).
41 | 343 | Ensure Cladding geometry is correct and all parts aligned.
41 | 343 | Ensure the suppliers Cladding list for both kits is put into the customers pack.
41 | 343 | Fit safety stickers to cladding (Ensure stickers language is as oulined in sales order).
382 | 6928 | Macro Result ($[StageRepeatability])
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
26 | 282 | Type in measured value (limit: <= 15um)
26 | 282 | Store Inlens to ESB image
9 | 103 | Log Aperture number as shown on pot on Aperture database (see link attached in 'Documents')
18 | 651 | Objective Coil
18 | 651 | Scan Coil
18 | 652 | Confirm there is no Electical breakdown between Scan coils and Objective body.
| 604 | Visually inspected for damage/dents
| 604 | Sealing surfaces free from scratches/ imperfections
2 | 743 | Are there any shortages shown on pick list?
2 | 743 | Shortage PN
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
26 | 283 | Check for ESB charging and Flashing.
16 | 653 | Objective Coil S/N :
16 | 653 | Scan Coil S/N :
16 | 653 | Pole Piece S/N :
41 | 343 | Ensure Cladding geometry is correct and all parts aligned.
41 | 343 | Ensure the suppliers Cladding list for both kits is put into the customers pack.
41 | 343 | Fit safety stickers to cladding (Stickers language should match any additional user manuals listed on order).
43 | 671 | 1kV =
43 | 671 | Store 1KV Resolution Images and Pate image
43 | 671 | 15kV: (Spec = 1.1nm)
43 | 671 | Store 15KV Resolution Images and Pate image
| 605 | Position A:
| 605 | Position B:
| 605 | Position C:
| 605 | Position D:
| 605 | Max Difference:
4 | 748 | Are there any shortages?
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
40 | 1551 | Grid working
40 | 1551 | Check Cut Off
41 | 343 | Ensure Cladding geometry is correct and all parts aligned.
41 | 343 | Ensure the suppliers Cladding list for both kits is put into the customers pack.
41 | 343 | Fit safety stickers to cladding (Sticker language should match any additional user manuals listed on order).
43 | 671 | 1kV (Spec = 1.1nm)
43 | 671 | Store 1KV Resolution Images and Pate image
43 | 671 | 15kV =
43 | 671 | Store 15KV Resolution Images and Pate image
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
16 | 655 | Visual Inspection
16 | 655 | Liner Fitment
16 | 655 | End Cap Fitment
16 | 656 | Continuity ok
16 | 657 | Visually inspected for damage/dents
16 | 657 | Sealing surfaces free from scratches/ imperfections
16 | 658 | Position A:
16 | 658 | Position B:
16 | 658 | Position C:
16 | 658 | Position D:
16 | 658 | Max Difference:
16 | 659 | Resin is uniform and not proud
16 | 659 | No exposed wires
707 | 11283 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11283 | Multi_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11287 | Measure aperture size, 120um (+/- 2um)
707 | 11287 | Multi_TS_120um - Passed inspection criteria, Image Saved
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
16 | 660 | Pressure 15 PSI for 5 minutes
16 | 661 | Stand Off 1:
16 | 661 | Stand Off 2:
16 | 661 | Stand Off 3:
16 | 661 | Stand Off 4:
16 | 661 | Max Difference:
16 | 662 | Ensure objective coil is concentric to pole piece using alignment jig
16 | 664 | Open Circuit (Coil – GND).
16 | 665 | Objective Coil
16 | 665 | Scan Coil
43 | 346 | 1kV =
43 | 346 | Store 1KV Resolution Images and Pate image
43 | 346 | 15kV =
43 | 346 | Store 1KV Resolution Images and Pate image
| 606 | Resin is uniform and not proud
| 606 | No exposed wires
10 | 678 | Position A:
10 | 678 | Position B:
10 | 678 | Position C:
10 | 678 | Position D:
10 | 678 | Max. Difference
5 | 750 | Are there any shortages?
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
16 | 666 | Confirm there is no Electical breakdown between Scan coils and Objective body.
389 | 3314 | Type of Detector
389 | 3314 | Detector Serial number
389 | 3314 | DDP Module Serial Number
389 | 3314 | Fit Flange & Mount Detector
389 | 3314 | Install EDS PCBs into PC
389 | 3314 | Installation of DPP Module & Cables
389 | 3314 | Configure PC Network & Test
389 | 3314 | Disable Windows10 Fast Startup
389 | 3314 | Install ApexKey.edx Licences
389 | 3314 | Install Genesis.edx Licences (if Applicable)
389 | 3314 | Fit Dongle
389 | 3314 | APEX Software Version
389 | 3314 | Install APEX Software
389 | 3314 | Install Remoting
389 | 3314 | Setup RConfigure
389 | 3314 | Run API Test
389 | 3314 | Configure imaging – Run SGShell
389 | 3314 | Calibrate Detector – Run eDPPShell
389 | 3314 | Save Calibration Report
389 | 3314 | Detector Resolution - ≥ 129 eV
389 | 3314 | Save calibration data on exit
389 | 3314 | Run APEX
389 | 3314 | Check Spectrum
389 | 3314 | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected
389 | 3314 | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected
389 | 3314 | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected
389 | 3314 | Check Mapping
389 | 3314 | Check Line Scanning
389 | 3314 | Save Aquired Spectrum, Mapping and Line Scanning Images to Image Folder
389 | 3314 | Save Calibration Data to the Licence Folderon the SmartSem USB stick.
389 | 3314 | Save Calibration Data to the Image Folder
82 | 1592 | Take and save images at Mag 200, 3kX and 35kX.
82 | 1592 | Mag is within +/- 2% for each.
82 | 1592 | Orthognality is within +/- 1 deg for each.
43 | 671 | 1kV (Spec = 1.1nm)
43 | 671 | Store 1KV Resolution Images and Pate image
43 | 671 | 15kV =
43 | 671 | Store 15KV Resolution Images and Pate image
| 607 | Pressure 15 PSI for 5 minutes
656 | 752 | Are there any shortages?
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
52 | 484 | Aperture Finder run succesfully
43 | 671 | Please record the time taken to achieve 1kV resolution
43 | 671 | 1kV (Spec = 1.1nm)
43 | 671 | Store 1KV Resolution Images and Pate image
43 | 671 | 15kV =
43 | 671 | Store 15KV Resolution Images and Pate image
6 | 1547 | PAT Checkbox Calibration Date
6 | 1547 | Earth Bond Test TP1 (0.1Ω - 0.09Ω)
6 | 1547 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
6 | 1547 | Insulation Test (2.0MΩ - 1.9MΩ)
6 | 1547 | Earth Leakage Test (0.9mA - 0.7mA)
| 608 | Stand Off 1:
| 608 | Stand Off 2:
| 608 | Stand Off 3:
| 608 | Stand Off 4:
| 608 | Max Difference:
| 609 | Ensure objective coil is concentric to pole piece ( this should be done as best as possible by eye as no jig yet available)
674 | 10819 | Ion Pump Cables
674 | 10819 | Column X1 Lens cable
674 | 10819 | Objective Lens Cable
674 | 10819 | Scan Cable
674 | 10819 | Penning Cable
20 | 407 | Fit labels to PC serial cable com port 4 to 7
13 | 167 | Dowell pins have been mounted correctly to the mounting plate
13 | 168 | Driver A – Liner (>1000MΩ)
13 | 168 | Driver B – Liner (>1000MΩ)
13 | 168 | Driver A – Driver B (>1000MΩ)
663 | 414 | No Leaks Detected (Prioir to running up Turbo)
663 | 414 | No Leaks Detected (With Turbo running) when in hgh sensitivity
667 | 10773 | No Leaks Detected (Prior to running up Turbo)
667 | 10773 | No Leaks Detected (With Turbo running) when in high sensitivity
678 | 10855 | No Leaks Detected (Prior to running up Turbo)
678 | 10855 | No Leaks Detected (With Turbo running) when in high sensitivity
4 | 11384 | No Leaks Detected (Prior to running up Turbo)
4 | 11384 | No Leaks Detected (With Turbo running) when in high sensitivity
6 | 68 | No Leaks Detected (Prior to running up Turbo)
6 | 68 | No Leaks Detected (With Turbo running) when in high sensitivity
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | R INS - ( > 0.5MΩ)
667 | 11413 | I HLNC - ( < 500µA)
667 | 11413 | I HLSF - ( < 3500µA)
667 | 11413 | Confirm EN 61010 Passed
667 | 11413 | Confirm Visual Inspection Criterea Passed
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | R INS - ( > 0.5MΩ)
667 | 11413 | I HLNC - ( < 500µA)
667 | 11413 | I HLSF - ( < 3500µA)
667 | 11413 | Confirm EN 61010 Passed
667 | 11413 | Confirm Visual Inspection Criterea Passed
667 | 11413 | Record maximum leakage current
102 | 667 | Confirm Plasma Cleaner config
13 | 11190 | Driver A – Liner (>1000MΩ)
663 | 414 | No Leaks Detected (Prioir to running up Turbo)
663 | 414 | No Leaks Detected (With Turbo running) when in high sensitivity
663 | 414 | No Leaks Detected (Prior to running up Turbo)
663 | 414 | No Leaks Detected (With Turbo running) when in high sensitivity
674 | 416 | No Leaks Detected (Prior to running up Turbo)
674 | 416 | No Leaks Detected (With Turbo running) when in high sensitivity
6 | 628 | Secutest SN:
6 | 628 | Secutest Calibration Date:
6 | 628 | PE Test - Confirm all test points (<0.1Ω)
6 | 628 | R INS - ( > 0.5MΩ)
6 | 628 | I HLNC - ( < 500µA)
6 | 628 | I HLSF - ( < 3500µA)
6 | 628 | Confirm EN 61010 Passed
6 | 628 | Confirm Visual Inspection Criterea Passed
6 | 628 | Record maximum leakage current
6 | 628 | Secutest SN:
6 | 628 | Secutest Calibration Date:
6 | 628 | PE Test - Confirm all test points (<0.1Ω)
6 | 628 | R INS - ( > 0.5MΩ)
6 | 628 | I HLNC - ( < 500µA)
6 | 628 | I HLSF - ( < 3500µA)
6 | 628 | Confirm EN 61010 Passed
6 | 628 | Confirm Visual Inspection Criterea Passed
6 | 628 | Record maximum leakage current
6 | 625 | HAL SN:
6 | 625 | HAL Calibration Date:
6 | 625 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
| 611 | Open Circuit (Coil – GND).
13 | 171 | Driver A – Liner (>1000MΩ)
13 | 171 | Driver B – Liner (>1000MΩ)
13 | 171 | Driver A – Driver B (>1000MΩ)
| 612 | Objective Coil
| 612 | Scan Coil
| 613 | Confirm there is no Electical breakdown between Scan coils and Objective body.
10 | 680 | Has Transzorb been fitted
17 | 697 | Open Circuit (Coil – GND).
646 | 10529 | BSD Diode Pre-test
646 | 10529 | Select Type
646 | 10529 | Select Amplifier Type
646 | 10529 | Black Level Calibration
646 | 10529 | Check Diode Segment Function
646 | 10529 | Check All Gain ranges
646 | 10529 | Check Auto range
646 | 10529 | Check Auto Level
646 | 10529 | Setup and Check TOPO
646 | 10529 | Take 3 images - Atomic, 5pA, 2KV
646 | 10529 | Store BSD Functional Images-Atomic, 5pA, 2KV
389 | 3309 | Connect Silver Box and Cables
389 | 3309 | Download latest Version Of Software
389 | 3309 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
389 | 3309 | Request Licence File from Fibics
389 | 3309 | Install Atlas Software
389 | 3309 | Install Licence File
389 | 3309 | Calibrate SEM
389 | 3309 | Obtain Mosaic
389 | 3309 | Set ATLAS Rotation Offset
389 | 3309 | Save ATLAS Licence Document
13 | 172 | Driver A – Driver A In
13 | 172 | Driver B – Driver B In
13 | 172 | Anode – Liner In
678 | 10855 | No Leaks Detected (Prior to running up Turbo)
678 | 10855 | No Leaks Detected (With Turbo running) when in high sensitivity
654 | 11420 | PAT Tester Serial Number
654 | 11420 | PAT Tester Cal Expiry Date
654 | 11420 | Earth resistance – measured value must be < 0.1 Ω
654 | 11420 | Insulation resistance – measured value must be > 1.0M Ω
654 | 11420 | Test Certificate Completed & Signed
| 11375 | Position A:
| 11375 | Position B:
| 11375 | Position C:
| 11375 | Position D:
| 11375 | Max. Difference
10 | 681 | Continuity (Close Circuit)
10 | 681 | Confirm that the coil is NOT short circuit to the condenser body. Do not use test box, check manually.
389 | 3309 | Connect Silver Box and Cables
389 | 3309 | Download latest Version Of Software (see link in docs)
389 | 3309 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
389 | 3309 | Request Licence File from Fibics
389 | 3309 | Install Atlas Software
389 | 3309 | Install Licence File
389 | 3309 | Calibrate SEM
389 | 3309 | Obtain Mosaic
389 | 3309 | Set ATLAS Rotation Offset
389 | 3309 | Save ATLAS Licence Document
642 | 8699 | System Mains Voltage
642 | 8699
642 | 8699 | Cable Up
642 | 8699 | Label Cables
642 | 8699 | Tape PC Configuration Image to PC
642 | 8699 | System Mains Voltage
642 | 8699 | Cable Up
642 | 8699 | Label Cables
642 | 8699 | Tape PC Configuration Image to PC
642 | 8699 | System Mains Voltage
642 | 8699 | IF LaB6 and 120V - Transformer Installed
642 | 8699 | Cable Up
642 | 8699 | Label Cables
642 | 8699 | Tape PC Configuration Image to PC
13 | 11190 | Driver A – Liner (>1000MΩ)
667 | 10773 | No Leaks Detected (Prior to running up Turbo)
667 | 10773 | No Leaks Detected (With Turbo running) when in high sensitivity
10 | 682 | Continuity (Close Circuit)
10 | 682 | Coils to GND (Open Circuit)
17 | 698 | Objective Coil
17 | 698 | Scan Coil
646 | 10548 | Connect Silver Box and Cables
646 | 10548 | Download latest Version Of Software (see link in docs)
646 | 10548 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
646 | 10548 | Request Licence File from Fibics
646 | 10548 | Install Atlas Software
646 | 10548 | Install Licence File
646 | 10548 | Calibrate SEM
646 | 10548 | Obtain Mosaic
646 | 10548 | Set ATLAS Rotation Offset
646 | 10548 | Save ATLAS Licence Document
40 | 10868 | Connect Silver Box and Cables
40 | 10868 | Download latest Version Of Software
40 | 10868 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
40 | 10868 | Request Licence File from Fibics
40 | 10868 | Install Atlas Software
40 | 10868 | Install Licence File
40 | 10868 | Calibrate SEM
40 | 10868 | Obtain Mosaic
40 | 10868 | Set ATLAS Rotation Offset
40 | 10868 | Save ATLAS Licence Document
646 | 10548 | Connect Silver Box and Cables
646 | 10548 | Download latest Version Of Software (see link in docs)
646 | 10548 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
646 | 10548 | Request Licence File from Fibics
646 | 10548 | Install Atlas Software
646 | 10548 | Install Licence File
646 | 10548 | Calibrate SEM
646 | 10548 | Obtain Mosaic
646 | 10548 | Set ATLAS Rotation Offset
646 | 10548 | Save ATLAS Licence Document
17 | 699 | Confirm there is no Electical breakdown between Scan coils and Objective body.
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
646 | 10545 | Correct PC - 347823-9024-660 Installed
646 | 10545 | Print Out SmartPI Check List (Link in Documents)
646 | 10545 | Install SmartSEM Licences
646 | 10545 | Install MIL Redistribution
646 | 10545 | Fit MIL Dongle
646 | 10545 | SmartPI Software Version
646 | 10545 | Install SmartPI Software
646 | 10545 | I/O-SCAN Card Fitted (Bruker Only)
646 | 10545 | I/O-SCAN Card Calibrated
646 | 10545 | Setup & Run Application
646 | 10545 | Check Results & Image Montage
646 | 10545 | Store Compressed Database,Montage Image, Log files etc
2 | 6 | Set regulator pressure indicator to 0.03MPa
2 | 6 | Set regulator pressure to 0.03MPa
20 | 252 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448 , then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Run ASAP Screening Program
20 | 252 | System Specific Datazone Loaded
20 | 11021 | PC Serial Number:
20 | 11021 | Connect Up Raith Test Tool PC
20 | 11021 | Unistall SmartSEM
20 | 11021 | Install SmartSEM Version
20 | 11021 | Install Latest Service Pack
20 | 11021 | Install Gemini Support Tools-Version
20 | 11021 | Run ASAP Screening Program
20 | 11021 | System Specific Datazone Loaded
382 | 6928 | Macro Result ($[StageRepeatability])
17 | 683 | Objective Coil S/N :
17 | 683 | Scan Coil S/N :
17 | 683 | Pole Piece S/N :
389 | 89 | Test Valves Using Raman Test Interlock Box
389 | 89 | Install SemRamanStage Software
389 | 89 | Run SEM Raman Correlation
646 | 10545 | Correct PC - 347823-9024-660 Installed
646 | 10545 | Print Out SmartPI Check List (Link in Documents)
646 | 10545 | Install SmartSEM Licences
646 | 10545 | Install MIL Redistribution
646 | 10545 | Fit MIL Dongle
646 | 10545 | SmartPI Software Version
646 | 10545 | Install SmartPI Software
646 | 10545 | I/O-SCAN Card Fitted (Bruker Only)
646 | 10545 | I/O-SCAN Card Calibrated
646 | 10545 | Setup & Run Application
646 | 10545 | Check Results & Image Montage
646 | 10545 | Store Compressed Database,Montage Image, Log files etc
20 | 252 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448 , then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Run ASAP Screening Program
20 | 252 | System Specific Datazone Loaded
46 | 1568 | Power system down.
46 | 1568 | Blow water out of the cooling system.
46 | 1568 | Put thread lock on pendulum screws and transit block Screws.
46 | 1568 | Chamber secured to vibration plate
46 | 1568 | Chamber Sercuring Screws Tightened to 25NM
46 | 1568 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 1568 | Secure Airfeet with foam blocks.
46 | 1568 | Fit foam blocks around column cladding.
46 | 1568 | Secure Turbo pump with foam blocks.
46 | 1568 | Fit Covers to all unused connectors.
46 | 1568 | Remove Cekon Plug From Mains Cable
46 | 1568 | Fit Live and Neutral Labels to Mains Cable
46 | 1568 | Mains and Earth cable stowed on Mains unit.
46 | 1568 | Fit shipping Eyes to transit blocks.
46 | 1568 | Confirm all column cladding is secure and screws Loctited in place.
46 | 1568 | All doors closed and secured.
46 | 1568 | All plinth panels removed and stored in transit trolley.
46 | 1568 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 1568 | Print shipping stickers.
46 | 1568 | Fit M8 screw to secure stage door to chamber.
25 | 704 | Gun X
25 | 704 | Gun Y
25 | 704 | Aper X
25 | 704 | Aper Y
389 | 89 | Test Valves Using Raman Test Interlock Box
389 | 89 | Install SemRamanStage Software
389 | 89 | Run SEM Raman Correlation
389 | 89 | InstaII M I L 10 Software
646 | 10654 | Print Out Mineralogic Check List (Link in Documents)
646 | 10654 | Correct PC - Installed
646 | 10654 | Install SmartSEM Licences
646 | 10654 | Instal MIL Redistribution
646 | 10654 | Fit MIL Dongle
646 | 10654 | Install Mineralogic Software
646 | 10654 | Calibrate I/O-SCAN card
646 | 10654 | Setup & Run Application
646 | 10654 | Check Results & Image Montage
646 | 10654 | 24 Hour Run
646 | 10654 | Store Compressed Database,Montage Image, Log files etc
20 | 11038 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
20 | 11038 | PC S/N
20 | 11038 | If SEM Enhanced PC(349535-9001-020) has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Hard disk Image (Version)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
46 | 11123 | Power system down.
46 | 11123 | Blow water out of the cooling system.
46 | 11123 | Chamber secured to vibration plate (N/A to XB 550L)
46 | 11123 | Chamber securing screws tightened to 25nm (N/A to XB 550L)
46 | 11123 | Chamber securing screws tightness checked. Type your initials (For XB 550L type N/A)
46 | 11123 | Secure Airfeet with foam blocks.
46 | 11123 | Secure Turbo pump with foam blocks.
46 | 11123 | Fit Covers to all unused connectors.
46 | 11123 | Mains and Earth cable stowed on Mains unit.
46 | 11123 | Fit shipping Eyes to transit blocks.
46 | 11123 | Put thread lock on pendulum screws and transit block Screws.
46 | 11123 | All doors closed and secured.
46 | 11123 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11123 | Print shipping stickers.
46 | 11123 | Fit 8x M6 x 55mm cap head screws to stage door
28 | 700 | Confirm EM Server screeening programme is installed and running
28 | 700 | Confirm no repeated EM Server screening errors have been flagged
20 | 425 | Record MAC adress (LAN 1)
20 | 425 | Record MAC adress (LAN 2)
671 | 754 | Are there any shortages?
389 | 89 | Test Valves Using Raman Test Interlock Box
389 | 89 | Install SemRamanStage Software
389 | 89 | Run SEM Raman Correlation
389 | 89 | InstaII M I L 10 Software
389 | 89 | Run Smart Stage Mapping
646 | 10654 | Print Out Mineralogic Check List (Link in Documents)
646 | 10654 | Correct PC - Installed
646 | 10654 | Install SmartSEM Licences
646 | 10654 | Instal MIL Redistribution
646 | 10654 | Fit MIL Dongle
646 | 10654 | Install Mineralogic Software
646 | 10654 | Calibrate I/O-SCAN card
646 | 10654 | Setup & Run Application
646 | 10654 | Check Results & Image Montage
646 | 10654 | 24 Hour Run
646 | 10654 | Store Compressed Database,Montage Image, Log files etc
646 | 10545 | Correct PC - 347823-9024-660 Installed
646 | 10545 | Print Out SmartPI Check List (Link in Documents)
646 | 10545 | Install SmartSEM Licences
646 | 10545 | Install MIL Redistribution
646 | 10545 | Fit MIL Dongle
646 | 10545 | SmartPI Software Version
646 | 10545 | Install SmartPI Software
646 | 10545 | I/O-SCAN Card Fitted (Bruker Only)
646 | 10545 | I/O-SCAN Card Calibrated
646 | 10545 | Setup & Run Application
646 | 10545 | Check Results & Image Montage
646 | 10545 | Store Compressed Database,Montage Image, Log files etc
46 | 11123 | Power system down.
46 | 11123 | Blow water out of the cooling system.
46 | 11123 | Chamber secured to vibration plate (N/A to XB 550L)
46 | 11123 | Chamber securing screws tightened to 25nm (N/A to XB 550L)
46 | 11123 | Chamber securing screws tightness checked. Type your initials (For XB 550L type N/A)
46 | 11123 | Secure Airfeet with foam blocks.
46 | 11123 | Secure Turbo pump with foam blocks.
46 | 11123 | Fit Covers to all unused connectors.
46 | 11123 | Mains and Earth cable stowed on Mains unit.
46 | 11123 | Fit shipping Eyes to transit blocks.
46 | 11123 | Put thread lock on pendulum screws and transit block Screws.
46 | 11123 | All doors closed and secured.
46 | 11123 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11123 | Print shipping stickers.
46 | 11123 | Fit 8x M6 x 55mm cap head screws to stage door
21 | 256 | Store System checker full report
17 | 685 | 12 hours drying time
389 | 89 | Test Valves Using Raman Test Interlock Box
389 | 89 | Install SemRamanStage Software
389 | 89 | Run SEM Raman Correlation
389 | 89 | Save Calibration File as Raman
389 | 89 | InstaII M I L 10 Software
389 | 89 | Run Smart Stage Mapping
389 | 3316 | Correct PC - 347823-9024-660 Installed
389 | 3316 | Print Out SmartPI Check List (Link in Documents)
389 | 3316 | Install SmartSEM Licences
389 | 3316 | Install MIL Redistribution
389 | 3316 | Fit MIL Dongle
389 | 3316 | SmartPI Software Version
389 | 3316 | Install SmartPI Software
389 | 3316 | Calibrate I/O-SCAN card
389 | 3316 | Setup & Run Application
389 | 3316 | Check Results & Image Montage
389 | 10874 | Print Out Mineralogic Check List (Link in Documents)
389 | 10874 | Correct PC - Installed
389 | 10874 | Install SmartSEM Licences
389 | 10874 | Instal MIL Redistribution
389 | 10874 | Fit MIL Dongle
389 | 10874 | Install Mineralogic Software
389 | 10874 | Calibrate I/O-SCAN card
389 | 10874 | Setup & Run Application
389 | 10874 | Check Results & Image Montage
389 | 10874 | 24 Hour Run
9 | 102 | Aperture cleaned using IPA and 1mm brush.
9 | 102 | Aperture run through ultra-sonic tank (in upright position)
9 | 102 | Aperture inspected and confimred clean. Eng 1 Initials:
9 | 102 | Aperture inspected and confimred clean. Eng 2 Initials:
3 | 28 | 346032-8030-100 - Please check continuity of EO Dynamic Cable using DVM ---- Pin 2 - Pin 2
3 | 28 | 346032-8027-100 - Please check continuity of EO Static Cable using DVM ---- Pin 2 - Pin 2
20 | 11021 | PC Serial Number:
20 | 11021 | Connect Up Raith Test Tool PC
20 | 11021 | Unistall SmartSEM
20 | 11021 | Install SmartSEM Version
20 | 11021 | Install Latest Service Pack
20 | 11021 | Install Gemini Support Tools-Version
20 | 11021 | PC Serial Number:
20 | 11021 | Connect Up Raith Test Tool PC
20 | 11021 | Unistall SmartSEM
20 | 11021 | Install SmartSEM Version
20 | 11021 | Install Latest Service Pack
20 | 11021 | Install Gemini Support Tools-Version
20 | 11021 | Run ASAP Screening Program
28 | 290 | Gun running to Datasheet
28 | 290 | Gun Vacuum acceptable (EHT on/off)
28 | 290 | Nominal Ext V acceptable
9 | 102 | Aperture cleaned using IPA and 1mm brush (000000-0695-322)
9 | 102 | Aperture run through ultra-sonic tank (in upright position)
9 | 102 | Aperture inspected and confimred clean. Eng 1 Initials:
9 | 102 | Aperture inspected and confimred clean. Eng 2 Initials:
9 | 102 | Aperture cleaned using IPA and 1mm brush
9 | 102 | Aperture run through ultra-sonic tank (in upright position)
9 | 102 | Aperture inspected and confimred clean. Eng 1 Initials:
9 | 102 | Aperture inspected and confimred clean. Eng 2 Initials:
46 | 11407 | Power system down.
46 | 11407 | Blow water out of the cooling system.
46 | 11407 | Chamber secured to vibration plate
46 | 11407 | Chamber Sercuring Screws Tightened to 25NM
46 | 11407 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 11407 | Check that the Stage bearing blocks screws are secure.
46 | 11407 | Secure Airfeet with foam blocks.
46 | 11407 | Secure Turbo pump with foam blocks.
46 | 11407 | Fit Covers to all unused connectors.
46 | 11407 | Remove Cekon Plug From Mains Cable
46 | 11407 | Fit Live and Neutral Labels to Mains Cable
46 | 11407 | Mains and Earth cable stowed on Mains unit.
46 | 11407 | Fit shipping Eyes to transit blocks.
46 | 11407 | Put thread lock on pendulum screws and transit block Screws.
46 | 11407 | All doors closed and secured.
46 | 11407 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11407 | Print shipping stickers.
46 | 1568 | Power system down.
46 | 1568 | Blow water out of the cooling system.
46 | 1568 | Put thread lock on pendulum screws and transit block Screws.
46 | 1568 | Chamber secured to vibration plate
46 | 1568 | Chamber Sercuring Screws Tightened to 25NM
46 | 1568 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 1568 | Secure Airfeet with foam blocks.
46 | 1568 | Fit foam blocks around column cladding.
46 | 1568 | Secure Turbo pump with foam blocks.
46 | 1568 | Fit Covers to all unused connectors.
46 | 1568 | Remove Cekon Plug From Mains Cable
46 | 1568 | Fit Live and Neutral Labels to Mains Cable
46 | 1568 | Mains and Earth cable stowed on Mains unit.
46 | 1568 | Fit shipping Eyes to transit blocks.
46 | 1568 | Confirm all column cladding is secure and screws Loctited in place.
46 | 1568 | All doors closed and secured.
46 | 1568 | All plinth panels removed and stored in transit trolley.
46 | 1568 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 1568 | Fit M8 screw to secure stage door to chamber.
46 | 1568 | Fit dust covers andsecure / remove cables.
46 | 1568 | Prepare column cladding transport jig .
46 | 1568 | Print shipping stickers .
46 | 1568 | Power system down.
46 | 1568 | Blow water out of the cooling system.
46 | 1568 | Put thread lock on pendulum screws and transit block Screws.
46 | 1568 | Chamber secured to vibration plate
46 | 1568 | Chamber Sercuring Screws Tightened to 25NM
46 | 1568 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 1568 | Secure Airfeet with foam blocks.
46 | 1568 | Fit foam blocks around column cladding.
46 | 1568 | Secure Turbo pump with foam blocks.
46 | 1568 | Fit Covers to all unused connectors.
46 | 1568 | Remove Cekon Plug From Mains Cable
46 | 1568 | Fit Live and Neutral Labels to Mains Cable
46 | 1568 | Mains and Earth cable stowed on Mains unit.
46 | 1568 | Fit shipping Eyes to transit blocks.
46 | 1568 | Confirm all column cladding is secure and screws Loctited in place.
46 | 1568 | All doors closed and secured.
46 | 1568 | All plinth panels removed and stored in transit trolley.
46 | 1568 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 1568 | Fit M8 screw to secure stage door to chamber.
46 | 1568 | Fit dust covers and secure / remove cables.
46 | 1568 | Prepare column cladding transport jig .
46 | 1568 | Print shipping stickers .
28 | 291 | Mechanical Alignment ok
28 | 291 | Column clamping secure
674 | 756 | Are there any shortages?
17 | 686 | Resistance:
28 | 409 | Gun X
28 | 409 | Gun Y
28 | 409 | Aper X
28 | 409 | Aper Y
28 | 292 | Cut off ok
28 | 292 | No Charging particles
28 | 294 | Run short image jump macro and analyse images
28 | 10647 | Confirm SEM Status that C2 & C3 Cal and System Align Offset Angle are sensible
403 | 6935 | Record MAC address (LAN 1)
403 | 6935 | Record MAC address (LAN 2)
403 | 6935 | Network access requested
403 | 6935 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
20 | 425 | Record MAC address (LAN 1)
20 | 425 | Record MAC address (LAN 2)
285 | 3242 | Select PC type
285 | 3242 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
285 | 3242 | PC S/No.
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Select PC type
285 | 3242 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
285 | 3242 | PC S/No.
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Upgrade ASAP Firmware (check Service Pack for latest version) and API
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
12 | 165 | Filament OK (Green LED Lit)
12 | 165 | Extractor - Suppressor ok, (No Red LED Lit)
12 | 165 | Gun head driven to four corners (No Red LED Lit)
382 | 6928 | Macro Result ($[StageRepeatability])
17 | 687 | Visual Inspection
17 | 687 | Liner Fitment
17 | 687 | End Cap Fitment
37 | 705 | Stigs: 30um Aperture, All EHT ≤20%
37 | 705 | Stigs: Off Axis Apertures, All EHT ≤30%
37 | 705 | Aper & Gun: 30um Aperture, All EHT ≤50%
37 | 705 | Stigs: 30um Aperture, All EHT ≤20%
37 | 705 | Stigs: Off Axis Apertures, All EHT ≤30%
37 | 705 | Aper & Gun: 30um Aperture, All EHT ≤50%
37 | 705 | Stigs: 30um Aperture, All EHT ≤20%
37 | 705 | Stigs: Off Axis Apertures, All EHT ≤30%
37 | 705 | Aper & Gun: 30um Aperture, All EHT ≤50%
285 | 708 | Record MAC address (LAN 1)
285 | 708 | Record MAC address (LAN 2)
20 | 425 | Record MAC address (LAN 1)
20 | 425 | Record MAC address (LAN 2)
651 | 10602 | Check calibration in each mag range
651 | 10602 | Check Calibration in each mag range
651 | 10602 | Check Stage Angle has been setup correctly
642 | 10485 | Approx. No. of Hours
6 | 710 | Actual
6 | 710 | Portable IGP controller fitted? (See work instruction for details)
382 | 6928 | Macro Result ($[StageRepeatability])
642 | 10485 | Approx. No. of Hours
20 | 1133 | Print Shipment Labels
20 | 1133 | Label Material Boxes
20 | 1133 | Place left over stickes in Document Holder
42 | 11047 | Actual X-Axis
42 | 11047 | Store X ripple shots and results
42 | 11047 | Actual Y-Axis
42 | 11047 | Store Y ripple shots and results
285 | 708 | Record MAC address (LAN 1)
285 | 708 | Record MAC address (LAN 2)
400 | 3405 | X-Ray Leakage Check
400 | 3405 | Quality Certificate
400 | 3405 | Gun/Aperture Align Sheet
400 | 3405 | Denka Data Sheet
400 | 3405 | Test Accessories Packing list
400 | 3405 | Standard Packing List
400 | 3405 | Items retained on Microscope Packing List
400 | 3405 | Items removed from Microscope Packing List
400 | 3405 | China RoHS Certificate
17 | 711 | Objective Coil S/N :
17 | 711 | Scan Coil S/N :
17 | 711 | Pole Piece S/N :
382 | 6928 | Macro Result ($[StageRepeatability])
18 | 397 | Check SCS Folder/Sales Order
18 | 397 | SCS Number (If applicable)
5 | 189 | Ensure the detector is fitted the correct way up
676 | 10826 | Serial Number
676 | 10826 | Shim Stage Door
676 | 10826 | Set Stage Clamp Height - Using Jig
| 522 | Ensure objective coil is concentric to pole piece ( this should be done as best as possible by eye as no jig yet available)
17 | 688 | Continuity ok
44 | 358 | All verification & Safety tests done and within threshold limits defined
44 | 358 | Sales Order is latest version
44 | 358 | All items on Sales order accounted for
44 | 358 | ESB SN matches, if applicable.
44 | 706 | Confirm ESB SN matches
676 | 10835 | SPP 360 Completed
676 | 10836 | SPP 360 Completed
5 | 11005 | SPP 360 Completed
665 | 10750 | SPP 360 Completed
709 | 11322 | Weight
709 | 11322 | Dimensions Length
709 | 11322 | Dimensions Width
709 | 11322 | Dimensions Height
709 | 11322 | Security Seals Serial No
709 | 11322 | Shock Watch Serial Number
709 | 11322 | Tilt Watch Serial No.
709 | 11322 | Packed By
17 | 689 | Visually inspected for damage/dents
17 | 689 | Sealing surfaces free from scratches/ imperfections
44 | 706 | Confirm ESB SN matches
40 | 1561 | EHT Cal - Note: The 'Stage Bias mirror Calib' values should be < 30 V
40 | 1561 | Spark Alarm
40 | 1561 | Resolution Preperation
40 | 1561 | Acquire a customer proof image @ 1kV landing energy 5kV bias on gold on carbon (Resolution not required)
285 | 3242 | Select PC type
285 | 3242 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
285 | 3242 | PC S/No.
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 688330
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 688330, then the PC needs reimaging – follow the attached WI. The correct image version is V1.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
17 | 713 | Visual Inspection
17 | 713 | Liner Fitment
17 | 713 | End Cap Fitment
25 | 427 | Confirm that the contents of Tablefiller are deleted before starting the procedure. Stored values can affect the alignment routine.
25 | 427 | Mechanically align the column using the Merlin Alignment Wizard
40 | 1562 | Run Clean Cycle
18 | 728 | Initials of 2nd Engineer confirming.
| 733 | Initials of 2nd Engineer confirming.
| 734 | Initials of 2nd Engineer confirming.
44 | 706 | Confirm ESB serial number matches.
44 | 706 | Confirm ESB SN matches
81 | 11112 | 10 kV, central aperture, High Current, 5mm WD. Specimen Current ≥400pA. Measured value =
81 | 11112 | Max Measured value =
102 | 1627 | Confrirm the 'Stage Bias mirror Calib' values are < 30 V
102 | 1627 | Confirm that the saved Resolution image is acceptable
17 | 714 | Continuity ok
17 | 715 | Visually inspected for damage/dents
17 | 715 | Sealing surfaces free from scratches/ imperfections
17 | 716 | Position A:
17 | 716 | Position B:
17 | 716 | Position C:
17 | 716 | Position D:
17 | 716 | Max Difference:
18 | 728 | Initials of 2nd Engineer confirming.
7 | 729 | Initials of 2nd Engineer confirming.
16 | 730 | Initials of 2nd Engineer confirming.
| 732 | Initials of 2nd Engineer confirming.
| 735 | Initials of 2nd Engineer confirming.
13 | 172 | Driver A – Driver A In
13 | 172 | Driver B – Driver B In
13 | 172 | Anode – Liner In
17 | 690 | Position A:
17 | 690 | Position B:
17 | 690 | Position C:
17 | 690 | Position D:
17 | 690 | Max Difference:
44 | 358 | All verification & Safety tests done and within threshold limits defined
44 | 358 | Sales Order is latest version
44 | 358 | All items on Sales order accounted for
81 | 11112 | 10 kV, central aperture, High Current, 5mm WD. Specimen Current ≥400pA. Measured value =
81 | 11112 | Max Probe Current. 30 kV, Aperture Largest and WD 5 mm (Analytic gun mode): Limit: 20nA Measured value =
81 | 11112 | 10 kV, central aperture, High Current, 5mm WD. Specimen Current ≥400pA. Measured value =
81 | 11112 | Max Probe Current. 30 kV, Aperture Largest and WD 5 mm (Analytic gun mode): Limit: 20nA Measured value =
17 | 717 | Resin is uniform and not proud
17 | 717 | No exposed wires
17 | 718 | Pressure 15 PSI for 5 minutes
17 | 719 | Stand Off 1:
17 | 719 | Stand Off 2:
17 | 719 | Stand Off 3:
17 | 719 | Stand Off 4:
17 | 719 | Max Difference:
17 | 731 | Initials of 2nd Engineer confirming.
13 | 171 | Driver A – Liner (>1000MΩ)
13 | 171 | Driver B – Liner (>1000MΩ)
13 | 171 | Driver A – Driver B (>1000MΩ)
| 524 | Open Circuit (Coil – GND).
676 | 757 | Are there any shortages?
17 | 691 | Position A:
17 | 691 | Position B:
17 | 691 | Position C:
17 | 691 | Position D:
17 | 691 | Max Difference:
44 | 706 | Confirm ESB serial number matches.
81 | 1587 | Measured value =
7 | 729 | Initials of 2nd Engineer confirming.
13 | 736 | Pumping adapter has been fitted before DC EBBU post fitment checks
| 526 | Objective Coil
| 526 | Scan Coil
| 525 | Confirm there is no Electical breakdown between Scan coils and Objective body.
17 | 714 | Continuity ok
17 | 718 | Pressure 15 PSI for 5 minutes
4 | 629 | Ensure Pirani gauge is isolated from the roughing line
4 | 629 | Ensure Pirani gauge is isolated from the roughing line
4 | 629 | Ensure Pirani gauge is isolated from the roughing line
4 | 629 | For the APG100-XM Remove the 2 metal clips, but for the APG200-XM Leave them in
707 | 11172 | SF_TS_HRQA - Passed inspection criteria, Image Saved
44 | 706 | All verification & Safety tests done and within threshold limits defined
44 | 706 | Sales Order is latest version
44 | 706 | All items on Sales order accounted for
81 | 1571 | Measured value =
17 | 725 | 12 hours drying time
17 | 726 | Resistance:
17 | 711 | Objective Coil S/N :
17 | 711 | Scan Coil S/N :
17 | 711 | Pole Piece S/N :
17 | 713 | Visual Inspection
17 | 713 | Liner Fitment
17 | 713 | End Cap Fitment
17 | 718 | Pressure 15 PSI for 5 minutes
17 | 719 | Stand Off 1:
17 | 719 | Stand Off 2:
17 | 719 | Stand Off 3:
17 | 719 | Stand Off 4:
17 | 719 | Max Difference:
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal
40 | 739 | Black Level Cal
40 | 739 | Mechanical Alignment
40 | 739 | Diode Segment Test
40 | 739 | Diode Mirror Test
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | All Gain Ranges working
40 | 739 | Check all channels in Invert mode
659 | 760 | Are there any shortages?
17 | 692 | Resin is uniform and not proud
17 | 692 | No exposed wires
17 | 693 | Pressure 15 PSI for 5 minutes
81 | 1588 | Measured value =
16 | 730 | Initials of 2nd Engineer confirming.
17 | 714 | Continuity ok
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal
40 | 739 | Black Level Cal
40 | 739 | Mechanical Alignment
40 | 739 | Diode Segment Test
40 | 739 | Diode Mirror Test
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | All Gain Ranges working
40 | 739 | Check all channels in Invert mode
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Black Level Cal
40 | 739 | Mechanical Alignment
40 | 739 | Diode Segment Test
40 | 739 | Diode Mirror Test
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | All Gain Ranges working
40 | 739 | Check all channels in Invert mode
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Mechanical Alignment
40 | 739 | Diode Segment Test
40 | 739 | Diode Mirror Test
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | All Gain Ranges working
40 | 739 | Check all channels in Invert mode
661 | 762 | Are there any shortages?
81 | 1589 | Measured value =
17 | 727 | Position A:
17 | 727 | Position B:
17 | 727 | Position C:
17 | 727 | Position D:
17 | 727 | Max Difference:
17 | 731 | Initials of 2nd Engineer confirming.
17 | 715 | Visually inspected for damage/dents
17 | 715 | Sealing surfaces free from scratches/ imperfections
17 | 716 | Position A:
17 | 716 | Position B:
17 | 716 | Position C:
17 | 716 | Position D:
17 | 716 | Max Difference:
17 | 717 | Resin is uniform and not proud
17 | 717 | No exposed wires
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Mechanical Alignment
40 | 739 | Diode Segment Test
40 | 739 | Diode Mirror Test
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | All Gain Ranges working
40 | 739 | Check all channels in Invert mode
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal to Noise Test (Find program in attached files)
40 | 739 | Diode Mirror Test
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | All Gain Ranges working
40 | 739 | Check all channels in Invert mode
17 | 716 | Position A:
17 | 716 | Position B:
17 | 716 | Position C:
17 | 716 | Position D:
17 | 716 | Max Difference:
17 | 694 | Stand Off 1:
17 | 694 | Stand Off 2:
17 | 694 | Stand Off 3:
17 | 694 | Stand Off 4:
17 | 694 | Max Difference:
81 | 11112 | Measured value =
81 | 11112 | Max Probe Current. 30 kV, Aperture Largest and WD 5 mm (Analytic gun mode): Limit: 20nA Measured value =
17 | 720 | Ensure objective coil is concentric to pole piece ( this should be done as best as possible by eye as no jig yet available)
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal to Noise Test (Find program in attached files)
40 | 739 | Diode Mirror Test
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | All Gain Ranges working
40 | 739 | All Gain Ranges working
40 | 739 | Check all channels in Invert mode
17 | 717 | Resin is uniform and not proud
17 | 717 | No exposed wires
17 | 714 | Continuity ok
707 | 11173 | Measure aperture size, 35um (+/- 2um)
663 | 764 | Are there any shortages?
665 | 766 | Are there any shortages?
17 | 695 | Ensure objective coil is concentric to pole piece ( this should be done as best as possible by eye as no jig yet available)
81 | 11112 | Measured value =
17 | 722 | Open Circuit (Coil – GND).
17 | 724 | Confirm there is no Electical breakdown between Scan coils and Objective body.
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal to Noise Test (Find program in attached files)
40 | 739 | Diode Mirror Test
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | All Gain Ranges working
40 | 739 | All Gain Ranges working
40 | 739 | Check all channels in Invert mode
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Diode Mirror Test
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | All Gain Ranges working
40 | 739 | All Gain Ranges working
40 | 739 | Check all channels in Invert mode
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal To Noise Test
40 | 739 | Diode Mirror Test
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | All Gain Ranges working
40 | 739 | All Gain Ranges working
40 | 739 | Check all channels in Invert mode
17 | 719 | Stand Off 1:
17 | 719 | Stand Off 2:
17 | 719 | Stand Off 3:
17 | 719 | Stand Off 4:
17 | 719 | Max Difference:
81 | 707 | Measured value =
17 | 723 | Objective Coil
17 | 723 | Scan Coil
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal To Noise Test
40 | 739 | Diode Mirror Test
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | All Gain Ranges working
40 | 739 | All Gain Ranges working
40 | 739 | Check all channels in Invert mode
20 | 11038 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
20 | 11038 | PC S/N
20 | 11038 | If SEM Enhanced PC(349535-9001-020) has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Hard disk Image (Version)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
707 | 758 | SF_TS_HRQA - Passed inspection criteria, Image Saved
40 | 739 | Record Diode Serial Number:
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal To Noise Test (Result >12%)
40 | 739 | Store Signal To Noise Result (Screen-shot)
40 | 739 | Set Up Diode Protection Plate
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | Fit Detector Mechanics Cover
81 | 707 | Measured value =
2 | 5 | Record tmeperature probe resistance
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal To Noise Test (Result >12%)
40 | 739 | Diode Mirror Test
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | All Gain Ranges working
40 | 739 | All Gain Ranges working
40 | 739 | Check all channels in Invert mode
40 | 739 | Record Diode Serial Number:
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal To Noise Test (Result >12%)
40 | 739 | Diode Mirror Test
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | All Gain Ranges working
40 | 739 | All Gain Ranges working
40 | 739 | Check all channels in Invert mode
40 | 739 | Record Diode Serial Number:
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal To Noise Test (Result >12%)
40 | 739 | Diode Mirror Test
40 | 739 | Set Up Diode Protection Plate
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | All Gain Ranges working
40 | 739 | All Gain Ranges working
40 | 739 | Check all channels in Invert mode
707 | 758 | Measure aperture size, 35um (+/- 2um)
40 | 739 | Record Diode Serial Number:
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal To Noise Test (Result >12%)
40 | 739 | Store Gun monitor trace as Probestability.mn2
40 | 739 | Diode Mirror Test
40 | 739 | Set Up Diode Protection Plate
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | All Gain Ranges working
40 | 739 | All Gain Ranges working
40 | 739 | Check all channels in Invert mode
707 | 11172 | SF_TS_HRQA - Passed inspection criteria, Image Saved
43 | 346 | 1kV =
43 | 346 | Store 1KV Resolution Images and Pate image
43 | 346 | 15kV =
43 | 346 | Store 1KV Resolution Images and Pate image
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
705 | 11057 | UPS Open X Interface (X121) is Functioning
705 | 11057 | Function Checked
705 | 11057 | Operation of UPS switch , V3 shuts, System Vents and System Shuts Down to Red Button
2 | 740 | Record tmeperature probe resistance
40 | 739 | Record Diode Serial Number:
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal To Noise Test (Result >12%)
40 | 739 | Store Signal To Noise Result Sheet
40 | 739 | Diode Mirror Test
40 | 739 | Set Up Diode Protection Plate
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | All Gain Ranges working
40 | 739 | All Gain Ranges working
40 | 739 | Check all channels in Invert mode
707 | 11173 | Measure aperture size, 35um (+/- 2um)
707 | 11173 | Measure aperture size, 35um (+/- 2um)
643 | 759 | Run EO Test
643 | 759 | Run EO Test
643 | 10488 | Run EO Test
643 | 10488 | X Gun Tilt ≤50%
643 | 10488 | Y Gun Tilt ≤50%
643 | 10488 | X Gun Shift ≤50%
643 | 10488 | Y Gun Shift ≤50%
643 | 10488 | X Gun Tilt ≤50%
643 | 10488 | Y Gun Tilt ≤50%
643 | 10488 | X Gun Shift ≤50%
643 | 10488 | Y Gun Shift ≤50%
46 | 11407 | Power system down.
46 | 11407 | Blow water out of the cooling system.
46 | 11407 | Chamber secured to vibration plate
46 | 11407 | Chamber Sercuring Screws Tightened to 25NM
46 | 11407 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 11407 | Check that the Stage bearing blocks screws are secure.
46 | 11407 | Remove test tool stage
46 | 11407 | Remove test tool stage PCB
46 | 11407 | Secure Airfeet with foam blocks.
46 | 11407 | Secure Turbo pump with foam blocks.
46 | 11407 | Fit Covers to all unused connectors.
46 | 11407 | Remove Cekon Plug From Mains Cable
46 | 11407 | Fit Live and Neutral Labels to Mains Cable
46 | 11407 | Mains and Earth cable stowed on Mains unit.
46 | 11407 | Fit shipping Eyes to transit blocks.
46 | 11407 | Put thread lock on pendulum screws and transit block Screws.
46 | 11407 | All doors closed and secured.
46 | 11407 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11407 | Print shipping stickers.
2 | 5 | Record the temp sensor probe resistance
40 | 739 | Record Diode Serial Number:
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal To Noise Test (Result >12%)
40 | 739 | Store Signal To Noise Result Sheet
40 | 739 | Set Up Diode Protection Plate
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | Record Diode Serial Number:
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal To Noise Test (Result >12%)
40 | 739 | Store Signal To Noise Result Sheet
40 | 739 | Set Up Diode Protection Plate
40 | 739 | Vac Degradation < 3x10-5 mbar
643 | 10488 | Ensure Gun Monitor is Running.
643 | 10488 | X Gun Tilt ≤50%
643 | 10488 | Y Gun Tilt ≤50%
643 | 10488 | X Gun Shift ≤50%
643 | 10488 | Y Gun Shift ≤50%
40 | 1562 | Run Clean Cycle
106 | 195 | Lindy Cables Supplied
40 | 739 | Record Diode Serial Number:
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal To Noise Test (Result >12%)
40 | 739 | Store Signal To Noise Result Sheet
40 | 739 | Diode Mirror Test
40 | 739 | Set Up Diode Protection Plate
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | Record Diode Serial Number:
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal To Noise Test (Result >12%)
40 | 739 | Store Signal To Noise Result (Screen-shot)
40 | 739 | Set Up Diode Protection Plate
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | Fit
643 | 10488 | Ensure Gun Monitor is Running.
643 | 10488 | Carry Out Initial Gun Run Up
643 | 10488 | X Gun Tilt ≤50%
643 | 10488 | Y Gun Tilt ≤50%
643 | 10488 | X Gun Shift ≤50%
643 | 10488 | Y Gun Shift ≤50%
643 | 10488 | Ensure Gun Monitor is Running and Test 96 Selected.
643 | 10488 | Carry Out Initial Gun Run Up
643 | 10488 | X Gun Tilt ≤50%
643 | 10488 | Y Gun Tilt ≤50%
643 | 10488 | X Gun Shift ≤50%
643 | 10488 | Y Gun Shift ≤50%
648 | 10553 | 30KV Res- ≤2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 1KV Res- ≤15nm
648 | 10553 | 200V Res- Not measured
648 | 10553 | Store Standard Resolution Images
40 | 1562 | Run Clean Cycle
663 | 10731 | EHT Serial Number
663 | 10731 | Torque Earth Connection to 3Nm
663 | 10731 | EHT Set Type
40 | 739 | Record Diode Serial Number:
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal To Noise Test (Result >12%)
40 | 739 | Store Signal To Noise Result (Screen-shot)
40 | 739 | Set Up Diode Protection Plate
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | Fit Detector Mechanics Cover
40 | 739 | Record Diode Serial Number:
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal To Noise Test (Result >12%)
40 | 739 | Store Signal To Noise Result (Screen-shot)
40 | 739 | Set Up Diode Protection Plate
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | Fit Detector Mechanics Cover
40 | 739 | Record Diode Serial Number:
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal To Noise Test (Result >12%)
40 | 739 | Store Signal To Noise Result (Screen-shot)
40 | 739 | Set Up Diode Protection Plate
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | Fit Detector Mechanics Cover
678 | 741 | Install Airlock FW Version 2.0.0
11 | 136 | A - B (< 1000Ω)
11 | 136 | C - D (< 10Ω)
11 | 136 | B - D (Open Circuit)
| 11371 | A - B (< 1000Ω)
| 11371 | C - D (< 10Ω)
| 11371 | B - D (Open Circuit)
12 | 163 | Tip S/N
12 | 163 | 2nd Engineer to inspect live tip prior to fitting gun head (Eng Initials)
663 | 10731 | EHT Serial Number
663 | 10731 | Torque Earth Connection to 3Nm
663 | 10731 | EHT Set Type Fitted
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
40 | 1562 | Run Clean Cycle
12 | 11208 | All cables Megger reads >1000Mohms at 1000V to ground
| 768 | Cable Megger reads >1000Mohms at 1000V to ground
37 | 11314 | Tablefiller successfully completed.
42 | 344 | Actual X-Axis
42 | 344 | Store X ripple shots and results
42 | 344 | Actual Y-Axis
42 | 344 | Store Y ripple shots and results
403 | 10464 | Disable Windows Fast Start UP
403 | 10464
403 | 10464 | Check That Power Fail is Set To Always On
403 | 10464 | Disable Windows Fast Start UP
28 | 700 | Confirm EM Server screeening programme is installed and running
28 | 700 | Confirm no repeated EM Server screening errors have been flagged
663 | 10731 | EHT Set Type Fitted
663 | 10731 | EHT Serial Number
663 | 10731 | Torque Earth Connection to 3Nm
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder to the Product Folder (O: drive)
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder to the Product Folder (O: drive)
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
646 | 11346
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder to the Product Folder (O: drive)
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
45 | 11059 | Copy- 1KV.Tif (Tiff customer image)
45 | 11059 | Copy- 15KV.Tif (Tiff customer image)
45 | 11059 | Copy- 30KV.Tif (Tiff customer image)
45 | 11059 | Copy- 200V.Tif (Tiff customer image)
45 | 11059 | Copy- 9876543.xml (XML Smartsem Licence)
45 | 11059 | Copy- Merlin Compact XX-XX S/O (Licence Rep File)
45 | 11059 | Copy- EmServer.accdb (Database)
45 | 11059 | Copy- EmServer.txt (Server Log File)
45 | 11059 | Copy- L500.ini (SmartSEM Config)
45 | 11059 | Copy- xx-xx.reg (Export of Registry File)
45 | 11059 | Copy- Mybackup.smartbackup (SmartSEM Backup)
45 | 11059 | Copy- Logbook (scanned)-PDF
45 | 11059 | Copy- DynamicEHT (mn2 logfile)
45 | 11059 | Copy- Probe Conditioning Stability (mn2 logfile)
45 | 11059 | Copy- 1KV Res.jpg (PATE Result)
45 | 11059 | Copy- 1KV Res.Tif (Tiff Res Image)
45 | 11059 | Copy- 15KV.jpg (PATE Results)
45 | 11059 | Copy- 15KV.Tif (Tiff Res Image)
45 | 11059 | Emission-Condenser OFF image from QA (Tiff Image)
45 | 11059 | Emission-Condenser ON image from QA (Tiff Image)
45 | 11059 | Merlin Compact User Manual (English) (PDF Manual)
45 | 11059 | Merlin Compact User Manual (German) (PDF Manual)
45 | 11059 | Merlin Compact Assembly (PDF Manual)
28 | 290 | Gun running to Datasheet
28 | 290 | Gun Vacuum acceptable (EHT on/off)
28 | 290 | Nominal Ext V acceptable
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448 , then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Run ASAP Screening Program
20 | 252 | System Specific Datazone Loaded
28 | 771 | Gun X
28 | 771 | Gun Y
28 | 771 | Aper X
28 | 771 | Aper Y
649 | 10555 | BSD Resolution ≤4nm.
649 | 10555 | Store BSD Resolution Image
646 | 10529 | BSD Diode Pre-test
646 | 10529 | Select Type
646 | 10529 | Select Amplifier Type
646 | 10529 | Black Level Calibration
646 | 10529 | Check Diode Segment Function
646 | 10529 | Check All Gain ranges
646 | 10529 | Check Auto range
646 | 10529 | Check Auto Level (Only functions on HDBSD)
646 | 10529 | Setup and Check TOPO
646 | 10529 | Take 3 images - Atomic, 5pA, 2KV
646 | 10529 | Store BSD Functional Images-Atomic, 5pA, 2KV
649 | 10555 | BSD Resolution ≤4nm.
649 | 10555 | Store BSD Resolution Image
389 | 783 | Check each individual segment in normal and inverse modes
389 | 783 | Analyze Segments
389 | 783 | Gain Ranges
389 | 783 | Pneumatics
389 | 783 | Alignment
389 | 783 | Vac Degradation < 3x10-5 mbar
389 | 783 | CZ Detector Cal
389 | 783 | Z Adjustment
28 | 291 | Mechanical Alignment ok
28 | 291 | Column clamping secure
28 | 409 | Gun X
28 | 409 | Gun Y
28 | 409 | Aper X
28 | 409 | Aper Y
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448 , then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Run ASAP Screening Program
20 | 252 | System Specific Datazone Loaded
40 | 1564 | Align Needle
40 | 1564 | Pneumatics ok
40 | 1564 | Pressure Test (CC Valve fully open)
40 | 1564 | Take image of paper
40 | 11085 | Black Level Calibration
40 | 11085 | Functional Check
40 | 11085 | Take Image
403 | 10464 | Check That AC Loss Policy Depend On, is Set To Power On
403 | 10464 | Disable Windows Fast Start UP
403 | 10464 | Check That AC Loss Policy Depend On, is Set To Power On
403 | 10464 | Disable Windows Fast Start UP
649 | 10555 | BSD Resolution ≤4nm.
649 | 10555 | Store BSD Resolution Image
389 | 783 | Check each individual segment in normal and inverse modes
389 | 783 | Analyze Segments
389 | 783 | Gain Ranges
389 | 783 | Pneumatics
389 | 783 | Alignment
389 | 783 | Vac Degradation < 3x10-5 mbar
389 | 783 | CZ Detector Cal
389 | 783 | Z Adjustment
40 | 11085 | Black Level Calibration
40 | 11085 | Functional Check
40 | 11085 | Take Image
285 | 3249 | Check That AC Loss Policy Depend On, is Set To Power On
285 | 3249 | Disable Windows Fast Start UP
646 | 10529 | BSD Diode Pre-test
646 | 10529 | Select Type
646 | 10529 | Select Amplifier Type
646 | 10529 | Black Level Calibration
646 | 10529 | Check Diode Segment Function
646 | 10529 | Check All Gain ranges
646 | 10529 | Check Auto range
646 | 10529 | Check Auto Level (Does Not Function With BSD1)
646 | 10529 | Setup and Check TOPO
646 | 10529 | Take 3 images - Atomic, 5pA, 2KV
646 | 10529 | Store BSD Functional Images-Atomic, 5pA, 2KV
389 | 783 | Check each individual segment in normal and inverse modes
389 | 783 | Analyze Segments
389 | 783 | Gain Ranges
389 | 783 | Pneumatics
389 | 783 | Alignment
389 | 783 | Vac Degradation < 3x10-5 mbar
389 | 783 | CZ Detector Cal
389 | 783 | Z Adjustment
37 | 11110 | 99v
37 | 11110 | 22kv
37 | 11110 | 30kv
37 | 11110 | 99v
37 | 11110 | 22kv
37 | 11110 | 30kv
37 | 329 | Stig: 30μm, All EHT’s is ≤ 20%
37 | 329 | Stig: Off Axis Aper, All EHT’s is ≤ 30%
37 | 329 | EHT <1kv: 30μm Aperture Gun/Aperture Aligns ≤ 60%
37 | 329 | EHT ≥1kv: 30μm Aperture Gun/Aperture Aligns ≤ 50%
37 | 11110 | 99v
37 | 11110 | 22kv
37 | 11110 | 30kv
37 | 329 | Stig: 30μm, All EHT’s is ≤ 20%
37 | 329 | Stig: Off Axis Aper, All EHT’s is ≤ 30%
37 | 329 | EHT <1kv: 30μm Aperture Gun/Aperture Aligns ≤ 60%
37 | 329 | EHT ≥1kv: 30μm Aperture Gun/Aperture Aligns ≤ 50%
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
389 | 783 | Check each individual segment in normal and inverse modes
389 | 783 | Analyze Segments
389 | 783 | Gain Ranges
389 | 783 | Pneumatics
389 | 783 | Alignment
389 | 783 | Vac Degradation < 3x10-5 mbar
389 | 783 | CZ Detector Cal
389 | 783 | Z Adjustment
37 | 329 | Stig: 30μm, All EHT’s is ≤ 20%
37 | 329 | Stig: Off Axis Aper, All EHT’s is ≤ 30%
37 | 329 | EHT <1kv: 30μm Aperture Gun/Aperture Aligns ≤ 60%
37 | 329 | EHT ≥1kv: 30μm Aperture Gun/Aperture Aligns ≤ 50%
40 | 785 | Collector1 Working
40 | 785 | Scintillator1 Working
40 | 785 | Collector2 Working
40 | 785 | Scintillator2 Working
| 594 | Open Circuit (Coil – GND).
| 592 | Ensure objective coil is concentric to pole piece ( this should be done as best as possible by eye as no jig yet available)
| 591 | Stand Off 1:
| 591 | Stand Off 2:
| 591 | Stand Off 3:
| 591 | Stand Off 4:
| 591 | Max Difference:
| 590 | Pressure 15 PSI for 5 minutes
| 589 | Resin is uniform and not proud
| 589 | No exposed wires
18 | 397 | Check SCS Folder/Sales Order
18 | 397 | SCS Number (If applicable)
389 | 783 | Setup
389 | 783 | Analyze Segments
389 | 783 | Gain Ranges
389 | 783 | Pneumatics
389 | 783 | Alignment
389 | 783 | Vac Degradation < 3x10-5 mbar
389 | 783 | CZ Detector Cal
389 | 783 | Z Adjustment
389 | 783 | Setup
389 | 783 | Calibration
389 | 783 | Fit beamsleeve funnel
389 | 783 | Mechanical alignment of beamsleeve
389 | 783 | set rise and fall height of beamsleeve
389 | 783 | Segmant test
389 | 783 | Diode mirror test
389 | 783 | Removing diode contamination with plasma cleaner
389 | 783 | Setup
389 | 783 | Calibration
389 | 783 | Fit beamsleeve funnel
389 | 783 | Mechanical alignment of beamsleeve
389 | 783 | set rise and fall height of beamsleeve
389 | 783 | Segment test
389 | 783 | Diode mirror test
389 | 783 | Removing diode contamination with plasma cleaner
27 | 430 | C1 Cal =
27 | 430 | C2 Cal =
27 | 430 | C3 Cal =
27 | 430 | Are Calibration factors typical? (see table of typical values in WI)
20 | 11021 | PC Serial Number:
20 | 11021 | Connect Up OEM Test Tool PC
20 | 11021 | Unistall SmartSEM
20 | 11021 | Install SmartSEM Version
20 | 11021 | Install Latest Service Pack
20 | 11021 | Install Gemini Support Tools-Version
20 | 11021 | Run ASAP Screening Program
20 | 11021 | System Specific Datazone Loaded
20 | 11021 | PC Serial Number:
20 | 11021 | Connect Up OEM Test Tool PC
20 | 11021 | Unistall SmartSEM
20 | 11021 | Install SmartSEM Version
20 | 11021 | Install Latest Service Pack
20 | 11021 | Install Gemini Support Tools-Version
20 | 11021 | Run ASAP Screening Program
20 | 11021 | System Specific Datazone Loaded
21 | 255 | Air (Pressure = 6 Bar)
21 | 255 | Water
21 | 255 | Mains
21 | 255 | Safety Earth
21 | 255 | N2
40 | 785 | Collector1 Working
40 | 785 | Scintillator1 Working
40 | 785 | Collector2 Working
40 | 785 | Scintillator2 Working
21 | 895 | Record Water Temperature (20ºC +/- 2ºC)
389 | 783 | Setup
389 | 783 | Calibration
389 | 783 | Fit beamsleeve funnel
389 | 783 | Mechanical alignment of beamsleeve
389 | 783 | set rise and fall height of beamsleeve
389 | 783 | Segment test
389 | 783 | Diode mirror test
389 | 783 | Removing diode contamination with plasma cleaner
21 | 256 | Store System checker full report
| 11391 | In-Lens and ESB
| 11391 | EO Dynamic PCB
| 11391 | EO Static PCB
| 11391 | Scan Gen PCB (ASAP or SVSEM)
| 11391 | Check that all column connections are securely fastened.
| 11391 | Ensure torque of gun earth is 20NM
| 588 | Position A:
| 588 | Position B:
| 588 | Position C:
| 588 | Position D:
| 588 | Max Difference:
| 587 | Visually inspected for damage/dents
| 587 | Sealing surfaces free from scratches/ imperfections
| 586 | Continuity ok
| 585 | Visual Inspection
| 585 | Liner Fitment
| 585 | End Cap Fitment
12 | 11354 | UHV Thermocouple resistance is 110Ὡ (+/-10Ὡ)
28 | 292 | Cut off ok
28 | 292 | No Charging particles
28 | 894 | Check pins are centered
28 | 894 | Dampers are concentric
28 | 894 | Rise and sink times are correct
21 | 257 | In Factory Tab: Omit Dongle Check
21 | 257 | In Factory Tab: Enter S/N
21 | 257 | In License Tab: Set All Licenses
21 | 187 | Z Move on Vent Enabled
21 | 258 | Confirm there are no Errors in EMserver (Apart from Dongle check and Stage Init)
21 | 258 | Confirm there are no Errors in EMserver (Apart from Dongle check and Stage Init)
21 | 385 | Select correct System type (HV / VP)
21 | 385 | If Pressure Stage is fitted Tick Box In Administrator
21 | 11026 | Carry out a 3hrs Plasma clean (Remove all samples from chamber!)
22 | 262 | Carry out overnight ‘Full service Bakeout’ and note actual duration.
22 | 262 | Confirm that all Thermocouples are reading.
22 | 262 | Confirm that all Ion pump heater elements and flange heaters are getting hot.
22 | 262 | Gun Vacuum post Bakeout. Ensure column has cooled prior to taking reading.
22 | 262 | After completion of Bakeout Torque the Gun head bolts to 25NM.
22 | 421 | Select "Tools and User Preferences": Set "Expert Gun Mode = Yes"
22 | 268 | Carry Out Gun Run UP
22 | 268 | Save Gun Monitor Log as Gun Run Up_TIP SN
23 | 269 | Run ESB conditioning MACRO prior to switching on EHT.
25 | 11401 | 30um aperture entered on central aperture
27 | 286 | Stage Angle =
718 | 11392 | Heater & Thermocouple cables plugged into AC unit
718 | 11392 | Column mounting bolts all secured
718 | 11392 | IGP bracket screws all secured
718 | 11392 | All column cabling cable tied neatly
30 | 297 | Check Mag cal errors are <1% in X & Y axes in all 3 Mag Ranges. Orthoganality ± 1deg.
30 | 297 | Store accuracy images (Mag 1, Mag 2, Mag 3 and Orthog)
42 | 344 | Actual X-Axis
42 | 344 | Store X ripple shots and results
42 | 344 | Actual Y-Axis
42 | 344 | Store Y ripple shots and results
50 | 480 | DC Tablefiller successfully completed.
| 582 | Objective Coil S/N :
| 582 | Scan Coil S/N :
| 582 | Pole Piece S/N :
12 | 11144 | Vacuum Check (-15PSI @ 5 Mins)
| 732 | Initials of 2nd Engineer confirming.
708 | 11236 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed.
708 | 504 | Select PC
708 | 504 | Hold aperture holder, look for holes on the sides.No particles are to be observed
28 | 293 | ESB concentricity in spec
28 | 293 | ESB-Inlens concenrticity in spec
21 | 895 | Record Water Temperature (20ºC +/- 2ºC)
39 | 339 | After conditioning confirm there are no EMServer error messages
39 | 339 | Store Gun monitor trace as 30kV Cycle Conditioning.mn2
33 | 316 | Actual for 1kV =
33 | 316 | Store 1KV Resolution Images and Pate image
33 | 316 | Actual for 15kV =
33 | 316 | Store 15KV Resolution Images and Pate image
35 | 319 | Aperture Finder successfully run
37 | 328 | Tablefiller successfully completed.
13 | 171 | Driver A – Liner (>1000MΩ)
13 | 171 | Driver B – Liner (>1000MΩ)
13 | 171 | Driver A – Driver B (>1000MΩ)
13 | 171 | Driver A – GROUND (>1000MΩ)
13 | 171 | Driver B – GROUND (>1000MΩ)
13 | 171 | Liner – GROUND (>1000MΩ)
12 | 11266 | Flange Heater Earths routed correctly (where applicable)
12 | 11266 | Thermocouple cables routed correctly
60 | 842 | Z Microswitch Position Set
60 | 842 | T Microswitch Position Set
60 | 842 | R Microswitch Position Set
22 | 264 | Tip No.
22 | 264 | Confirm that the Tip S/N on the Denka datasheet is the same as the Tip S/N on the column sticker.
28 | 700 | Confirm EM Server screeening programme is installed and running
28 | 700 | Confirm no repeated EM Server screening errors have been flagged
| 11375 | Position A:
| 11375 | Position B:
| 11375 | Position C:
| 11375 | Position D:
| 11375 | Max. Difference
| 11378 | Continuity (Close Circuit)
| 11378 | Coils to GND (Open Circuit)
43 | 346 | 1kV =
43 | 346 | Store 1KV Resolution Images and Pate image
43 | 346 | 15kV =
43 | 346 | Store 1KV Resolution Images and Pate image
43 | 347 | 1kV Inlens
43 | 347 | Store 1kV Inlens Customer Image
43 | 347 | 15kV Inlens
43 | 347 | Store 15kV Inlens Customer Image
43 | 11341 | 1kV ESB
43 | 11341 | Store 1kV ESB Customer Image
46 | 886 | Remove the sample holder from the system.
46 | 886 | Remove Test Tool SE2 and fit to warehouse stand. Fit SE2 blank plate to chamber and SHV connector to Spellman (00000-0302-367).
46 | 886 | Confirm that Raith top mounted SE2 detector is mounted to a warehouse stand and that a protective cap is fitted over the light guide.
46 | 886 | Set Stage position to Z=41mm & X & Y =65mm.
46 | 886 | Remove SCM and pack for shipment (include screws)
46 | 886 | Remove Docking Brackets and fit Transit clamps.
46 | 886 | Airfeet deflated and free to move.
46 | 886 | Backing lines re-configured so that PVSV valve is attached to flexi hose. Secure valve to frame.
46 | 886 | Confirm that the 3 'o' rings are fitted to the flexi hose backing line
46 | 886 | Select 'Partial vent on standby' in Smartsem.
46 | 888 | X-Ray Certificate ----------------------------------------------------------------------------------
46 | 888 | Quality Certificate ---------------------------------------------------------------------------------
46 | 888 | Quality Feedback Form
46 | 888 | DENKA Data Sheet
46 | 888 | Options Retained packing List
46 | 888 | Standard Packing List
46 | 888 | Declaration of Conformity
46 | 888 | Raith Installation Report
46 | 888 | Software Licence Agreement
28 | 294 | Run short image jump macro and analyse images
21 | 895 | Record Water Temperature (20ºC +/- 2ºC)
| 11379 | Continuity (Close Circuit)
| 11379 | Using Megger tester check there is not short circuit between lens and the condenser casing
| 11380 | S/N =
43 | 11341 | 1kV ESB
43 | 11341 | Store 1kV ESB Customer Image
44 | 348 | Chamber Door Interlock
44 | 348 | Service Switch
44 | 348 | Lens De-activation
44 | 348 | Detector De-activation
44 | 348 | EHT + IPPSU De-activation
644 | 10497 | Gun Chamber has been cleaned and inspected using UV light
644 | 10497 | EHT set socket well has been cleaned and inspected using UV light
644 | 10497 | Clean Firing Unit & Anode in Plasma Oven for 20 Minutes
644 | 10497 | Lab6 Firing Unit Fitted
644 | 10497 | Overnight Bakeout
46 | 11408 | Remove the sample holder from the system.
46 | 11408 | Remove Test Tool SE2 and fit to warehouse stand. Fit SE2 blank plate to chamber and SHV connector to Spellman (00000-0302-367).
46 | 11408 | Confirm that Raith top mounted SE2 detector is mounted to a warehouse stand and that a protective cap is fitted over the light guide.
46 | 11408 | Set Stage position to Z=41mm & X & Y =65mm.
46 | 11408 | Remove SCM and pack for shipment (include screws)
46 | 11408 | Remove Docking Brackets and fit Transit clamps.
46 | 11408 | Airfeet deflated and free to move.
46 | 11408 | Backing lines re-configured so that PVSV valve is attached to flexi hose. Secure valve to frame.
46 | 11408 | Confirm that the 3 'o' rings are fitted to the flexi hose backing line
46 | 11408 | Select 'Partial vent on standby' in Smartsem.
45 | 889 | Copy- 1KV.Tif (Tiff customer image)
45 | 889 | Copy- 15KV.Tif (Tiff customer image)
45 | 889 | Copy- 30KV.Tif (Tiff customer image)
45 | 889 | Copy- 200V.Tif (Tiff customer image)
45 | 889 | Copy- 9876543.xml (XML Smartsem Licence)
45 | 889 | Copy- XX-XX S/O (Licence Rep File)
45 | 889 | Copy- EmServer.accdb (Database)
45 | 889 | Copy- EmServer.txt (Server Log File)
45 | 889 | Copy- L500.ini (SmartSEM Config)
45 | 889 | Copy- xx-xx.reg (Export of Registry File)
45 | 889 | Copy- Mybackup.smartbackup (SmartSEM Backup)
45 | 889 | Copy- Logbook (scanned)-PDF
45 | 889 | Copy- DynamicEHT (mn2 logfile)
45 | 889 | Copy- Probe Conditioning Stability (mn2 logfile)
45 | 889 | Copy- 1KV Res.jpg (PATE Result)
45 | 889 | Copy- 1KV Res.Tif (Tiff Res Image)
45 | 889 | Copy- 15KV.jpg (PATE Results)
45 | 889 | Copy- 15KV.Tif (Tiff Res Image)
45 | 889 | Emission-Condenser OFF image from QA (Tiff Image)
45 | 889 | Emission-Condenser ON image from QA (Tiff Image)
28 | 10647 | Confirm SEM Status that (C1 Double only), C2 & C3 Cal and System Align Offset Angle are sensible
649 | 10555 | BSD Resolution ≤4nm.
649 | 10555 | Store BSD Resolution Image
21 | 895 | Record Water Temperature (20ºC +/- 2ºC)
55 | 776 | Check Carroed Out By:
83 | 701 | Confirm no repeated EM Server screening errors have been flagged
57 | 807 | Check Carried Out By:
59 | 828
59 | 826 | Select Stage Type & Travel
59 | 827 | Select Stage Type & Travel
59 | 828 | Select Stage Type & Travel
102 | 11054 | Function check
59 | 832 | Stage Kit Present
59 | 832 | Select Type Of Stage Kit Present
45 | 11024 | Em Server.db
45 | 11024 | Em Server.txt Log
45 | 11024 | Smart Backup.log
45 | 11024 | L500.ini
45 | 11024 | License file.xml
45 | 11024 | Registry file.reg
45 | 11024 | CalibDC.dat
45 | 11024 | Export Tablefiller
59 | 834 | Quality Sign Off Carried Out By:
21 | 895 | Record Water Flow Temperature (20ºC +/- 2ºC)
56 | 793 | Z Motor Fitted
56 | 793 | Z Motor Serial Number:
56 | 835 | T Motor Fitted
56 | 835 | T Motor Serial Number:
59 | 831 | Duration
45 | 11058 | Em Server.db
45 | 11058 | Em Server.log Text File
45 | 11058 | Optiprobe Calibration.Log Text File
45 | 11058 | Mybackup.Smartbackup
45 | 11058 | Image Folder
45 | 11058 | L500.ini
45 | 11058 | Dynamic EHT.MN2
45 | 11058 | Probe Stability.MN2
45 | 11058 | License file.xml
45 | 11058 | Registry file.reg
45 | 11058 | System Checker Full Report
45 | 11058 | Export Alignment Data From Table Editor(.MERLINCDB)
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Chamber secured to vibration plate
46 | 887 | Chamber Sercuring Screws Tightened to 25NM
46 | 887 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 887 | Check that the Stage bearing blocks screws are secure.
46 | 887 | Remove test tool stage
46 | 887 | Remove test tool stage PCB
46 | 887 | Secure Airfeet with foam blocks.
46 | 887 | Secure Turbo pump with foam blocks.
46 | 887 | Fit Covers to all unused connectors.
46 | 887 | Remove Cekon Plug From Mains Cable
46 | 887 | Fit Live and Neutral Labels to Mains Cable
46 | 887 | Mains and Earth cable stowed on Mains unit.
46 | 887 | Fit shipping Eyes to transit blocks.
46 | 887 | Put thread lock on pendulum screws and transit block Screws.
46 | 887 | All doors closed and secured.
46 | 887 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 887 | Print shipping stickers.
45 | 11059 | Copy- 1KV.Tif (Tiff customer image)
45 | 11059 | Copy- 15KV.Tif (Tiff customer image)
45 | 11059 | Copy- 30KV.Tif (Tiff customer image)
45 | 11059 | Copy- 200V.Tif (Tiff customer image)
45 | 11059 | Copy- 9876543.xml (XML Smartsem Licence)
45 | 11059 | Copy- Merlin Compact XX-XX S/O (Licence Rep File)
45 | 11059 | Copy- EmServer.accdb (Database)
45 | 11059 | Copy- EmServer.txt (Server Log File)
45 | 11059 | Copy- L500.ini (SmartSEM Config)
45 | 11059 | Copy- xx-xx.reg (Export of Registry File)
45 | 11059 | Copy- Mybackup.smartbackup (SmartSEM Backup)
45 | 11059 | Copy- Logbook (scanned)-PDF
45 | 11059 | Copy- DynamicEHT (mn2 logfile)
45 | 11059 | Copy- Probe Conditioning Stability (mn2 logfile)
45 | 11059 | Copy- 1KV Res.jpg (PATE Result)
45 | 11059 | Copy- 1KV Res.Tif (Tiff Res Image)
45 | 11059 | Copy- 15KV.jpg (PATE Results)
45 | 11059 | Copy- 15KV.Tif (Tiff Res Image)
45 | 11059 | Emission-Condenser OFF image from QA (Tiff Image)
45 | 11059 | Emission-Condenser ON image from QA (Tiff Image)
45 | 11059 | Merlin Compact User Manual (English) (PDF Manual)
45 | 11059 | Merlin Compact User Manual (German) (PDF Manual)
45 | 11059 | Merlin Compact Assembly (PDF Manual)
649 | 10555 | BSD Resolution ≤4nm.
649 | 10555 | Store BSD Resolution Image
45 | 11059 | Copy- 1KV.Tif (Tiff customer image)
45 | 11059 | Copy- 15KV.Tif (Tiff customer image)
45 | 11059 | Copy- 30KV.Tif (Tiff customer image)
45 | 11059 | Copy- 200V.Tif (Tiff customer image)
45 | 11059 | Copy- 9876543.xml (XML Smartsem Licence)
45 | 11059 | Copy- Merlin Compact XX-XX S/O (Licence Rep File)
45 | 11059 | Copy- EmServer.accdb (Database)
45 | 11059 | Copy- EmServer.txt (Server Log File)
45 | 11059 | Copy- L500.ini (SmartSEM Config)
45 | 11059 | Copy- xx-xx.reg (Export of Registry File)
45 | 11059 | Copy- Mybackup.smartbackup (SmartSEM Backup)
45 | 11059 | Copy- Logbook (scanned)-PDF
45 | 11059 | Copy- DynamicEHT (mn2 logfile)
45 | 11059 | Copy- Probe Conditioning Stability (mn2 logfile)
45 | 11059 | Copy- 1KV Res.jpg (PATE Result)
45 | 11059 | Copy- 1KV Res.Tif (Tiff Res Image)
45 | 11059 | Copy- 15KV.jpg (PATE Results)
45 | 11059 | Copy- 15KV.Tif (Tiff Res Image)
45 | 11059 | Emission-Condenser OFF image from QA (Tiff Image)
45 | 11059 | Emission-Condenser ON image from QA (Tiff Image)
45 | 11059 | Merlin Compact User Manual (English) (PDF Manual)
45 | 11059 | Merlin Compact User Manual (German) (PDF Manual)
45 | 11059 | Merlin Compact Assembly (PDF Manual)
45 | 368 | Analytic Fil I
45 | 368 | Analytic Extractor V:
55 | 776 | Checked Carried Our By:
60 | 842 | Z Microswitch Position Set
60 | 842 | T Microswitch Position Set
60 | 842 | R Microswitch Position Set
| 607 | pressure 15 PSI for 5 minutes
644 | 10498 | Gun Monitor Setup
644 | 10498 | Test No. 96 Selected
644 | 10498 | Run up to 32KV and leave for 15 minute
644 | 10498 | Shutdown and run up to 32KV, Repeat 5 times
644 | 10498 | Run up to 32KV and leave for 30 minutes
644 | 10498 | 32KV Gun Monitor Traced Saved
644 | 10498 | Gun Monitor Setup
644 | 10498 | Test No. 96 Selected
644 | 10498 | Run up to 32KV and leave for 15 minute
644 | 10498 | Shutdown and run up to 32KV, Repeat 5 times
644 | 10498 | Run up to 32KV and leave for 30 minutes (There must be no spikes)
644 | 10498 | 32KV Gun Monitor Traced Saved
56 | 836 | R Motor Fitted
56 | 836 | R Motor Serail Number
45 | 367 | High Res Fil I:
45 | 367 | High Res Extractor V:
46 | 11060 | X-Ray Certificate ----------------------------------------------------------------------------------
46 | 11060 | Quality Certificate ---------------------------------------------------------------------------------
46 | 11060 | Quality Feedback Form
46 | 11060 | DENKA Data Sheet
46 | 11060 | Options Retained packing List
46 | 11060 | Standard Packing List
46 | 11060 | Declaration of Conformity
46 | 11060 | Raith Installation Report
46 | 11060 | Software Licence Agreement
707 | 11283 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11283 | Multi_TS_HRQA - Passed inspection criteria, Image Saved
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
21 | 895 | Record Water Flow Temperature (20ºC +/- 4ºC)
2 | 5 | Record the temp sensor probe resistance
2 | 896 | Record the temp sensor probe resistance
2 | 6 | Set regulator pressure indicator to 0.03MPa
2 | 6 | Set regulator pressure to 0.03MPa
55 | 776 | Check Carried Out By:
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
707 | 11285 | Measure aperture size, 30um (+/- 2um)
707 | 11285 | Multi_TS_30um - Passed inspection criteria, Image Saved
2 | 7 | Record IPPSU S/N
2 | 7 | Set DIP switches
2 | 7 | Qualification Marks Checked
390 | 900 | Carry out safety tests using the Seaward Prime 100 Tester.
390 | 900 | Fill out completely the Product Safety Tests form in Section 13.
60 | 845 | Checked Carried Our By:
60 | 843 | Z Motor SeriaL Number:
60 | 843 | Z Motor SeriaL Number:
60 | 843 | T Motor Serial Number:
56 | 795 | Check Carried Out By:
707 | 11160 | Measure aperture size, 120um (+/- 2um)
707 | 11160 | Multi_TS_120um - Passed inspection criteria, Image Saved
2 | 20 | Install Compressed Air Line
2 | 20 | Set foot height
55 | 776 | Checked Carried Our By:
60 | 843 | Z Motor SeriaL Number:
60 | 843 | T Motor Serial Number:
60 | 843 | R Motor Serial Number:
56 | 792 | Set Micro-Switch Position .
707 | 11290 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11290 | Multi_US_HRQA - Passed inspection criteria, Image Saved
707 | 11319 | Anode_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11319 | Anode_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11327 | Anode_TS_Bore1 - Passed inspection criteria, Image Saved
56 | 792 | Set Each Micro-Switch Position .
707 | 11318 | Anode_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11328 | Anode_TS_Bore1 - Passed inspection criteria, Image Saved
2 | 21 | Set Valve Screw torque IN (1 turn)
2 | 21 | Set Valve Screw torque OUT (1 turn)
2 | 21 | Set foot height
2 | 21 | Install Compressed Air Line
2 | 21 | Install Control Cable
390 | 900 | Carry out safety tests using the Seaward Prime 100 Tester.
390 | 900 | Fill out completely the Product Safety Tests form in Section 13.
42 | 344 | Actual X-Axis
42 | 344 | Store X ripple shots and results
42 | 344 | Actual Y-Axis
42 | 344 | Store Y ripple shots and results
30 | 906 | Store Images in QA folder
38 | 11315 | Measure I Probe with beam blanked - Spec <=2pA
38 | 11315 | Difference between EHT on and off < 1pA.
707 | 11330 | Anode_TS_Bore2 - Passed inspection criteria, Image Saved
707 | 11331 | Anode_TS_Bore2 - Passed inspection criteria, Image Saved
707 | 11165 | Measure aperture size, 40um (+/- 2um)
707 | 11165 | Anode_TS_40um - Passed inspection criteria, Image Saved
707 | 11166 | Measure aperture size, 55um (+/- 2um)
707 | 11166 | Anode_TS_55um - Passed inspection criteria, Image Saved
707 | 11334 | Anode_US_HRQA - Passed inspection criteria, Image Saved
707 | 11333 | Anode_US_HRQA - Passed inspection criteria, Image Saved
2 | 22 | Set Valve Screw torque IN (.5 turn)
2 | 22 | Set Valve Screw torque OUT(1 turn)
2 | 22 | Set foot height
2 | 22 | Install Compressed Air Line
2 | 22 | Install Control Cable
390 | 900 | Record Prime Test 100 SN)
390 | 900 | Record Prime Test 100 Calibration Date)
30 | 907 | XB Tilt Test carried out
59 | 833 | Enter Stage Serial Number
707 | 758 | SF_TS_HRQA - Passed inspection criteria, Image Saved
707 | 11173 | Measure aperture size, 35um (+/- 2um)
5 | 11275 | Verify the ASAP cables are connected to the customer port configuration
5 | 11275 | Verify the ASAP cables are connected to the customer port configuration
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
646 | 10545 | Correct PC - 349535-9001-020 Installed
646 | 10545 | Print Out SmartPI Check List (Link in Documents)
646 | 10545 | Install SmartSEM Licences
646 | 10545 | Install MIL Redistribution
646 | 10545 | Fit MIL Dongle
646 | 10545 | SmartPI Software Version
646 | 10545 | Install SmartPI Software
646 | 10545 | I/O-SCAN Card Fitted (Bruker Only)
646 | 10545 | I/O-SCAN Card Calibrated
646 | 10545 | Setup & Run Application
646 | 10545 | Check Results & Image Montage
646 | 10545 | Store Compressed Database,Montage Image, Log files etc
2 | 23 | Set Valve Screw torque IN (.5 turn)
2 | 23 | Set Valve Screw torque OUT(1 turn)
2 | 23 | Set foot height
2 | 23 | Install Compressed Air Line
2 | 23 | Install Control Cable
285 | 708 | Record MAC address (Internet not MIC-NET)
24 | 273 | Level Frame
24 | 273 | Check pendulum height (with spacer)
24 | 273 | Set up ‘float’ height
24 | 273 | Check inlet / outlet valves
24 | 273 | Set up docking pins
24 | 273 | Level Frame
24 | 273 | Check pendulum height (with spacer)
24 | 273 | Set up ‘float’ height
24 | 273 | Check inlet / outlet valves
24 | 273 | Set up docking pins
40 | 902 | Dowel pins have been mounted correctly to the mounting plate
644 | 10497 | System Vacuum is better than 3.0 x 10-6
644 | 10497 | Gun Chamber has been cleaned and inspected using UV light
644 | 10497 | EHT set socket well has been cleaned and inspected using UV light
644 | 10497 | Clean Firing Unit & Anode in Plasma Oven for 20 Minutes
644 | 10497 | Lab6 Firing Unit Fitted
644 | 10497 | Overnight Bakeout
2 | 896 | Record the temp sensor probe resistance
285 | 708 | Record MAC address (Internet not MIC-NET)
4 | 11389 | Fit chamber pumping pipe
4 | 11389 | Fit chamber to damper
4 | 11389 | Fit chamber bolts
4 | 11389 | Fit earthquake bolts
4 | 10937 | Screws fitted by
4 | 39 | Turbo Pump S/No.
4 | 39 | Qualification Marks Checked
20 | 425 | Record MAC address (Internet not MIC-NET)
4 | 37 | Shim Stage
4 | 37 | Set Mag Clamp
4 | 37 | Fit Gun Interlock Key to Stage Door
30 | 296 | Calibration wizard completed.
63 | 871 | Check Set Point
63 | 870 | Check Set Point
20 | 11038 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
20 | 11038 | PC S/N
20 | 11038 | If SEM Enhanced PC(349535-9001-020) has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Hard disk Image (Version)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
79 | 1585 | Check wiring is neat and tidy and properly secured.
79 | 1585 | Check cable ties have been cut back sufficiently and no sharp edges remain.
80 | 1570 | Reading (≤2.00e-06) =
651 | 915 | Check that the 200V image ia at a good quality.
651 | 915 | Check that the 1KV Resolution Image meets specification and of a good quality.
403 | 890 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
20 | 425 | Record MAC address (Internet not MIC-NET)
285 | 708 | Record MAC address (Internet not MIC-NET)
399 | 892 | Remove MAC address (Internet)
399 | 892 | Remove MAC address (MIC-NET)
648 | 10553 | 30KV Res- ≤2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 1KV Res- ≤15nm
648 | 10553 | 200V Res- Not measured
648 | 10553 | 1KV Oscillation Check-EHT=1KV, Scan Speed=2, WD=10mmm, Mag=10Kx
648 | 10553 | Store Standard Resolution Images + Oscillation Check
648 | 10553 | 30KV Res- ≤2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 1KV Res- ≤15nm
648 | 10553 | 200V Res- Not measured
648 | 10553 | 1KV Oscillation Check - EHT=1KV, Scan Speed=2, WD=10mmm, Mag=10Kx
648 | 10553 | Store Standard Resolution Images + Oscillation Check
646 | 10654 | Print Out Mineralogic Check List (Link in Documents)
646 | 10654 | Correct PC - Installed
646 | 10654 | Install SmartSEM Licences
646 | 10654 | Instal MIL Redistribution
646 | 10654 | Fit MIL Dongle
646 | 10654 | Install Mineralogic Software
646 | 10654 | Calibrate I/O-SCAN card
646 | 10654 | Setup & Run Application
646 | 10654 | Check Results & Image Montage
646 | 10654 | 24 Hour Run
646 | 10654 | Store Compressed Database,Montage Image, Log files etc
403 | 6935 | Record MAC address (Internet not MIC-NET)
654 | 893 | Remove MAC address (Internet)
654 | 893 | Remove MAC address (MIC-NET)
403 | 6935 | Record MAC address (Internet not MIC-NET)
648 | 10553 | 30KV Res- ≤2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 1KV Res- ≤15nm
648 | 10553 | 200V Res- Not measured
648 | 10553 | 1KV Oscillation Check - EHT=1KV, Scan Speed=2, WD=10mmm, Mag=10Kx
648 | 10553 | Store Standard Resolution Images + Oscillation Check
646 | 10654 | Print Out Mineralogic Check List (Link in Documents)
646 | 10654 | Correct PC - Installed
646 | 10654 | Install SmartSEM Licences
646 | 10654 | Instal MIL Redistribution
646 | 10654 | Fit MIL Dongle
646 | 10654 | Install Mineralogic Software
646 | 10654 | Calibrate I/O-SCAN card
646 | 10654 | Setup & Run Application
646 | 10654 | Check Results & Image Montage
646 | 10654 | 24 Hour Run
646 | 10654 | Store Compressed Database,Montage Image, Log files etc
646 | 10654 | Print Out Mineralogic Check List (Link in Documents)
646 | 10654 | Correct PC - Installed
646 | 10654 | Install SmartSEM Licences
646 | 10654 | Instal MIL Redistribution
646 | 10654 | Fit MIL Dongle
646 | 10654 | Install Mineralogic Software
646 | 10654 | Calibrate I/O-SCAN card
646 | 10654 | Setup & Run 20 min Rapid MAC STD Measure
646 | 10654 | Check Results & Image Montage - 20 min Rapid MAC STD Measure
646 | 10654 | Run 90 min High Mag Rapid MAC STD Measure
646 | 10654 | 10 Hour Overnight Run
646 | 10654 | Store Compressed Database,Montage Image, Log files etc
5 | 50 | Fit Camera
63 | 873 | Check Set Point
45 | 891 | Remove MAC address (Internet)
45 | 891 | Remove MAC address (MIC-NET)
399 | 892 | Remove MAC address (Internet)
399 | 892 | Remove MAC address (MIC-NET)
45 | 891 | Remove MAC address (Internet)
45 | 891 | Remove MAC address (MIC-NET)
646 | 10654 | Print Out Mineralogic Check List (Link in Documents)
646 | 10654 | Correct PC - Installed
646 | 10654 | Install SmartSEM Licences
646 | 10654 | Instal MIL Redistribution
646 | 10654 | Fit MIL Dongle
646 | 10654 | Install Mineralogic Software
646 | 10654 | Calibrate I/O-SCAN card
646 | 10654 | Setup & Run 20 min Rapid MAC STD Measure
646 | 10654 | Check Results & Image Montage - 20 min Rapid MAC STD Measure
646 | 10654 | Run 90 min High Mag Rapid MAC STD Measure
646 | 10654 | Check Results and Compare with 20 min Run
646 | 10654 | 10 Hour Overnight Run
646 | 10654 | Check Results and Compare with 20 and 90 min Run
646 | 10654 | Store Compressed Database,Montage Image, Log files etc
646 | 10654 | Print Out Mineralogic Check List (Link in Documents)
646 | 10654 | Correct PC - Installed
646 | 10654 | Install SmartSEM Licences
646 | 10654 | Instal MIL Redistribution
646 | 10654 | Fit MIL Dongle
646 | 10654 | Install Mineralogic Software
646 | 10654 | Calibrate I/O-SCAN card
646 | 10654 | Setup & Run 20 min Rapid MAC STD Measure
646 | 10654 | Check Results & Image Montage - 20 min Rapid MAC STD Measure
646 | 10654 | Run 90 min High Mag Rapid MAC STD Measure
646 | 10654 | Check Results and Compare with 20 min Run
646 | 10654 | 10 Hour Overnight Run
646 | 10654 | Check Results and Compare with 20 and 90 min Run
646 | 10654 | Store Compressed Database,Montage Image, Log files etc
646 | 10654 | Print Out Mineralogic Check List (Link in Documents)
646 | 10654 | Correct PC - Installed
646 | 10654 | Install SmartSEM Licences
646 | 10654 | Instal MIL Redistribution
646 | 10654 | Fit MIL Dongle
646 | 10654 | Install Mineralogic Software
646 | 10654 | Install MAC Standard Recipe Files
646 | 10654 | Calibrate I/O-SCAN card
646 | 10654 | MAC Standard & Stage Postion Set
646 | 10654 | Adjust Recipe for the System
646 | 10654 | Adjust SEM Macro
646 | 10654 | Run 20 min Rapid MAC STD Measure
646 | 10654 | Check Results & Image Montage - 20 min Rapid MAC STD Measure
646 | 10654 | Run 90 min High Mag Rapid MAC STD Measure
646 | 10654 | Check Results and Compare with 20 min Run
646 | 10654 | 10 Hour Overnight Run
646 | 10654 | Check Results and Compare with 20 and 90 min Run
646 | 10654 | Store Compressed Database,Montage Image, Log files etc
46 | 886 | Remove the sample holder from the system.
46 | 886 | If fitted remove Test Tool SE2 and fit to warehouse stand. Fit SE2 blank plate to chamber and SHV connector to Spellman (00000-0302-367).
46 | 886 | Remove Docking Brackets and fit Transit clamps.
46 | 886 | Airfeet deflated and free to move.
63 | 874 | Check Set point (-10.8º)
45 | 11058 | Em Server.db
45 | 11058 | Em Server.log Text File
45 | 11058 | Optiprobe Calibration.Log Text File
45 | 11058 | Mybackup.Smartbackup
45 | 11058 | Image Folder
45 | 11058 | L500.ini
45 | 11058 | Dynamic EHT.MN2
45 | 11058 | Probe Stability.MN2
45 | 11058 | License file.xml
45 | 11058 | Registry file.reg
45 | 11058 | System Checker Full Report
45 | 11058 | Export Alignment Data From Table Editor(.MERLINCDB)
651 | 915 | Check that the 200V image ia at a good quality.
651 | 915 | Check 1KV Resolution Image is in specification and of a good quality.
646 | 10654 | Print Out Mineralogic Check List (Link in Documents)
646 | 10654 | Correct PC - Installed
646 | 10654 | Install SmartSEM Licences
646 | 10654 | Instal MIL Redistribution
646 | 10654 | Fit MIL Dongle
646 | 10654 | Install Mineralogic Software
646 | 10654 | Install MAC Standard Recipe Files
646 | 10654 | Calibrate I/O-SCAN card
646 | 10654 | MAC Standard & Stage Postion Set
646 | 10654 | Adjust Recipe for the System
646 | 10654 | Run 20 min Rapid MAC STD Measure
646 | 10654 | Check Results & Image Montage - 20 min Rapid MAC STD Measure
646 | 10654 | Run 90 min High Mag Rapid MAC STD Measure
646 | 10654 | Check Results and Compare with 20 min Run
646 | 10654 | 10 Hour Overnight Run
646 | 10654 | Check Results and Compare with 20 and 90 min Run
646 | 10654 | Store Compressed Database,Montage Image, Log files etc
389 | 10874 | Print Out Mineralogic Check List (Link in Documents)
389 | 10874 | Correct PC - Installed
389 | 10874 | Install SmartSEM Licences
389 | 10874 | Instal MIL Redistribution
389 | 10874 | Fit MIL Dongle
389 | 10874 | Install Mineralogic Software
389 | 10874 | Install MAC Standard Recipe Files
389 | 10874 | Calibrate I/O-SCAN card
389 | 10874 | MAC Standard & Stage Postion Set
389 | 10874 | Adjust Recipe for the System
389 | 10874 | Run 20 min Rapid MAC STD Measure
389 | 10874 | Check Results & Image Montage - 20 min Rapid MAC STD Measure
389 | 10874 | Run 90 min High Mag Rapid MAC STD Measure
389 | 10874 | Check Results and Compare with 20 min Run
389 | 10874 | 10 Hour Overnight Run
389 | 10874 | Check Results and Compare with 20 and 90 min Run
389 | 10874 | Setup & Run Application
389 | 10874 | Check Results & Image Montage
389 | 10874 | 24 Hour Run
389 | 10874 | Print Out Mineralogic Check List (Link in Documents)
389 | 10874 | Correct PC - Installed
389 | 10874 | Install SmartSEM Licences
389 | 10874 | Instal MIL Redistribution
389 | 10874 | Fit MIL Dongle
389 | 10874 | Install Mineralogic Software
389 | 10874 | Install MAC Standard Recipe Files
389 | 10874 | Calibrate I/O-SCAN card
389 | 10874 | MAC Standard & Stage Postion Set
389 | 10874 | Adjust Recipe for the System
389 | 10874 | Run 20 min Rapid MAC STD Measure
389 | 10874 | Check Results & Image Montage - 20 min Rapid MAC STD Measure
389 | 10874 | Run 90 min High Mag Rapid MAC STD Measure
389 | 10874 | Check Results and Compare with 20 min Run
389 | 10874 | 10 Hour Overnight Run
389 | 10874 | Check Results and Compare with 20 and 90 min Run
61 | 860 | Checked Carried Our By:
63 | 875 | Check Stage Initialises Correctly
63 | 876 | Select Stage Type
63 | 876 | Select Stage Type
63 | 876 | Select Stage Type and Travel
63 | 877 | Select Stage Type and Travel
63 | 877 | Select Stage Type and Travel
63 | 878 | Select Stage Type and Travel
63 | 878 | Check Stage Type and Travel
63 | 876 | Select Stage Type and Check Travel
61 | 860 | Checked Carried Our By:
63 | 875 | Check Stage Initialises Correctly
63 | 876 | Select Stage Type
63 | 876 | Select Stage Type and Travel
63 | 878 | Select Stage Type and Travel
63 | 877 | Check Stage Type and Travel
63 | 877 | Select Stage Type and Check Travel
18 | 911 | Objective Coil S/N :
18 | 911 | Scan Coil S/N :
18 | 911 | Pole Piece S/N :
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Remove and pack all power and system cables (not stage cables)
46 | 887 | Remove and pack all Fibre optic cables (not stage cables)
46 | 887 | Remove and pack all cooling pipes (not stage)
46 | 887 | Remove PCB frame complete
46 | 887 | Remove AC unit complete with mains cable
46 | 887 | Remove PSU unit
46 | 887 | Remove
46 | 887 | Chamber secured to vibration plate
46 | 887 | Chamber Sercuring Screws Tightened to 25NM
46 | 887 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 887 | Mains and Earth cable stowed on Mains unit.
46 | 887 | Print shipping stickers.
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Remove and pack all power and system cables (not stage cables)
46 | 887 | Remove and pack all Fibre optic cables (not stage cables)
46 | 887 | Remove and pack all cooling pipes (not stage)
46 | 887 | Remove PCB frame complete
46 | 887 | Remove AC unit complete with mains cable
46 | 887 | Remove Festo valve cluster
46 | 887 | Remove PSU unit
46 | 887 | Remove EHT Set
46 | 887 | Remove IGP PSU
46 | 887 | Remove column / support frame complete and load into transit box
46 | 887 | Chamber secured to vibration plate
46 | 887 | Chamber Securing Screws Tightened to 25NM
46 | 887 | Print shipping stickers.
45 | 362 | Em Server.db
45 | 362 | Em Server.txt Log
45 | 362 | Optiprobe.txt Log
45 | 362 | Smart Backup.log
45 | 362 | L500.ini
45 | 362 | License file.xml
45 | 362 | Registry file.reg
45 | 362 | CalibSc.dat
102 | 1626 | Function of Internal Specimen Current Monitor checked.
102 | 1626 | Confirm when Beam Blanked measured Specimen Current <= 2pa
102 | 1625 | Function of CCD cameras/ focus.
102 | 1625 | Confirm that camera toggle function works.
63 | 876 | Check Stage Type and Travel
63 | 878 | Select Stage Type and Check Travel
63 | 879 | Cycle Test with 0.5Kg Weight for 10 Mins
63 | 881 | Run Axis Cycle Test
63 | 881 | Run Axis Cycle Test
63 | 882 | Stage Kit
63 | 882 | 27mm Spacer Block
63 | 882 | 54mm Spacer Block
63 | 882 | 81mm Spacer Block - Fitted
63 | 882 | Stage Kit
63 | 882 | 27mm Spacer Block
63 | 882 | 54mm Spacer Block
63 | 882 | 81mm Spacer Block - Fitted
63 | 903 | Stage Kit
63 | 903 | 27mm Spacer Block Fitted
63 | 884 | Quality Sign Off Carried Out By:
63 | 905 | Select Stage Type
63 | 903 | Stage Kit
63 | 903 | 27mm Spacer Block Fitted
45 | 11024 | Em Server.db
45 | 11024 | Em Server.txt Log
45 | 11024 | Smart Backup.log
45 | 11024 | L500.ini
45 | 11024 | License file.xml
45 | 11024 | Registry file.reg
45 | 11024 | CalibDC.dat
45 | 11024 | Export Tablefiller
45 | 364 | Copy all sub directories and files from the Image directory
644 | 10497 | System Vacuum & Gun Vacuum is within Spec
644 | 10497 | Gun Chamber has been cleaned and inspected using UV light
644 | 10497 | EHT set socket well has been cleaned and inspected using UV light
644 | 10497 | Clean Firing Unit & Anode in Plasma Oven for 20 Minutes
644 | 10497 | Lab6 Firing Unit Fitted
644 | 10497 | Overnight Bakeout
651 | 915 | Check that the 200V image is of a good quality.
651 | 915 | Check 1KV Resolution Image is in specification and of a good quality.
651 | 915 | Check 3KV Resolution Image is in specification and of a good quality.
6 | 70 | Camera Cables Routing outside Plinth
6 | 70 | Obtained an image on the camera, ensure no flickering / distortions occur.
63 | 879 | Cycle Test with 0.5Kg Weight for 10 Mins
63 | 903 | Stage Kit
63 | 903 | 27mm Spacer Block Fitted
63 | 905 | Select Stage Type
63 | 905 | Stage Serial Number:
63 | 905 | Select Stage Type
63 | 905 | Stage Serial Number:
45 | 11058 | Em Server.db
45 | 11058 | Em Server.log Text File
45 | 11058 | Optiprobe Calibration.Log Text File
45 | 11058 | Mybackup.Smartbackup
45 | 11058 | Image Folder
45 | 11058 | L500.ini
45 | 11058 | Dynamic EHT.MN2
45 | 11058 | Probe Stability.MN2
45 | 11058 | License file.xml
45 | 11058 | Registry file.reg
45 | 11058 | System Checker Full Report
45 | 11058 | Export Alignment Data From Table Editor(.MERLINCDB)
46 | 888 | X-Ray Certificate ----------------------------------------------------------------------------------
46 | 888 | Quality Certificate ---------------------------------------------------------------------------------
46 | 888 | Quality Feedback Form
46 | 888 | DENKA Data Sheet
46 | 888 | Standard Packing List
46 | 888 | Software Licence Agreement
45 | 362 | EM Server.db
45 | 362 | Em Server.Log
45 | 362 | Optiprobe Calibration.Log
45 | 362 | MyBackup.smartbackup
45 | 362 | L500.ini
45 | 362 | License file.xml
45 | 362 | Registry file.reg
45 | 362 | CalibSC.dat
45 | 11024 | EM Server.db
45 | 11024 | Em Server.Log
45 | 11024 | MyBackup.smartbackup
45 | 11024 | L500.ini
45 | 11024 | License file.xml
45 | 11024 | Registry file.reg
45 | 11024 | CalibDC.dat
45 | 11024 | Export Tablefiller
45 | 889 | Copy- WO.xml (XML Smartsem Licence)
45 | 889 | Copy- S/N W/O (Licence Rep File)
45 | 889 | Copy- EmServer.db (Database)
45 | 889 | Copy- EmServer.log (Server Log File)
45 | 889 | Copy- L500.ini (SmartSEM Config)
45 | 889 | Copy- Registery S/N.reg (Export of Registry File)
45 | 889 | Copy- Mybackup.smartbackup (SmartSEM Backup)
46 | 886 | Remove the sample holder from the system.
46 | 886 | If fitted remove Test Tool SE2 and fit to warehouse stand. Fit SE2 blank plate to chamber and SHV connector to Spellman (00000-0302-367).
46 | 886 | Remove Docking Brackets and fit Transit clamps.
46 | 886 | Airfeet deflated and free to move.
46 | 886 | Vent the system as column and penning will need to removed.
651 | 915 | Check that the 200V image is of a good quality.
651 | 915 | Check 1KV Resolution Image meets specification and of a good quality.
651 | 915 | Check 3KV Resolution Image meets specification and of a good quality.
651 | 915 | Check 30KV Resolution Image meets specification and of a good quality.
63 | 879 | Check Tilt Travel -10° to +90°
63 | 903 | Stage Kit
63 | 903 | 27mm Spacer Block Fitted
63 | 905 | Select Stage Type
63 | 905 | Stage Serial Number:
63 | 905 | Select Stage Type
63 | 905 | Stage Serial Number:
45 | 889 | Copy- WO.xml (XML Smartsem Licence)
45 | 889 | Copy- S/N W/O (Licence Rep File)
45 | 889 | Copy- EmServer.db (Database)
45 | 889 | Copy- EmServer.log (Server Log File)
45 | 889 | Copy- L500.ini (SmartSEM Config)
45 | 889 | Copy- Registery S/N.reg (Export of Registry File)
45 | 889 | Copy- Mybackup.smartbackup (SmartSEM Backup)
45 | 889 | Copy - Optiprobe Calibration.Log
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Remove and pack all power and system cables (not stage cables)
46 | 887 | Remove and pack all Fibre optic cables (not stage cables)
46 | 887 | Remove and pack all cooling pipes (not stage)
46 | 887 | Remove PCB frame complete
46 | 887 | Remove AC unit complete with mains cable
46 | 887 | Remove Festo valve cluster
46 | 887 | Remove PSU unit
46 | 887 | Remove EHT Set
46 | 887 | Remove IGP PSU
46 | 887 | Remove EBBU Electonics from column and fit shorting plug
46 | 887 | Remove Water flow and Temperature sensor assembly
46 | 887 | Remove column / support frame complete and load into transit box
46 | 887 | Chamber secured to vibration plate
46 | 887 | Chamber Securing Screws Tightened to 25NM
46 | 887 | Print shipping stickers.
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Remove and pack all power and system cables (not stage cables)
46 | 887 | Remove and pack all Fibre optic cables (not stage cables)
46 | 887 | Remove and pack all cooling pipes (not stage)
46 | 887 | Remove PCB frame complete
46 | 887 | Remove AC unit complete with mains cable
46 | 887 | Remove Festo valve cluster
46 | 887 | Remove PSU unit
46 | 887 | Remove EHT Set
46 | 887 | Remove IGP PSU
46 | 887 | Remove EBBU Electonics from column and fit shorting plug
46 | 887 | Remove Water flow and Temperature sensor assembly
46 | 887 | Remove column / support frame complete and attach to column storage frame
46 | 887 | Chamber secured to vibration plate
46 | 887 | Chamber Securing Screws Tightened to 25NM
46 | 887 | Print shipping stickers.
651 | 915 | Check that the 200V image is of a good quality.
651 | 915 | Check 1KV Resolution Image meets specification and of a good quality.
651 | 915 | Check 3KV Resolution Image meets specification and of a good quality.
651 | 915 | Check 30KV Resolution Image meets specification and of a good quality.
6 | 60 | Validate temperature probe knurled nut seal has no leak
6 | 60 | No Pressure loss in 2mins
63 | 880 | Stage Axis Cycle Test with 0.5Kg Weight for 10 Mins
46 | 888 | X-Ray Certificate ----------------------------------------------------------------------------------
46 | 888 | Quality Certificate ---------------------------------------------------------------------------------
46 | 888 | DENKA Data Sheet
46 | 888 | Standard Packing List
106 | 195 | Lindy Cables Supplied
24 | 273 | Level Frame
24 | 273 | Check pendulum height (with spacer)
24 | 273 | Set up ‘float’ height
24 | 273 | Check inlet / outlet valves
24 | 273 | Set up docking pins
24 | 273 | Level Frame
24 | 273 | Check pendulum height (with spacer)
24 | 273 | Set up ‘float’ height
24 | 273 | Check inlet / outlet valves
24 | 273 | Set up docking pins
46 | 1568 | Remove docking brackets and fit transit brackets
46 | 1568 | Power system down.
46 | 1568 | Blow water out of the cooling system.
46 | 1568 | Put thread lock on pendulum screws and transit block Screws.
46 | 1568 | Chamber secured to vibration plate
46 | 1568 | Chamber Sercuring Screws Tightened to 25NM
46 | 1568 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 1568 | Secure Airfeet with foam blocks.
46 | 1568 | Fit foam blocks around column cladding.
46 | 1568 | Secure Turbo pump with foam blocks.
46 | 1568 | Fit Covers to all unused connectors.
46 | 1568 | Remove Cekon Plug From Mains Cable
46 | 1568 | Fit Live and Neutral Labels to Mains Cable
46 | 1568 | Mains and Earth cable stowed on Mains unit.
46 | 1568 | Fit shipping Eyes to transit blocks.
46 | 1568 | Confirm all column cladding is secure and screws Loctited in place.
46 | 1568 | All doors closed and secured.
46 | 1568 | All plinth panels removed and stored in transit trolley.
46 | 1568 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 1568 | Fit M8 screw to secure stage door to chamber.
46 | 1568 | Fit dust covers and secure / remove cables.
46 | 1568 | Prepare column cladding transport jig .
46 | 1568 | Print shipping stickers .
389 | 114 | Black Level Cal
389 | 114 | C2D Bias Target Check
389 | 114 | Gain Range Check
389 | 114 | Auto Level Check
389 | 114 | Auto Range Check
389 | 114 | Function Image
389 | 114 | Store C2D Functional Image
63 | 880 | Stage Axis Cycle Test with 0.5Kg Weight for 10 Mins
63 | 903 | Stage Kit
63 | 903 | 27mm Spacer Block Fitted
46 | 887 | Remove docking brackets and fit transit brackets
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Remove and pack all power and system cables (not stage cables)
46 | 887 | Remove and pack all Fibre optic cables (not stage cables)
46 | 887 | Remove and pack all cooling pipes (not stage)
46 | 887 | Remove PCB frame complete
46 | 887 | Remove AC unit complete with mains cable
46 | 887 | Remove Festo valve cluster
46 | 887 | Remove PSU unit
46 | 887 | Remove EHT Set
46 | 887 | Remove IGP PSU
46 | 887 | Remove EBBU Electonics from column and fit shorting plug
46 | 887 | Remove Water flow and Temperature sensor assembly
46 | 887 | Remove column / support frame complete and attach to column storage frame
46 | 887 | Chamber secured to vibration plate
46 | 887 | Chamber Securing Screws Tightened to 25NM
46 | 887 | Print shipping stickers.
46 | 887 | Remove docking brackets and fit transit brackets
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Remove and pack all power and system cables (not stage cables)
46 | 887 | Remove and pack all Fibre optic cables (not stage cables)
46 | 887 | Remove and pack all cooling pipes (not stage)
46 | 887 | Remove PCB frame complete
46 | 887 | Remove AC unit complete with mains cable
46 | 887 | Remove Cekon Plug From Mains Cable
46 | 887 | Remove Festo valve cluster
46 | 887 | Remove PSU unit
46 | 887 | Remove EHT Set
46 | 887 | Remove IGP PSU
46 | 887 | Remove EBBU Electonics from column and fit shorting plug
46 | 887 | Remove Water flow and Temperature sensor assembly
46 | 887 | Remove column / support frame complete and attach to column storage frame
46 | 887 | Chamber secured to vibration plate
46 | 887 | Chamber Securing Screws Tightened to 25NM
46 | 887 | Print shipping stickers.
650 | 10567 | Install Customer Licences
650 | 10567 | Fit SmartSEM Dongle
650 | 10567
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer.db to the Product Folder
650 | 10567 | Store EM Server.db
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
650 | 10567 | Add Licence to the SmartSEM USB Stick
650 | 10567 | Add EM Server.db to the SmartSEM USB Stick
389 | 114 | Black Level Cal
389 | 114 | C2D Bias Target Check
389 | 114 | Gain Range Check
389 | 114 | Auto Level Check
389 | 114 | Auto Range Check
389 | 114 | Function Image
389 | 114 | Store C2D Functional Image
46 | 887 | Remove docking brackets and fit transit brackets
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Remove and pack all power and system cables (not stage cables)
46 | 887 | Remove and pack all Fibre optic cables (not stage cables)
46 | 887 | Remove and pack all cooling pipes (not stage)
46 | 887 | Remove PCB frame complete
46 | 887 | Remove AC unit complete with mains cable
46 | 887 | Remove Festo valve cluster
46 | 887 | Remove PSU unit
46 | 887 | Remove EHT Set
46 | 887 | Remove IGP PSU
46 | 887 | Remove EBBU Electonics from column and fit shorting plug
46 | 887 | Remove Water flow and Temperature sensor assembly
46 | 887 | Remove column / support frame complete and attach to column storage frame
46 | 887 | Chamber secured to vibration plate
46 | 887 | Chamber Securing Screws Tightened to 25NM
46 | 887 | Print shipping stickers.
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder to the Product Folder (O: drive)
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder to the Product Folder (O: drive)
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
389 | 115 | Setup Ethernet Connection
389 | 115 | Install Aztec Software
389 | 115 | Aztec Software Version
389 | 115 | Licence Installation
389 | 115 | Run Aztec Assist+
389 | 115 | Install Detector
389 | 115 | Imaging Setup
389 | 115 | Stage Calibration Setup
389 | 115 | Calibrate Detector
389 | 115 | Save Aztec Sign Off PDF and Calibration Folder
12 | 11354 | UHV Thermocouple resistance is 110Ὡ (+/-10Ὡ)
63 | 904 | Stage Kit
46 | 887 | Remove docking brackets and fit transit brackets
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Remove and pack all power and system cables (not stage cables)
46 | 887 | Remove and pack all Fibre optic cables (not stage cables)
46 | 887 | Remove and pack all cooling pipes (not stage)
46 | 887 | Remove PCB frame complete
46 | 887 | Remove AC unit complete with mains cable
46 | 887 | Remove Festo valve cluster
46 | 887 | Remove PSU unit
46 | 887 | Remove EHT Set
46 | 887 | Remove IGP PSU
46 | 887 | Remove EBBU Electonics from column and fit shorting plug
46 | 887 | Remove Water flow and Temperature sensor assembly
46 | 887 | Remove column / support frame complete and attach to column storage frame
46 | 887 | Chamber secured to vibration plate
46 | 887 | Chamber Securing Screws Tightened to 25NM
46 | 887 | Print shipping stickers.
646 | 10653 | Number of Detectors
646 | 10653 | Detector Serial Number
646 | 10653 | Detector Resolution
646 | 10653 | Detector Area
646 | 10653 | Detector Installed
646 | 10653 | In/Out Stops Set
646 | 10653 | Isolation Check of Detector & Column >2MΩ
646 | 10653 | Install Remoting & Setup Network Connection
646 | 10653 | Bruker Esprit Software Version
646 | 10653 | Install Bruker Esprit Software
646 | 10653 | Bruker EDS Manager Version
646 | 10653 | Install EDS Manager
646 | 10653 | Install Bruker Licences
646 | 10653 | Setup Pulse Processor
646 | 10653 | Check Standard Timing Cards Fitted
646 | 10653 | List Optional Timing Cards Fitted
646 | 10653 | Install Detector Data
646 | 10653 | Check Microscope driver
646 | 10653 | Setup SEM Conditions
646 | 10653 | Determine Analytical Working Distance
646 | 10653 | Calibrate Detector
646 | 10653 | Calibrate Image in Espirit
389 | 10873 | Number of Detectors
389 | 10873 | Detector Serial Number(s)
389 | 10873 | Detector Resolution
389 | 10873 | Detector Area
389 | 10873 | Detector Installed
389 | 10873 | In/Out Stops Set
389 | 10873 | Isolation Check of Detector & Column >2MΩ
389 | 10873 | Install Remoting & Setup Network Connection
389 | 10873 | Bruker Esprit Software Version
389 | 10873 | Install Bruker Esprit Software
389 | 10873 | Bruker EDS Manager Version
389 | 10873 | Install EDS Manager
389 | 10873 | Install Bruker Licences
389 | 10873 | Setup Pulse Processor
389 | 10873 | Check Standard Timing Cards Fitted
389 | 10873 | List Optional Timing Cards Fitted
389 | 10873 | Install Detector Data
389 | 10873 | Check Microscope driver
389 | 10873 | Setup SEM Conditions
389 | 10873 | Determine Analytical Working Distance
389 | 10873 | Calibrate Detector
389 | 10873 | Calibrate Image in Espirit
46 | 887 | Remove docking brackets and fit transit brackets
46 | 887 | Vent the Chamber.
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Remove and pack all power and system cables (not stage cables)
46 | 887 | Remove and pack all Fibre optic cables (not stage cables)
46 | 887 | Remove and pack all cooling pipes (not stage)
46 | 887 | Remove PCB frame complete
46 | 887 | Remove AC unit complete with mains cable
46 | 887 | Remove Cekon Plug From Mains Cable
46 | 887 | Remove Festo valve cluster
46 | 887 | Remove PSU unit
46 | 887 | Remove EHT Set
46 | 887 | Remove IGP PSU
46 | 887 | Remove EBBU Electonics from column and fit shorting plug
46 | 887 | Remove Water flow and Temperature sensor assembly
46 | 887 | Remove column / support frame complete and attach to column storage frame
46 | 887 | Chamber secured to vibration plate
46 | 887 | Chamber Securing Screws Tightened to 25NM
46 | 887 | Print shipping stickers.
46 | 11123 | Remove docking brackets and fit transit brackets
46 | 11123 | Power system down.
46 | 11123 | Blow water out of the cooling system.
46 | 11123 | Chamber secured to vibration plate (N/A to XB 550L)
46 | 11123 | Chamber securing screws tightened to 25nm (N/A to XB 550L)
46 | 11123 | Chamber securing screws tightness checked. Type your initials (For XB 550L type N/A)
46 | 11123 | Secure Airfeet with foam blocks.
46 | 11123 | Secure Turbo pump with foam blocks.
46 | 11123 | Fit Covers to all unused connectors.
46 | 11123 | Mains and Earth cable stowed on Mains unit.
46 | 11123 | Fit shipping Eyes to transit blocks.
46 | 11123 | Put thread lock on pendulum screws and transit block Screws.
46 | 11123 | All doors closed and secured.
46 | 11123 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11123 | Print shipping stickers.
46 | 11123 | Fit 8x M6 x 55mm cap head screws to stage door
644 | 10497 | System Vacuum & Gun Vacuum is within Spec
644 | 10497 | Gun Chamber has been cleaned and inspected using UV light
644 | 10497 | EHT set socket well has been cleaned and inspected using UV light
644 | 10497 | Clean Firing Unit & Anode in Plasma Oven for 20 Minutes
644 | 10497 | Lab6 Firing Unit Fitted
644 | 10497 | Overnight Bakeout
650 | 10567 | Install Customer Licences
650 | 10567 | Fit SmartSEM Dongle
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer.db to the Product Folder
650 | 10567 | Store EM Server.db
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
650 | 10567 | Add Licence to the SmartSEM USB Stick
650 | 10567 | Add EM Server.db to the SmartSEM USB Stick
650 | 10567 | Install Customer Licences
650 | 10567 | Fit SmartSEM Dongle
650 | 10567 | Deselect Omit Dongle Check in Administrator, Factory tab
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer.db to the Product Folder
650 | 10567 | Store EM Server.db
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
650 | 10567 | Add Licence to the SmartSEM USB Stick
650 | 10567 | Add EM Server.db to the SmartSEM USB Stick
44 | 923
46 | 11407 | Remove docking brackets and fit transit brackets
46 | 11407 | Power system down.
46 | 11407 | Blow water out of the cooling system.
46 | 11407 | Chamber secured to vibration plate
46 | 11407 | Chamber Sercuring Screws Tightened to 25NM
46 | 11407 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 11407 | Check that the Stage bearing blocks screws are secure.
46 | 11407 | Remove test tool stage
46 | 11407 | Remove test tool stage PCB
46 | 11407 | Secure Airfeet with foam blocks.
46 | 11407 | Secure Turbo pump with foam blocks.
46 | 11407 | Fit Covers to all unused connectors.
46 | 11407 | Remove Cekon Plug From Mains Cable
46 | 11407 | Fit Live and Neutral Labels to Mains Cable
46 | 11407 | Mains and Earth cable stowed on Mains unit.
46 | 11407 | Fit shipping Eyes to transit blocks.
46 | 11407 | Put thread lock on pendulum screws and transit block Screws.
46 | 11407 | All doors closed and secured.
46 | 11407 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11407 | Print shipping stickers.
650 | 10567 | Install Customer Licences
650 | 10567 | Fit SmartSEM Dongle
650 | 10567 | Deselect Omit Dongle Check in Administrator, Factory tab
650 | 10567 | In Administrator, Service check that the Database and Dongle serial numbers match.
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer.db to the Product Folder
650 | 10567 | Store EM Server.db
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
650 | 10567 | Add Licence to the SmartSEM USB Stick
650 | 10567 | Add EM Server.db to the SmartSEM USB Stick
44 | 356 | Check System And Dongle S/No. Match In Service Panel
44 | 356 | Un-tick Omit Dongle In Factory Panel
44 | 356 | Check Dongle fits in USB HUB (“ICY BOX”) and replace if needed.
390 | 3322 | Check System And Dongle S/No. Match In Service Panel
390 | 3322 | Un-tick Omit Dongle In Factory Panel
390 | 3322 | Check Dongle fits in USB HUB (“ICY BOX”) and replace if needed.
389 | 11074 | Check each individual segment in normal and inverse modes
389 | 11074 | Analyze Segments
389 | 11074 | Gain Ranges
389 | 11074 | Pneumatics
389 | 11074 | Alignment
389 | 11074 | Vac Degradation < 3x10-5 mbar
389 | 11074 | CZ Detector Cal
389 | 11074 | Z Adjustment
20 | 425 | Record MAC address (Internet not MIC-NET)
| 614 | Enter pole shoe S/N =
62 | 868 | Checked Carried Our By:
63 | 884
46 | 11407 | Remove docking brackets and fit transit brackets
46 | 11407 | Power system down.
46 | 11407 | Blow water out of the cooling system.
46 | 11407 | Chamber secured to vibration plate
46 | 11407 | Chamber Sercuring Screws Tightened to 25NM
46 | 11407 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 11407 | Check that the Stage bearing blocks screws are secure.
46 | 11407 | Remove test tool stage
46 | 11407 | Remove test tool stage PCB
46 | 11407 | Secure Airfeet with foam blocks.
46 | 11407 | Secure Turbo pump with foam blocks.
46 | 11407 | Fit Covers to all unused connectors.
46 | 11407 | Remove Cekon Plug From Mains Cable
46 | 11407 | Fit Live and Neutral Labels to Mains Cable
46 | 11407 | Mains and Earth cable stowed on Mains unit.
46 | 11407 | Fit shipping Eyes to transit blocks.
46 | 11407 | Put thread lock on pendulum screws and transit block Screws.
46 | 11407 | All doors closed and secured.
46 | 11407 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11407 | Print shipping stickers.
650 | 914 | Check Dongle fits in USB HUB (“ICY BOX”) and replace if needed.
650 | 914 | Check Dongle fits in USB HUB (“ICY BOX”) and replace if needed.
46 | 887 | Remove docking brackets and fit transit brackets
46 | 887 | Vent the Chamber.
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Remove and pack all power and system cables (not stage cables)
46 | 887 | Remove and pack all Fibre optic cables (not stage cables)
46 | 887 | Remove and pack all cooling pipes (not stage)
46 | 887 | Remove PCB frame complete
46 | 887 | Remove AC unit complete with mains cable
46 | 887 | Remove Cekon Plug From Mains Cable
46 | 887 | Remove Festo valve cluster
46 | 887 | Remove PSU unit
46 | 887 | Remove EHT Set
46 | 887 | Remove IGP PSU
46 | 887 | Remove EBBU Electonics from column and fit shorting plug
46 | 887 | Remove Water flow and Temperature sensor assembly
46 | 887 | Remove column / support frame complete and attach to column storage frame
46 | 887 | Chamber secured to vibration plate
46 | 887 | Chamber Securing Screws Tightened to 25NM
46 | 887 | Print shipping stickers.
46 | 887 | Remove docking brackets and fit transit brackets
46 | 887 | Vent the Chamber.
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Remove On / Off switch panel
46 | 887 | Remove and pack all power and system cables (not stage cables)
46 | 887 | Remove and pack all Fibre optic cables (not stage cables)
46 | 887 | Remove and pack all cooling pipes (not stage)
46 | 887 | Remove PCB frame complete
46 | 887 | Remove AC unit complete with mains cable
46 | 887 | Remove Cekon Plug From Mains Cable
46 | 887 | Remove Festo valve cluster
46 | 887 | Remove PSU unit
46 | 887 | Remove EHT Set
46 | 887 | Remove IGP PSU
46 | 887 | Remove EBBU Electonics from column and fit shorting plug
46 | 887 | Remove Water flow and Temperature sensor assembly
46 | 887 | Remove column / support frame complete and attach to column storage frame
46 | 887 | Chamber secured to vibration plate
46 | 887 | Chamber Securing Screws Tightened to 25NM
46 | 887 | Print shipping stickers.
46 | 887 | Remove docking brackets and fit transit brackets
46 | 887 | Vent the Chamber.
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Remove On / Off switch panel
46 | 887 | Print and fit DJEL S/N Label
46 | 887 | Remove and pack all power and system cables (not stage cables)
46 | 887 | Remove and pack all Fibre optic cables (not stage cables)
46 | 887 | Remove and pack all cooling pipes (not stage)
46 | 887 | Remove PCB frame complete
46 | 887 | Remove AC unit complete with mains cable
46 | 887 | Remove Cekon Plug From Mains Cable
46 | 887 | Remove Festo valve cluster
46 | 887 | Remove PSU unit
46 | 887 | Remove EHT Set
46 | 887 | Remove IGP PSU
46 | 887 | Remove EBBU Electonics from column and fit shorting plug
46 | 887 | Remove Water flow and Temperature sensor assembly
46 | 887 | Remove column / support frame complete and attach to column storage frame
46 | 887 | Chamber secured to vibration plate
46 | 887 | Chamber Securing Screws Tightened to 25NM
46 | 887 | Print shipping stickers.
46 | 11407 | Remove docking brackets and fit transit brackets
46 | 11407 | Power system down.
46 | 11407 | Blow water out of the cooling system.
46 | 11407 | Chamber secured to vibration plate
46 | 11407 | Chamber Sercuring Screws Tightened to 25NM
46 | 11407 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 11407 | Check that the Stage bearing blocks screws are secure.
46 | 11407 | Remove test tool stage
46 | 11407 | Remove test tool stage PCB
46 | 11407 | Secure Airfeet with foam blocks.
46 | 11407 | Secure Turbo pump with foam blocks.
46 | 11407 | Fit Covers to all unused connectors.
46 | 11407 | Remove Cekon Plug From Mains Cable
46 | 11407 | Fit Live and Neutral Labels to Mains Cable
46 | 11407 | Mains and Earth cable stowed on Mains unit.
46 | 11407 | Fit shipping Eyes to transit blocks.
46 | 11407 | Put thread lock on pendulum screws and transit block Screws.
46 | 11407 | All doors closed and secured.
46 | 11407 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11407 | Print shipping stickers.
40 | 916 | Unpack Volutome package and check it's completeness
40 | 916 | Remove transit clamps from Volutome stage
40 | 916 | Ensure hardware has been configured according to work instructions
40 | 916 | Fit Volutome stage
40 | 916 | Connect Volutome stage and controller to the SEM
40 | 916 | Install Volutome software installation package
40 | 916 | Configure Volutome stage motors
40 | 916 | Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator
40 | 916 | Set up the Volutome workbench in Zen
40 | 916 | Record SenseBSD serial number
40 | 916 | Record Volutome stage serial number
40 | 916 | Record SenseBSD amplifier serial number
40 | 916 | Record SenseBSD mechanics serial number
40 | 916 | Record C-885 PI controller serial number
40 | 916 | Record E-727 PI controller serial number
40 | 916 | Function check the Volutome stage light
40 | 916 | Adjust Volutome stage centre (Crosshairs should be within 4 central squares)
40 | 916 | Install Nav cam in-line connections (If applicable)
40 | 916 | Perform the microtome and stage functionality checks
40 | 916 | Fit and align the SenseBSD diode
40 | 916 | Check SenseBSD alignment repeatability
40 | 916 | Check all gain ranges of the SenseBSD (normal/inverted)
40 | 916 | Perform Signal/noise ratio check
40 | 916 | Ensure BSD diode protector is fitted and adjusted in chamber
40 | 916 | Install cable tie to loop BSD signal cable
40 | 916 | Check SCM with Volutome stage
40 | 916 | Ripple X Spec 2nm (600Kx Mag 15KV WD2)
40 | 916 | Ripple Y Spec 2nm (600Kx Mag 15KV WD2)
40 | 916 | Resolution 1KV Spec same as standard resolution
40 | 916 | Resolution 15KV Spec same as standard resolution
40 | 916 | Contamination
40 | 916 | Perform X-ray check
40 | 916 | Pump Down Time
40 | 916 | Ultimate Vacuum
40 | 916 | Focal CC fitted to Volutome stage
40 | 916 | Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar
40 | 916 | Generate live image from Axiocam
40 | 916 | Re-package Volutome and check it's completeness
40 | 916 | ManEng/ Senior Eng/ Production Manger Sign off
45 | 891 | Remove MAC address (Internet)
45 | 891 | Remove MAC address (MIC-NET)
12 | 157 | Postion A
12 | 157 | Postion B
12 | 157 | Postion C
12 | 157 | Postion D
12 | 157 | Max. Difference:
12 | 157 | Postion A
12 | 157 | Postion B
12 | 157 | Postion C
12 | 157 | Postion D
12 | 157 | Max. Difference:
| 515 | Continuity ok
709 | 925 | Backing Pump Type
709 | 925 | Serial Number
40 | 916 | Unpack Volutome package and check it's completeness
40 | 916 | Remove transit clamps from Volutome stage
40 | 916 | Ensure hardware has been configured according to work instructions
40 | 916 | Fit Volutome stage
40 | 916 | Connect Volutome stage and controller to the SEM
40 | 916 | Install Volutome software installation package
40 | 916 | Configure Volutome stage motors
40 | 916 | Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator
40 | 916 | Set up the Volutome workbench in Zen
40 | 916 | Record SenseBSD serial number
40 | 916 | Record Volutome stage serial number
40 | 916 | Record SenseBSD amplifier serial number
40 | 916 | Record SenseBSD mechanics serial number
40 | 916 | Record C-885 PI controller serial number
40 | 916 | Record E-727 PI controller serial number
40 | 916 | Function check the Volutome stage light
40 | 916 | Adjust Volutome stage centre (Crosshairs should be within 4 central squares)
40 | 916 | Install Nav cam in-line connections (If applicable)
40 | 916 | Perform the microtome and stage functionality checks
40 | 916 | Fit and align the SenseBSD diode
40 | 916 | Check SenseBSD alignment repeatability
40 | 916 | Check all gain ranges of the SenseBSD (normal/inverted)
40 | 916 | Perform Signal/noise ratio check
40 | 916 | Ensure BSD diode protector is fitted and adjusted in chamber
40 | 916 | Install cable tie to loop BSD signal cable
40 | 916 | Check SCM with Volutome stage
40 | 916 | Ripple X Spec 2nm (600Kx Mag 15KV WD2)
40 | 916 | Ripple Y Spec 2nm (600Kx Mag 15KV WD2)
40 | 916 | Resolution 1KV Spec same as standard resolution
40 | 916 | Resolution 15KV Spec same as standard resolution
40 | 916 | Contamination
40 | 916 | Perform X-ray check
40 | 916 | Pump Down Time
40 | 916 | Ultimate Vacuum
40 | 916 | Focal CC fitted to Volutome stage
40 | 916 | Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar
40 | 916 | Generate live image from Axiocam
40 | 916 | Re-package Volutome and check it's completeness
40 | 916 | ManEng/ Senior Eng/ Production Manger Sign off
| 515 | Continuity ok
46 | 887 | Remove docking brackets and fit transit brackets
46 | 887 | Vent the Chamber.
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Remove On / Off switch panel
46 | 887 | Print and fit DJEL S/N Label
46 | 887 | Remove and pack all power and system cables (not stage cables)
46 | 887 | Remove and pack all Fibre optic cables (not stage cables)
46 | 887 | Remove and pack all cooling pipes and cooling flanges.
46 | 887 | Remove PCB frame complete
46 | 887 | Remove AC unit complete with mains cable
46 | 887 | Remove Cekon Plug From Mains Cable
46 | 887 | Remove Festo valve cluster
46 | 887 | Remove PSU unit
46 | 887 | Remove EHT Set
46 | 887 | Remove IGP PSU
46 | 887 | Remove EBBU Electonics from column and fit shorting plug
46 | 887 | Remove Water flow and Temperature sensor assembly
46 | 887 | Remove column / support frame complete and attach to column storage frame
46 | 887 | Chamber secured to vibration plate - Only required if test tool is being shipped to Bollore
46 | 887 | Chamber Securing Screws Tightened to 25NM - Only required if test tool is being shipped to Bollore
46 | 887 | Secure the pendulum feet using foam blocks and strap - Only required if test tool is being shipped to Bollore
46 | 887 | Print shipping stickers.
709 | 925 | Backing Pump Type
709 | 925 | Serial Number
709 | 925 | Backing Pump Type
4 | 11389 | Fit chamber pumping pipe
4 | 11389 | Fit chamber to damper
4 | 11389 | Fit chamber bolts
4 | 11389 | Fit earthquake bolts
4 | 11389 | Record Chamber SN
4 | 11389 | Fit chamber pumping pipe
4 | 11389 | Fit chamber to damper
4 | 11389 | Fit chamber bolts
4 | 11389 | Fit earthquake bolts
4 | 11389 | Record Chamber SN
30 | 906 | Store Images in QA folder
40 | 916 | Unpack Volutome package and check it's completeness
40 | 916 | Remove transit clamps from Volutome stage
40 | 916 | Ensure hardware has been configured according to work instructions
40 | 916 | Fit Volutome stage
40 | 916 | Connect Volutome stage and controller to the SEM
40 | 916 | Install Volutome software installation package
40 | 916 | Configure Volutome stage motors
40 | 916 | Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator
40 | 916 | Set up the Volutome workbench in Zen
40 | 916 | Record VolumeBSD serial number
40 | 916 | Record Volutome stage serial number
40 | 916 | Record VolumeBSD amplifier serial number
40 | 916 | Record VolumeBSD mechanics serial number
40 | 916 | Record C-885 PI controller serial number
40 | 916 | Record E-727 PI controller serial number
40 | 916 | Function check the Volutome stage light
40 | 916 | Adjust Volutome stage centre (Crosshairs should be within 4 central squares)
40 | 916 | Install Nav cam in-line connections (If applicable)
40 | 916 | Perform the microtome and stage functionality checks
40 | 916 | Fit and align the VolumeBSD diode
40 | 916 | Check VolumeBSD alignment repeatability
40 | 916 | Check all gain ranges of the VolumeBSD (normal/inverted)
40 | 916 | Perform Signal/noise ratio check
40 | 916 | Ensure BSD diode protector is fitted and adjusted in chamber
40 | 916 | Install cable tie to loop BSD signal cable
40 | 916 | Check SCM with Volutome stage
40 | 916 | Ripple X Spec 2nm (600Kx Mag 15KV WD2)
40 | 916 | Ripple Y Spec 2nm (600Kx Mag 15KV WD2)
40 | 916 | Resolution 1KV Spec same as standard resolution
40 | 916 | Resolution 15KV Spec same as standard resolution
40 | 916 | Contamination
40 | 916 | Perform X-ray check
40 | 916 | Pump Down Time
40 | 916 | Ultimate Vacuum
40 | 916 | Focal CC fitted to Volutome stage
40 | 916 | Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar
40 | 916 | Generate live image from Axiocam
40 | 916 | Re-package Volutome and check it's completeness
40 | 916 | ManEng/ Senior Eng/ Production Manger Sign off
46 | 886 | Remove the sample holder from the system.
46 | 886 | If fitted remove the test tool SE2 and fit to warehouse stand and fit blank to chamber.
46 | 886 | Remove Docking Brackets and fit Transit clamps.
46 | 886 | Airfeet deflated and free to move.
46 | 886 | Vent the system as column and penning will need to removed.
46 | 887 | Remove docking brackets and fit transit brackets
46 | 887 | Vent the Chamber.
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Remove On / Off switch panel
46 | 887 | Print and fit DJEL S/N Label
46 | 887 | Remove and pack all power and system cables (not stage cables)
46 | 887 | Remove and pack all Fibre optic cables (not stage cables)
46 | 887 | Remove and pack all cooling pipes and cooling flanges.
46 | 887 | Remove PCB frame complete
46 | 887 | Remove AC unit complete with mains cable
46 | 887 | Remove Cekon Plug From Mains Cable
46 | 887 | Remove Festo valve cluster
46 | 887 | Remove PSU unit
46 | 887 | Remove EHT Set
46 | 887 | Remove IGP PSU
46 | 887 | Remove EBBU Electonics from column and fit shorting plug
46 | 887 | Remove Water flow and Temperature sensor assembly
46 | 887 | Remove column / support frame complete and attach to column storage frame
46 | 887 | Chamber secured to vibration plate
46 | 887 | Chamber Securing Screws Tightened to 25NM
46 | 887 | Print shipping stickers.
43 | 347 | 1kV Inlens
43 | 347 | Store 1kV Inlens Customer Image
43 | 347 | 15kV Inlens
43 | 347 | Store 15kV Inlens Customer Image
43 | 11341 | 1kV ESB
43 | 11341 | Store 1kV ESB Customer Image
709 | 925 | Ensure Chiller Type is 'AIR' cooled.
46 | 11407 | Remove docking brackets and fit transit brackets
46 | 11407 | Power system down.
46 | 11407 | Blow water out of the cooling system.
46 | 11407 | Chamber secured to vibration plate
46 | 11407 | Chamber Sercuring Screws Tightened to 25NM
46 | 11407 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 11407 | Check that the Stage bearing blocks screws are secure.
46 | 11407 | Remove test tool stage
46 | 11407 | Remove test tool stage PCB
46 | 11407 | Secure Airfeet with foam blocks.
46 | 11407 | Secure Turbo pump with foam blocks.
46 | 11407 | Confirm there is no cooling fluid laying in the floor of the plinth
46 | 11407 | Fit Covers to all unused connectors.
46 | 11407 | Remove Cekon Plug From Mains Cable
46 | 11407 | Fit Live and Neutral Labels to Mains Cable
46 | 11407 | Mains and Earth cable stowed on Mains unit.
46 | 11407 | Fit shipping Eyes to transit blocks.
46 | 11407 | Put thread lock on pendulum screws and transit block Screws.
46 | 11407 | All doors closed and secured.
46 | 11407 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11407 | Print shipping stickers.
5 | 423 | Confirm Liner cable routed as specified
4 | 668 | Fit chamber to damper
4 | 668 | Fit chamber bolts item 10 and 12 and Torque to 42Nm.
4 | 668 | Fit chamber bolts item 13 and 14 and Torque to 21Nm.
4 | 668 | Fit chamber pumping tube and Toruqe bolts item 7 to 12Nm.
4 | 668 | Record Chamber SN
30 | 906 | Store Images in QA folder
648 | 10611 | 30KV Res- ≤2nm
648 | 10611 | 3KV Res- ≤6nm
648 | 10611 | 1KV Res- ≤9nm
648 | 10611 | 200V Res- Not Measured
648 | 10611 | 1KV Oscillation Check - EHT=1KV, Scan Speed=2, WD=10mmm, Mag=10Kx
648 | 10611 | Store Standard LaB6 Resolution + Oscillation Images
399 | 931 | PAT Tester Serial Number
399 | 931 | PAT Tester Cal Expiry Date
399 | 931 | Earth resistance – measured value must be < 0.1 Ω
399 | 931 | Insulation resistance – measured value must be > 1.0M Ω
399 | 931 | Test Certificate Completed & Signed
676 | 10838 | Install SCD Amplifier
676 | 10838 | Install SCD Amplifier
676 | 10838 | SCD Relay PCB-Sigma500 Only
709 | 925 | Ensure Chiller Type is 'AIR' cooled.
709 | 925 | Ensure Chiller Type is 'AIR' cooled.
709 | 925 | Ensure Chiller Type is 'AIR' cooled - (345980-9023-100)
709 | 926 | Ensure Chiller Type is 'WATER' cooled - (345980-9023-100)
709 | 926 | Ensure Chiller Type is 'WATER' cooled - (345980-9024-100)
25 | 198 | Gun X
25 | 198 | Gun Y
25 | 198 | Aper X
25 | 198 | Aper Y
25 | 219 | Normally no cut off is visible @ 10kV, SS9 on the central aperture of the emission image (Condenser off and alignment corrections zeroed). However, some cut off is allowable if certain criteria are met. Please refer to WI for details.
661 | 10710 | Torque Chamber Bolts to 35Nm
661 | 10710 | Record Chamber SN
671 | 10797 | Torque Chamber Bolts To 35Nm
671 | 10797 | Record Chamber SN
40 | 85 | Install ZEN core EM.Exe (This is NOT ZEN Core)
40 | 85 | ZEN core Software Version
40 | 85 | Setup Exchange Position
40 | 85 | Capture Image
46 | 887 | Remove docking brackets and fit transit brackets
46 | 887 | Vent the Chamber.
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Remove On / Off switch panel
46 | 887 | Print and fit DJEL S/N Label
46 | 887 | Remove and pack all power and system cables (not stage cables)
46 | 887 | Remove and pack all Fibre optic cables (not stage cables)
46 | 887 | Remove and pack all cooling pipes and cooling flanges.
46 | 887 | Remove PCB frame complete
46 | 887 | Remove AC unit complete with mains cable
46 | 887 | Remove Cekon Plug From Mains Cable
46 | 887 | Remove Festo valve cluster
46 | 887 | Remove PSU unit
46 | 887 | Remove EHT Set
46 | 887 | Remove IGP PSU
46 | 887 | Remove EBBU Electonics from column and fit shorting plug
46 | 887 | Remove customer SE2 detector (460F only)
46 | 887 | Fit customer SE2 Detectors (2 off) to transport vessels (460 F only).
46 | 887 | Remove BSD4 complete from chamber and fit protective cover over Diode (460F only).
46 | 887 | Pack customer accessories as shown WI (460F Only)
46 | 887 | Remove Water flow and Temperature sensor assembly
46 | 887 | Remove column / support frame complete and attach to column storage frame
46 | 887 | Fit blank plate to column bypass pipe so that lower column can be shipped at vacuum (460F Only).
46 | 887 | Pump down lower column as shown in WI.
46 | 887 | Remove EBBU Electonics from column and fit shorting plug
46 | 887 | Chamber secured to vibration plate - Only required if test tool is being shipped to Bollore
46 | 887 | Chamber Securing Screws Tightened to 25NM - Only required if test tool is being shipped to Bollore
46 | 887 | Secure the pendulum feet using foam blocks and strap - Only required if test tool is being shipped to Bollore
46 | 887 | Print shipping stickers.
676 | 10838 | Install SCD Amplifier
676 | 10838 | SCD Relay PCB-Sigma500 Only
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448 , then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Run ASAP Screening Program
20 | 252 | System Specific Datazone Loaded
44 | 348 | Chamber Door Interlock
44 | 348 | Service Switch
44 | 348 | Lens De-activation
44 | 348 | Detector De-activation
44 | 348 | EHT + IPPSU De-activation
6 | 61 | Safety Earth
6 | 61 | Compressed Air (Set to 0.5MPa)
6 | 61 | PVP Valve
6 | 61 | Power
6 | 61 | Rotary Pump
6 | 61 | Nitrogen (Set to 0.05 MPa)
6 | 61 | EHT Lens output resistance
6 | 61
6 | 61 | Safety Earth
6 | 61 | Compressed Air (Set to 0.5MPa)
6 | 61 | PVP Valve
6 | 61 | Power
6 | 61 | Rotary Pump
6 | 61 | Nitrogen (Set to 0.05 MPa)
6 | 61 | Cooling Water
6 | 61 | EHT Lens output resistance
103 | 940 | The Gun vacuum should be better than 9 x10-10 mb. If this cannot be achieved consult with Man Eng.
44 | 349 | Record Seaward PRIMETEST 100 SN:
44 | 349 | Calibration Due Date:
44 | 349 | Tester verification test; Confirm earth short circuit = (<0.03Ω)
44 | 349 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
44 | 349 | Insulation Test (2.0MΩ - 1.9MΩ)
44 | 349 | Earth Leakage Test (0.9mA - 0.7mA)
44 | 349 | Record Seaward PRIMETEST 100 SN:
44 | 349 | Calibration Due Date:
44 | 349 | Tester verification test; Confirm earth short circuit = (<0.03Ω)
44 | 349 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
44 | 349 | Record Seaward PRIMETEST 100 SN:
44 | 349 | Calibration Due Date:
44 | 349 | Tester verification test; Confirm earth short circuit = (<0.03Ω)
44 | 349 | Test all earth points as shown in procedure. (Max = 0.1Ω) Record max resistance shown
44 | 349 | Record SEAWARD PRIMETEST 100 SN:
44 | 349 | Calibration Due Date:
44 | 349 | Tester verification test; Confirm earth short circuit = (<0.03Ω)
44 | 349 | Test all earth points as shown in procedure. (Max = 0.1Ω) Record max resistance shown
44 | 349 | Confirm correct SN label applied to product: (CSA Compliant / Non-CSA compliant)
390 | 928 | Record SEAWARD PRIMETEST 100 SN:
390 | 928 | Calibration Due Date:
390 | 928 | Tester verification test; Confirm earth short circuit = (<0.03Ω)
390 | 928 | Test all earth points as shown in procedure. (Max = 0.1Ω) Record max resistance shown
390 | 928 | Confirm correct SN label applied to product: (CSA Compliant / Non-CSA compliant)
678 | 930 | Fit China RoHS Compliance Label
44 | 349 | Record Seaward PRIMETEST 100 SN:
44 | 349 | Calibration Due Date:
44 | 349 | Tester verification test; Confirm earth short circuit = (<0.03Ω)
44 | 349 | Test all earth points as shown in procedure. (Max = 0.1Ω) Record max resistance shown
44 | 349 | Confirm correct SN label applied to product: (CSA Compliant / Non-CSA compliant)
678 | 929 | Print and fit the relevant S/N label to the system. See WI for details.
676 | 10838 | Install SCD Amplifier
676 | 10838 | SCD Relay PCB-Sigma500 Only
676 | 10838 | Install SCD Amplifier
676 | 10838 | Fit SCD Relay PCB-Sigma500 Only
4 | 668 | Fit chamber to damper
4 | 668 | Fit chamber bolts item 10 and 12 and Torque to 42Nm.
4 | 668 | Fit chamber bolts item 13 and 14 and Torque to 21Nm.
4 | 668 | Fit chamber pumping tube and Toruqe bolts item 7 to 12Nm.
4 | 668 | Fit Turbo guard with non silver vent screw 000000-0539-927.
4 | 668 | Record Chamber SN
6 | 61 | Safety Earth
6 | 61 | Compressed Air (Set to 0.5MPa)
6 | 61 | PVP Valve
6 | 61 | Power
6 | 61 | Rotary Pump
6 | 61 | Nitrogen (Set to 0.05 MPa)
6 | 61 | Cooling Water
6 | 61 | EHT Set Serial Number SN:
6 | 61 | EHT Lens output resistance ( < 2.2kΩ)
390 | 928 | Record SEAWARD PRIMETEST 100 SN:
390 | 928 | Calibration Due Date:
390 | 928 | Tester verification test; Confirm earth short circuit = (<0.03Ω)
390 | 928 | Test all earth points as shown in procedure. (Max = 0.1Ω) Record max resistance shown
390 | 928 | Confirm correct SN label applied to product: (CSA Compliant / Non-CSA compliant)
676 | 10838 | Install SCD Amplifier
676 | 10838 | Fit SCD Relay PCB - Sigma500 Only
646 | 10654 | Print Out Mineralogic Check List (Link in Documents)
646 | 10654 | Correct PC - Installed
646 | 10654 | Install SmartSEM Licences
646 | 10654 | Instal MIL Redistribution
646 | 10654 | Fit MIL Dongle
646 | 10654 | Install Mineralogic Software
646 | 10654 | Install MAC Standard Recipe Files
646 | 10654 | Calibrate I/O-SCAN card
646 | 10654 | MAC Standard & Stage Postion Set
646 | 10654 | Adjust Recipe for the System
646 | 10654 | Run 20 min Rapid MAC STD Measure
646 | 10654 | Check Results & Image Montage - 20 min Rapid MAC STD Measure
646 | 10654 | Run 90 min High Mag Rapid MAC STD Measure
646 | 10654 | Check Results and Compare with 20 min Run
646 | 10654 | 10 Hour Overnight Run
646 | 10654 | Check Results and Compare with 20 and 90 min Run
646 | 10654 | Store Compressed Database,Montage Image, Log files etc
103 | 940 | Vacuum should be ≤9.00e-10 mbar. Please consult Man Eng if this cannot be achieved.
79 | 941 | Ensure CCD Cables are correctly routed (Not Far RHS of cable rack).
79 | 941 | Ensure the HV cable is corrctly ruted (inside of Ion Pump Bracket)
22 | 262 | Carry out overnight ‘Full service Bakeout’ and note actual duration.
22 | 262 | Confirm that all Thermocouples are reading.
22 | 262 | Confirm that all Ion pump heater elements and flange heaters are getting hot.
22 | 262 | Gun Vacuum post Bakeout. Ensure column has cooled prior to taking reading. If vacuum is not <9.00e-10 please consult Man Eng.
22 | 262 | After completion of Bakeout Torque the Gun head bolts to 25NM.
5 | 1544 | Fit A/L Key to Stage Door
5 | 1544 | Check distance of 200mm door extended = 228mm
651 | 10592 | Check Diode Segments
651 | 10592 | Check Gain Ranges
651 | 10592 | Check Auto Range
651 | 10592 | Check TOPO
40 | 1557 | CCD
40 | 1557 | Focus
390 | 928 | Record SEAWARD PRIMETEST 100 SN:
390 | 928 | Calibration Due Date:
390 | 928 | Tester verification test; Confirm earth short circuit = (<0.03Ω)
390 | 928 | Test all earth points as shown in procedure. (Max = 0.1Ω) Record max resistance shown
390 | 928 | Confirm correct SN label applied to product: (CSA Compliant / Non-CSA compliant)
389 | 10874 | Print Out Mineralogic Check List (Link in Documents)
389 | 10874 | Correct PC - Installed
389 | 10874 | Install SmartSEM Licences
389 | 10874 | Instal MIL Redistribution
389 | 10874 | Fit MIL Dongle
389 | 10874 | Install Mineralogic Software
389 | 10874 | Install MAC Standard Recipe Files
389 | 10874 | Calibrate I/O-SCAN card
389 | 10874 | MAC Standard & Stage Postion Set
389 | 10874 | Adjust Recipe for the System
389 | 10874 | Run 20 min Rapid MAC STD Measure
389 | 10874 | Check Results & Image Montage - 20 min Rapid MAC STD Measure
389 | 10874 | Run 90 min High Mag Rapid MAC STD Measure
389 | 10874 | Check Results and Compare with 20 min Run
389 | 10874 | 10 Hour Overnight Run
389 | 10874 | Check Results and Compare with 20 and 90 min Run
103 | 940 | Vacuum should be ≤9.00e-10 mbar. Please consult Man Eng if this cannot be achieved.
4 | 37 | Shim Stage
4 | 37 | Set Mag Clamp
4 | 37 | Fit Gun Interlock Key to Stage Door
4 | 37 | Record Stage SN (Can be found on right hand side of XY platform
644 | 10497 | System Vacuum & Gun Vacuum is within Spec
644 | 10497 | Gun Chamber has been cleaned and inspected using UV light
644 | 10497 | EHT set socket well has been cleaned and inspected using UV light
644 | 10497 | Clean Anode in Plasma Oven for 20 Minutes
644 | 10497 | Lab6 Firing Unit Fitted
644 | 10497 | Overnight Bakeout
20 | 11038 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
20 | 11038 | PC S/N
20 | 11038 | If SEM Enhanced PC(349535-9001-020) has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Hard disk Image (Version)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
4 | 37 | Shim Stage
4 | 37 | Set Mag Clamp
4 | 37 | Fit Gun Interlock Key to Stage Door
4 | 37 | Record Stage SN (Can be found on right hand side of XY platform)
40 | 945 | Fit the HT-BSD
40 | 945 | Software setup of the HTBSD
40 | 945 | Test the HTBSD
40 | 945 | X-ray check HT-BSD
40 | 945 | System vac in spec (<2.0e-06mbar)
40 | 945 | Remove detector before shipping
678 | 11417 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
678 | 11417 | HAL SN:
678 | 11417 | HAL Calibration Date:
678 | 11417 | DC IR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
678 | 11417 | DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 11417 | Confirm System Passed Hi-Pot Tests
678 | 951 | Secutest SN:
678 | 951 | Secutest Calibration Date:
678 | 951 | PE Test - Confirm all test points (<0.1Ω)
678 | 951 | R INS - ( > 0.5MΩ)
678 | 951 | I HLNC - ( < 500µA)
678 | 951 | I HLSF - ( < 3500µA)
678 | 951 | Confirm EN 61010 Passed
678 | 951 | Confirm Visual Inspection Criterea Passed
678 | 951 | Record maximum leakage current
46 | 371 | X-Ray Certificate
46 | 371 | Quality Certificate
46 | 371 | Quality Feedback Form
46 | 371 | DENKA Data Sheet
46 | 371 | Options Removed packing List
46 | 371 | Options Retained packing List
46 | 371 | Standard Packing List
46 | 371 | PC Inferface assignment
46 | 371 | Declaration of conformity
46 | 371 | Installation report
46 | 371 | Acceptance test record
46 | 371 | Licence Generator report
12 | 1655 | S/N =
12 | 1655 | Please confirm UHV supplier
20 | 952 | Please print 'Declaration of Conformity' from URL link
20 | 952 | Please print 'Declaration of Conformity' from URL link
20 | 952 | Please print 'Declaration of Conformity' from URL link
20 | 953 | Please print 'Declaration of Conformity' from URL link
20 | 953 | Please print 'Declaration of Conformity' from URL link
40 | 945 | Fit the HT-BSD
40 | 945 | Software setup of the HTBSD
40 | 945 | Test the HTBSD
40 | 945 | X-ray check HT-BSD
40 | 945 | System vac in spec (<2.0e-06mbar)
40 | 945 | Remove detector before shipping
40 | 945 | Fit the HT-BSD
40 | 945 | Software setup of the HTBSD
40 | 945 | Test the HTBSD
40 | 945 | X-ray check HT-BSD
40 | 945 | System vac in spec (<2.0e-06mbar)
40 | 945 | Remove detector before shipping
| 468 | Visually inspected for damage/dents
| 468 | Sealing surfaces free from scratches/ imperfections
| 468 | Check Polepiece for correct position of threaded clip holes
| 468 | Visually inspected for damage/dents
| 468 | Sealing surfaces free from scratches/ imperfections
| 468 | Check Polepiece for correct position of threaded clip holes
| 468 | Visually inspected for damage/dents
| 468 | Sealing surfaces free from scratches/ imperfections
| 468 | Check Polepiece for correct position of threaded clip holes
40 | 948 | Fit the Insitu substage
40 | 948 | Test the Insitu substage
40 | 948 | Test the Insitu substage
40 | 948 | Remove detector before shipping
40 | 948 | Fit the Insitu substage
40 | 948 | Test the Insitu substage
40 | 948 | System vac <2.0e-06mbar
40 | 948 | X-ray test
40 | 948 | Repackage all hardware excluding the chamber feed through in Deben packaging
40 | 945 | Fit the HT-BSD
40 | 945 | Software setup of the HTBSD
40 | 945 | Test the HTBSD
40 | 945 | X-ray check HT-BSD
40 | 945 | System vac in spec (<2.0e-06mbar)
40 | 945 | Remove detector before shipping
40 | 948 | Fit the Insitu substage
40 | 948 | Test the Insitu substage
40 | 948 | System vac <2.0e-06mbar
40 | 948 | X-ray test
40 | 948 | Repackage all hardware excluding the chamber feed through in Deben packaging
678 | 11417 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
678 | 11417 | HAL SN:
678 | 11417 | HAL Calibration Date:
678 | 11417 | DC IR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
678 | 11417 | DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 11417 | Confirm System Passed Hi-Pot Tests
46 | 11060 | X-Ray Certificate ----------------------------------------------------------------------------------
46 | 11060 | Quality Certificate ---------------------------------------------------------------------------------
46 | 11060 | Quality Feedback Form
46 | 11060 | DENKA Data Sheet
46 | 11060 | Options Retained packing List
46 | 11060 | Standard Packing List
46 | 11060 | Declaration of Conformity
46 | 11060 | Raith Installation Report
46 | 11060 | Software Licence Agreement
20 | 953 | Please print 'Declaration of Conformity' from URL link
20 | 953 | Please print 'Declaration of Conformity' from URL link
20 | 954 | Please print 'Declaration of Conformity' from URL link
285 | 956 | Please print 'Declaration of Conformity' from URL link
285 | 956 | Please print 'Declaration of Conformity' from URL link
285 | 957 | Please print 'Declaration of Conformity' from URL link
285 | 958 | Please print 'Declaration of Conformity' from URL link
285 | 958 | Please print 'Declaration of Conformity' from URL link
28 | 293 | ESB concentricity in spec
28 | 293 | ESB-Inlens concentricity in spec
28 | 293 | ESB concentricity in spec
28 | 293 | ESB-Inlens concentricity in spec
81 | 1587 | Measured value =
389 | 949 | Fit the HT-BSD
389 | 949 | Software setup of the HTBSD
389 | 949 | Test the HTBSD
389 | 949 | X-ray check HT-BSD
389 | 949 | System vac in spec (<2.0e-06mbar)
389 | 949 | Remove detector before shipping
389 | 950 | Fit the Insitu substage
389 | 950 | Test the Insitu substage
389 | 950 | System vac <2.0e-06mbar
389 | 950 | X-ray test
389 | 950 | Repackage all hardware excluding the chamber feed through in Deben packaging
389 | 917 | Unpack Volutome package and check it's completeness
389 | 917 | Remove transit clamps from Volutome stage
389 | 917 | Ensure hardware has been configured according to work instructions
389 | 917 | Fit Volutome stage
389 | 917 | Connect Volutome stage and controller to the SEM
389 | 917 | Install Volutome software installation package
389 | 917 | Configure Volutome stage motors
389 | 917 | Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator
389 | 917 | Set up the Volutome workbench in Zen
389 | 917 | Record VolumeBSD serial number
389 | 917 | Record Volutome stage serial number
389 | 917 | Record VolumeBSD amplifier serial number
389 | 917 | Record VolumeBSD mechanics serial number
389 | 917 | Record C-885 PI controller serial number
389 | 917 | Record E-727 PI controller serial number
389 | 917 | Function check the Volutome stage light
389 | 917 | Adjust Volutome stage centre (Crosshairs should be within 4 central squares)
389 | 917 | Install Nav cam in-line connections (If applicable)
389 | 917 | Perform the microtome and stage functionality checks
389 | 917 | Fit and align the VolumeBSD diode
389 | 917 | Check VolumeBSD alignment repeatability
389 | 917 | Check all gain ranges of the VolumeBSD (normal/inverted)
389 | 917 | Perform Signal/noise ratio check
389 | 917 | Ensure BSD diode protector is fitted and adjusted in chamber
389 | 917 | Install cable tie to loop BSD signal cable
389 | 917 | Check SCM with Volutome stage
389 | 917 | Ripple X Spec 2nm (600Kx Mag 15KV WD2)
389 | 917 | Ripple Y Spec 2nm (600Kx Mag 15KV WD2)
389 | 917 | Resolution 1KV Spec same as standard resolution
389 | 917 | Resolution 15KV Spec same as standard resolution
389 | 917 | Contamination
389 | 917 | Perform X-ray check
389 | 917 | Pump Down Time
389 | 917 | Ultimate Vacuum
389 | 917 | Focal CC fitted to Volutome stage
389 | 917 | Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar
389 | 917 | Generate live image from Axiocam
389 | 917 | Re-package Volutome and check it's completeness
389 | 917 | ManEng/ Senior Eng/ Production Manger Sign off
4 | 37 | Shim Stage
4 | 37 | Set Mag Clamp
4 | 37 | Fit Gun Interlock Key to Stage Door
4 | 37 | Record Stage SN (Can be found on right hand side of XY platform)
2 | 7 | Record IPPSU S/N
2 | 7 | Set DIP switches
2 | 7 | Qualification Marks Checked
4 | 37 | Shim Stage
4 | 37 | Set Mag Clamp
4 | 37 | Fit Gun Interlock Key to Stage Door
4 | 37 | Record Stage SN (Can be found on right hand side of XY platform)
285 | 956 | Please print 'Declaration of Conformity' from URL link
285 | 957 | Please print 'Declaration of Conformity' from URL link
285 | 958 | Please print 'Declaration of Conformity' from URL link
285 | 959 | Please print 'Declaration of Conformity' from URL link
285 | 959 | Please print 'Declaration of Conformity' from URL link
403 | 960 | Please print 'Declaration of Conformity' from URL link
81 | 1587 | Measured value =
6 | 921 | Record Water Flow Temperature (20ºC +/- 5ºC)
654 | 11420 | PAT Tester Serial Number
654 | 11420 | PAT Tester Cal Expiry Date
654 | 11420 | Earth resistance – measured value must be < 0.1 Ω
654 | 11420 | Insulation resistance – measured value must be > 1.0M Ω
654 | 11420 | Test Certificate Completed & Signed
654 | 11420 | PAT Tester Serial Number
654 | 11420 | PAT Tester Cal Expiry Date
654 | 11420 | Earth resistance – measured value must be < 0.1 Ω
654 | 11420 | Insulation resistance – measured value must be > 1.0M Ω
654 | 11420 | Test Certificate Completed & Signed
654 | 11420 | PAT Tester Serial Number
654 | 11420 | PAT Tester Cal Expiry Date
654 | 11420 | Earth resistance – measured value must be < 0.1 Ω
654 | 11420 | Insulation resistance – measured value must be > 1.0M Ω
654 | 11420 | Test Certificate Completed & Signed
403 | 960 | Please print 'Declaration of Conformity' from URL link
403 | 961 | Please print 'Declaration of Conformity' from URL link
403 | 961 | Please print 'Declaration of Conformity' from URL link
28 | 963 | Cut off ok
28 | 963 | No Charging particles
28 | 963 | Click through Gemini / Crossbeam alignment wizard and confirm mechanical alignment is good.
28 | 292 | No Charging particles
28 | 292 | Ensure no charging particles outside of specification
28 | 294 | Review image jumps run and confirm acceptable
28 | 963 | Click through Gemini / Crossbeam alignment wizard and confirm mechanical alignment is good.
28 | 963 | Confirm Min- Max current deviation = <10um
81 | 964 | Actual
81 | 964 | Actual
81 | 964 | Actual
82 | 1592 | Take and save images at Mag 200, 3kX and 35kX.
82 | 1592 | Mag is within +/- 2% for each.
82 | 1592 | Orthognality is within +/- 0.1 deg for each.
30 | 11071 | Check Mag cal errors are <2% in X & Y axes in all 3 Mag Ranges. Orthog. ± 0.1deg
30 | 11071 | Store accuracy images (Mag 1, Mag 2, Mag 3 and Orthog)
82 | 1591 | Take and save images at Mag 200, 2kX and 35kX.
82 | 1591 | Mag is within +/- 1% for each.
82 | 1591 | Orthognality is within +/- 0.1 deg for each.
28 | 963 | Cut off ok
28 | 963 | No Charging particles
83 | 701 | Confirm no repeated EM Server screening errors have been flagged
30 | 297 | Check Mag cal errors are <1% in X & Y axes in all 3 Mag Ranges. Orthoganality ± 0.1deg.
30 | 297 | Store accuracy images (Mag 1, Mag 2, Mag 3 and Orthog)
82 | 1592 | Take and save images at Mag 200, 3kX and 35kX.
82 | 1592 | Mag is within +/- 2% for each.
82 | 1592 | Orthognality is within +/- 0.1 deg
667 | 11247 | Print Plinth Seperators
667 | 11247 | Sort Documentation with Seperators
667 | 11247 | Print Plinth Seperators
667 | 11247 | Sort Documentation with Seperators
37 | 330 | Stig: 20μm Aper, 10kV is ≤ 35%
37 | 330 | Stig: Off Axis Aper, 10kV is ≤ 45%
37 | 330 | Stig: 20μm Aper, 1kV is ≤ 20%
37 | 330 | Stig: Off Axis Aper, 1kV is ≤ 30%
37 | 330 | Aper/Gun: 20μm, All EHT’s is ≤ 60%
21 | 11321 | Correct Firmware installed
393 | 6934 | PC S/N
393 | 6934 | SmartSEM Version
393 | 6934 | Stage Firmware Version
393 | 6934 | Vacuum Firmware Version
393 | 6934 | EO Firmware Version
285 | 403 | Fit labels to PC serial cable com port 4 to 7
66 | 994 | Z Motor SeriaL Number:
66 | 994 | T Motor Serial Number:
66 | 994 | R Motor Serial Number:
403 | 962 | Please print 'Declaration of Conformity' from URL link
27 | 434 | Image Shift (um) =
671 | 10797 | Torque Chamber Bolts To 10Nm
671 | 10797 | Record Chamber SN
22 | 262 | Carry out overnight ‘Full service Bakeout’ and note actual duration.
22 | 262 | Confirm that all Thermocouples are reading.
22 | 262 | Confirm that all Ion pump heater elements and flange heaters are getting hot.
22 | 262 | Gun Vacuum post Bakeout. Ensure column has cooled prior to taking reading.
22 | 262 | After completion of Bakeout Torque the Gun head bolts to 25NM.
21 | 387 | Order revision in use or 'n/a'
33 | 316 | Actual for 1kV =
33 | 316 | Store 1KV Resolution Images and Pate image
33 | 316 | Actual for 15kV =
33 | 316 | Store 15KV Resolution Images and Pate image
648 | 10611 | 30KV Res- ≤2nm
648 | 10611 | 3KV Res- ≤6nm
648 | 10611 | 1KV Res- ≤9nm
648 | 10611 | 200V Res- Not Measured
648 | 10611 | 1KV Oscillation Check - EHT=1KV, Scan Speed=2, WD=10mmm, Mag=10Kx
648 | 10611 | Store Standard LaB6 Resolution + Oscillation Images
393 | 6934 | PC S/N
393 | 6934 | SmartSEM Version
393 | 6934 | Confirm Stage, EO and Vac FW versions are latest release
403 | 962 | Please print 'Declaration of Conformity' from URL link
82 | 1591 | Take and save images at Mag 200, 2kX and 35kX.
82 | 1591 | Mag is within +/- 1% for each.
82 | 1591 | Orthognality is within +/- 0.1 deg.
84 | 1574 | 5 and 60 Pa without beamsleeve.
84 | 1574 | 5 and 150 Pa in NanoVP Mode
84 | 1574 | 5 and 500 Pa in XVP Mode
44 | 354 | Check Installed Licences Against Sales Order
393 | 6934 | PC S/N
393 | 6934 | SmartSEM Version
393 | 6934 | Confirm Stage, EO and Vac FW versions are latest release
55 | 966 | Checked Carried Our By:
393 | 3345 | EHT set remains powered on yellow button
393 | 3345 | After restart EM Server boots with no errors
393 | 3346 | Ensure all customer licences are loaded in the Administrator Panel and functions correctly
393 | 3346 | Ensure all customer licences are loaded in the Administrator Panel and functions correctly
393 | 3346 | Ensure all customer licences are loaded in the Administrator Panel and functions correctly
393 | 3346 | Ensure all customer licences are loaded in the Administrator Panel and functions correctly
393 | 3346 | Ensure all customer licences are loaded in the Administrator Panel and functions correctly
393 | 3346 | Ensure the Omit dongle tick box is un-ticked in the Factory tab in Administrator Panel.
393 | 3346 | Ensure all customer licences are loaded in the Administrator Panel and functions correctly
393 | 3346 | Ensure all customer licences are loaded in the Administrator Panel and functions correctly
393 | 3346 | Ensure the Omit dongle tick box is un-ticked in the Factory tab in Administrator Panel.
393 | 3346 | Under the Service Tab ensure the Hardware Dongle and System S/No. Match.
394 | 3351 | Chamber vacuum can be changed continuously between 10 and 133 PA (Beam On)
382 | 6933 | Spacer Height Setup
382 | 6933 | Stage stops when touch alarm is initiated or F12 is pressed
382 | 6933 | Stage stops when touch alarm is initiated or F12 is pressed
396 | 3360 | Filament Serial number
396 | 3360 | Filament heating Current
396 | 3360 | Extrator Power at 0.2 mA/sr
396 | 3360 | Total Current at 0.2 mA/sr
396 | 3361 | Measured Value
60 | 842 | Z Microswitch Position Set
60 | 842 | T Microswitch Position Set
60 | 842 | R Microswitch Position Set
396 | 10885 | Check That all detectors are listed under and signed off in X-Ray checks, If any extra detectors have been fitted perform X-Ray checks on those detectors
396 | 10885 | Sign XRay Certificate
396 | 10885 | Check That all detectors are listed under and signed off in X-Ray checks, If any extra detectors have been fitted perform X-Ray checks on those detectors
396 | 10885 | Sign XRay Certificate
66 | 986 | Checked Carried Our By:
66 | 993 | Z Microswitch Position Set
66 | 993 | T Microswitch Position Set
66 | 993 | R Microswitch Position Set
403 | 890 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
4 | 11384 | Warning -ONLY VENT using the electronic vent valve in the chamber
4 | 11384 | No Leaks Detected (Prior to running up Turbo)
4 | 11384 | No Leaks Detected (With Turbo running) when in high sensitivity
5 | 10959 | Warning -ONLY VENT using the electronic vent valve in the chamber
5 | 10959 | No Leaks Detected (Prior to running up Turbo)
5 | 10959 | No Leaks Detected (With Turbo running) when in high sensitivity
4 | 11384 | Warning -ONLY VENT using the electronic vent valve in the chamber
4 | 11384 | No Leaks Detected (Prior to running up Turbo)
4 | 11384 | No Leaks Detected (With Turbo running) when in high sensitivity
5 | 10959 | Warning -ONLY VENT using the electronic vent valve in the chamber
5 | 10959 | No Leaks Detected (Prior to running up Turbo)
5 | 10959 | No Leaks Detected (With Turbo running) when in high sensitivity
6 | 68 | Warning -ONLY VENT using the electronic vent valve in the chamber
6 | 68 | No Leaks Detected (Prior to running up Turbo)
6 | 68 | No Leaks Detected (With Turbo running) when in high sensitivity
20 | 972 | Service centre S/W installed on Customer PC.
6 | 1546 | HAL SN:
6 | 1546 | HAL Calibration Date:
6 | 1546 | IGP Heaters (346022-8120-000). . .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 710 | Opening of manual valves required to achieve gun vacuum
6 | 710 | Actual IGP1 vacuum
6 | 710 | Portable IGP controller fitted? (See work instruction for details)
399 | 969 | Clean Front Panel
399 | 969 | Prepare Logo
399 | 969 | Fitting Logo
654 | 970 | Clean Front Panel
654 | 970 | Prepare Logo
654 | 970 | Fitting Logo
654 | 970 | Clean Front Panel
654 | 970 | Prepare Logo
654 | 970 | Fitting Logo
654 | 970 | Clean Front Panel
654 | 970 | Prepare Logo
654 | 970 | Fitting Logo
650 | 971 | Clean Front Panel
650 | 971 | Prepare Logo
650 | 971 | Fitting Logo
6 | 710 | Opening of manual valves required to achieve gun vacuum
6 | 710 | Actual IGP1 vacuum
6 | 710 | Portable IGP controller fitted? (See work instruction for details)
6 | 710 | Opening of column manual valves required to achieve gun vacuum
6 | 710 | Actual IGP1 vacuum
6 | 710 | Portable IGP controller fitted? (See work instruction for details)
377 | 3261 | Opening of column manual valve/s required to achieve gun vacuum
377 | 3261 | Bake Out Hours
397 | 3363 | Confirm no noise can be seen on fastest scan speeds
397 | 3364 | Changing the collector bias voltage bias results in a change of the signal intensity
397 | 3364 | Collector - Changing the collector bias voltage bias results in a change of the signal intensity
397 | 3364 | Scintillator - Confirm switching scintillator on/off results in a change of the signal intensity
654 | 970 | Clean Front Panel
654 | 970 | Prepare Logo
654 | 970 | Fitting Logo
399 | 969 | Clean Front Panel
399 | 969 | Prepare Logo
399 | 969 | Fitting Logo
22 | 262 | Opening of column manual valve/s required to achieve gun vacuum
22 | 262 | Carry out overnight ‘Full service Bakeout’ and note actual duration.
22 | 262 | Confirm that all Thermocouples are reading.
22 | 262 | Confirm that all Ion pump heater elements and flange heaters are getting hot.
22 | 262 | Gun Vacuum post Bakeout. Ensure column has cooled prior to taking reading.
22 | 262 | After completion of Bakeout Torque the Gun head bolts to 25NM.
397 | 3366 | Check each gain ranges and functionality of each segment
397 | 3366 | Confirm there is no rotation of the BSD Diode when inserting the rod.
397 | 3366 | Confirm is smooth action when rod inserted
6 | 1546 | HAL SN:
6 | 1546 | HAL Calibration Date:
6 | 1546 | IGP Heaters (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
9 | 1125 | Aperture cleaned using IPA and 1mm brush
9 | 1125 | Aperture run through ultra-sonic tank (in upright position)
9 | 1125 | Aperture inspected and confimred clean. Eng 1 Initials:
9 | 1125 | Aperture inspected and confimred clean. Eng 2 Initials:
9 | 1125 | Aperture cleaned using IPA and 1mm brush
9 | 1125 | Aperture run through ultra-sonic tank (in upright position)
9 | 1125 | Aperture inspected and confimred clean. Eng 1 Initials:
9 | 1125 | Aperture inspected and confimred clean. Eng 2 Initials:
9 | 1125 | Confirm differential aperture is long bore PN:
397 | 3366 | Confirm there is no rotation of the BSD Diode when inserting the rod.
397 | 3366 | Confirm is smooth action when rod inserted
397 | 3366 | Check each gain ranges and functionality of each segment
397 | 3366 | Check quality of BSD images
397 | 10914 | Confirm a signal intesity change can be seen when changing bias voltage.
397 | 10914 | Confirm a signal intesity change can be seen when changing bias voltage.
285 | 403 | Fit labels to PC serial cable com port 4 to 7
389 | 3303 | Function Test
389 | 3303 | Disable Confirmation Check
389 | 3303 | Turbo Load Check
389 | 3303 | Save .CSV of quite mode cycle time
396 | 3358 | Reading (220 +/- 10pA)
396 | 3358 | Beam Blank - Confirm
396 | 3358 | Reading (220 +/- 10pA)
396 | 3358 | Beam Blank - Confirm reading goes under-range when beam blanked.
396 | 3359 | Reading (300 +/- 10pA)
396 | 3359 | Beam Blank - Confirm reading goes under-range when beam blanked.
397 | 3376 | All controls function correctly
397 | 3376 | Ensure keyboard language is correct
397 | 10913 | Run Up & Obtain An Image
397 | 10913 | Confirm Image Stitching is acceptable
397 | 10915 | Run quick sample clean and confirm no errors
397 | 10916 | Sample can load without any issues
397 | 10916 | Exchange and Resume Macros function correctly from conrol panel / airlock panel
398 | 3391 | 10kV All apertures - Confirm wobble can be achieved in user align window <20%
396 | 3358 | Reading (220 +/- 10pA)
396 | 3358 | Confirm Ext V is running <500V over data sheet
396 | 3358 | Beam Blank - Confirm reading goes under-range when beam blanked.
82 | 1593 | 10kV, 30um - Aperture and Guns <50%
397 | 10914 | Confirm a signal intesity change can be seen when changing bias voltage.
397 | 10914 | No breakdown of the detector can be seen at max bias voltage
397 | 10916 | Run Up & Obtain An Image
397 | 10921 | Check Aquire Spectrum image
397 | 10921 | Check Line Scanning
397 | 10921 | Check Mapping
397 | 10921 | Check EDX Icon Is Present In All User Toolbars
397 | 10921 | EDAX Calibration File Saved To SmartSEM USB Stick
396 | 3358 | Reading (220 +/- 10pA)
396 | 3358 | Confirm Ext V is running <500V over data sheet @220pA
396 | 3358 | Beam Blank - Confirm reading goes under-range when beam blanked.
82 | 1593 | 10kV, 30um - Aperture and Guns <50%
6 | 625 | HAL SN:
6 | 625 | HAL Calibration Date:
6 | 625 | IGP Heaters (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | HAL SN:
6 | 625 | HAL Calibration Date:
6 | 625 | IGP Heaters (IGP#1) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | IGP Heaters (IGP#2) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
9 | 1125 | Aperture cleaned using IPA and 1mm brush
9 | 1125 | Aperture run through ultra-sonic tank (in upright position)
9 | 1125 | Aperture inspected and confimred clean. Eng 1 Initials:
9 | 1125 | Aperture inspected and confimred clean. Eng 2 Initials:
9 | 1125 | Confirm differential aperture is long bore PN: 349520-0086-020
397 | 10912 | Check correct operation of Quiet Mode. Deselect Quiet Mode and ensure rotary pump starts
397 | 10912 | Load previously saved .CSV file to confirm acceptable cycle times
397 | 10922 | Load project image and confirm ok
82 | 1593 | (10kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (10kV, Outer Apertures) - Stigmation
82 | 1594 | Image with Condenser ON
82 | 1594 | Check table editor graph for outliers (Stigs, Apertures, Guns)
82 | 1594 | Check table editor graph plaots for outliers
82 | 1594 | Check table editor graph for outliers (Stigs, Apertures, Guns)
82 | 1594 | Confirm only 2mm & 5mm WD entries in table
82 | 1594 | Check table editor graph for outliers (Stigs, Apertures, Guns)
82 | 1594 | Confirm only 2mm & 5mm WD entries stored in table
82 | 1594 | Check table editor graph for outliers (Stigs, Apertures, Guns)
82 | 1594 | Confirm only 2mm & 5mm WD entries stored in table
82 | 1594 | Confirm only 1 30kV entry in table
82 | 974 | Check table editor graph for grouping and outliers (Apertures, Guns)
82 | 974 | Wall of Stigs - Confirm no outliers and
82 | 974 | Check table editor graph for grouping and outliers (Apertures, Guns)
82 | 974 | Wall of Stigs - Confirm no outliers and uniform grouping
82 | 1594 | Check table editor graph for outliers (Stigs, Apertures, Guns)
82 | 1594 | Confirm only 2mm & 5mm WD entries stored in table
82 | 1594 | Confirm only 1 30kV entry in table
82 | 1594 | Save PDF of table editor plots
82 | 1594 | Check table editor graph for outliers (Stigs, Apertures, Guns)
82 | 1594 | Confirm only 2mm & 5mm WD entries stored in table
82 | 1594 | Confirm only 1 30kV entry in table
82 | 1594 | Save PDF of table editor plots
82 | 1594 | Check table editor graph for outliers (Stigs, Apertures, Guns)
82 | 1594 | Confirm only 2mm & 5mm WD entries stored in table
82 | 1594 | Confirm only 1 30kV entry in table
82 | 1594 | Save PDF of table editor plots
397 | 10923 | Confirm oxford software is loaded
82 | 974 | Check table editor graph for grouping and outliers (Apertures, Guns)
82 | 974 | Wall of Stigs - Confirm no outliers and uniform grouping
82 | 974 | Save PDF of table editor plots
34 | 318 | Jumps - Limit is <= 1 Event / 2 Hours............................................... Actual =
34 | 318 | Stigmation change - Limit <= 1% change / 12 Hours....................... Actual =
44 | 923 | Remove all old / invalid data and images from the PC.
44 | 923 | Ensure all image folders are labelled accordingly.
44 | 923 | Ensure all 'return to stock' or 'reject parts have been processed.
44 | 358 | Confirm all deviations have been cleared on OST - The Andon lamp should be green
44 | 358 | All verification & Safety tests done and within threshold limits defined
44 | 358 | Sales Order is latest version
44 | 358 | All items on Sales order accounted for
26 | 1549 | Objective Cap - No particles visible
26 | 1549 | Lower Aperture - No large particles visible (See WI for details)
26 | 1549 | Upper Aperture - No cut-off on image and no particles visible.
650 | 971 | Clean Front Panel
650 | 971 | Prepare Logo
650 | 971 | Fitting Logo
397 | 10923 | Confirm oxford software is loaded
397 | 10923 | Check Licenses
81 | 1588 | Measured value =
81 | 1588 | Save
81 | 1588 | Measured value =
81 | 1588 | Save annotated image of Max Current
82 | 974 | Check table editor graph for grouping and outliers (Apertures, Guns)
82 | 974 | Wall of Stigs - Confirm no outliers and uniform grouping
82 | 974 | Review DC tablefiller images
82 | 974 | Save PDF of table editor plots
82 | 1593 | (10kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (10kV, Central Aperture) - Stigs <20%
82 | 1593 | (1kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (10kV, Outer Apertures) - Stigs <30%
82 | 1593 | (1kV, Outer Apertures) - Stigs <30%
82 | 1593 | (10kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (10kV, Central Aperture) - Stigs <20%
82 | 1593 | (1kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (10kV, Outer Apertures) - Stigs <30%
82 | 1593 | (1kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (1kV, Outer Apertures) - Stigs <30%
82 | 1593 | (10kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (10kV, Central Aperture) - Stigs <20%
82 | 1593 | (10kV, Outer Apertures) - Stigs <30%
82 | 1593 | (1kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (1kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (1kV, Outer Apertures) - Stigs <30%
82 | 1593 | (10kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (10kV, Central Aperture) - Stigs <20%
82 | 1593 | (10kV, Outer Apertures) - Stigs <30%
82 | 1593 | (1kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (1kV, Central Aperture) - Stigs <20%
82 | 1593 | (1kV, Outer Apertures) - Stigs <30%
82 | 1593 | (10kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (10kV, Central Aperture) - Stigs <20%
82 | 1593 | (10kV, Outer Apertures) - Stigs <30%
82 | 1593 | (1kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (1kV, Central Aperture) - Stigs <20%
82 | 1593 | (1kV, Outer Apertures) - Stigs <30%
82 | 1593 | <10% user aperture wobble required on all apertures
68 | 1056 | Stage Axis Cycle Test with 0.5Kg Weight for 10 Mins
34 | 318 | Jumps - Limit is <= 1 Event / 2 Hours............................................... Actual =
34 | 318 | Stigmation change - Limit <= 1% change / 12 Hours....................... Actual =
13 | 1129 | Ensure the EBBU pins do not protrude from the EBBU (See work instruction for details.
82 | 975 | Confirm VP alignment values look reasonable in table editor
82 | 1593 | (10kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (10kV, Central Aperture) - Stigs <20%
82 | 1593 | (10kV, Outer Apertures) - Stigs <30%
82 | 1593 | (1kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (1kV, Central Aperture) - Stigs <20%
82 | 1593 | (1kV, Outer Apertures) - Stigs <30%
82 | 1593 | <10% user align wobble required on all apertures
68 | 1048 | Checked Carried Our By:
68 | 1050 | Check Set Point X Axis Micro Switch
68 | 1050 | Check Set Point Y Axis Micro Switch
68 | 1050 | Check Set Point Z Axis Micro Switch
68 | 1050 | Check Set point T Axis (-10.8º)
68 | 1051 | Check Stage Initialises Correctly
68 | 1052 | Select Stage Type and Check Travel
68 | 1055 | Check Tilt Travel -10° to +90°
68 | 1057 | Run Axis Cycle Test
68 | 1062 | Select Stage Type
68 | 1062 | Stage Serial Number:
68 | 1063 | Quality Sign Off Carried Out By:
82 | 975 | Confirm VP alignment values are present and look reasonable in table editor
82 | 973 | (10kV, Central Aperture) - Aperture and Guns <50%
82 | 973 | (10kV, Central Aperture) - Stigs <20%
82 | 973 | (10kV, Outer Apertures) - Stigs <30%
82 | 973 | (1kV, Central Aperture) - Aperture and Guns <50%
82 | 973 | (1kV, Central Aperture) - Stigs <20%
82 | 973 | (1kV, Outer Apertures) - Stigs <30%
82 | 973 | <10% user align wobble required on all apertures
28 | 290 | Gun running to Datasheet
28 | 290 | Gun Vacuum better than 5e-9mbar (EHT on/off)
28 | 290 | Nominal Ext V acceptable
28 | 10647 | Confirm SEM Status that (C1 Double only), C2 & C3 Cal
28 | 10647 | System Align Offset Angle Approx: (-21 for X-Beam, -51 for Sphinx, -110 for Raith, N/A for D/C)
28 | 10647 | Confirm SEM Status that (C1 Double only), C2 & C3 Cal
28 | 10647 | System Align Offset Angle Approx: (-21 for X-Beam) (-51 for Sphinx) (-110 for Raith) (N/A for D/C)
28 | 10647 | Confirm SEM Status that (C1 Double only), C2 & C3 Cal are reasonable
28 | 10647 | System Align Offset Angle Approx: (-21 for X-Beam) (-51 for Sphinx) (-110 for Raith) (N/A for D/C)
82 | 1594 | Check table editor graph for outliers (Stigs, Apertures, Guns)
82 | 1594 | Confirm only 2mm & 5mm WD entries stored in table
82 | 1594 | Confirm only 1 30kV entry in table
82 | 1594 | Confirm widefiled entry (Gem 560 ONLY)
82 | 1594 | Save PDF of table editor plots
82 | 1594 | Check table editor graph for outliers (Stigs, Apertures, Guns)
82 | 1594 | Confirm only 2mm & 5mm WD entries stored in table
82 | 1594 | Confirm only 1x 30kV entry in table
82 | 1594 | Confirm widefiled entry (Gem 560 ONLY)
82 | 1594 | Save PDF of table editor plots
82 | 975 | Confirm VP alignment values are present and look reasonable in table editor
82 | 973 | (10kV, Central Aperture) - Aperture and Guns <50%
82 | 973 | (10kV, Central Aperture) - Stigs <35%
82 | 973 | (10kV, Outer Apertures) - Stigs <30%
82 | 973 | (1kV, Central Aperture) - Aperture and Guns <50%
82 | 973 | (1kV, Central Aperture) - Stigs <20%
82 | 973 | (1kV, Outer Apertures) - Stigs <30%
82 | 973 | <10% user align wobble required on all apertures
68 | 1053 | Select Stage Type and Check Travel
68 | 1058 | Stage Kit
68 | 1058 | 27mm Spacer Block
68 | 1058 | 54mm Spacer Block
68 | 1058 | 81mm Spacer Block - Fitted
6 | 625 | HAL SN:
6 | 625 | HAL Calibration Date:
6 | 625 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 625 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | HAL SN:
6 | 1546 | HAL Calibration Date:
6 | 1546 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
382 | 1126 | Store Orth Picture Mag Range 1
382 | 1126 | Store Orth Picture Mag Range 2
82 | 973 | (10kV, Central Aperture) - Aperture and Guns <50%
82 | 973 | (10kV, Central Aperture) - Stigs <35%
82 | 973 | (10kV, Outer Apertures) - Stigs <45%
82 | 973 | (1kV, Central Aperture) - Aperture and Guns <50%
82 | 973 | (1kV, Central Aperture) - Stigs <20%
82 | 973 | (1kV, Outer Apertures) - Stigs <30%
82 | 973 | <10% user align wobble required on all apertures
68 | 1054 | Select Stage Type and Check Travel
68 | 1059 | Stage Kit
68 | 1059 | 27mm Spacer Block Fitted
70 | 1073 | Checked Carried Our By:
68 | 1060 | Stage Kit
70 | 1079 | Z Microswitch Position Set
70 | 1079 | T Microswitch Position Set
70 | 1079 | R Microswitch Position Set
70 | 1080 | Z Motor SeriaL Number:
70 | 1080 | T Motor Serial Number:
70 | 1080 | R Motor Serial Number:
70 | 1082 | Checked Carried Our By:
34 | 318 | Jumps - Limit is <= 1 Event / 2 Hours............................................... Actual =
34 | 318 | Stigmation change - Limit <= 1% change / 12 Hours....................... Actual =
34 | 318 | Jumps - Limit is <= 1 Event / 2 Hours............................................... Actual =
34 | 318 | Stigmation change - Limit <= 1% change / 12 Hours....................... Actual =
68 | 1052 | Confirm X Travel (0 - 130mm)
68 | 1052 | Confirm Y Travel (0 - 130mm)
68 | 1052 | Confirm Z Travel (0 - 50mm)
68 | 1052 | Check Tilt Travel -10° to +90°
66 | 1008 | Checked Carried Our By:
66 | 994 | Z Motor SeriaL Number:
66 | 994 | T Motor Serial Number:
66 | 994 | R Motor Serial Number:
66 | 996 | Checked Carried Our By:
382 | 1126 | Store Orth Picture Mag Range 1
382 | 1126 | Store Orth Picture Mag Range 2
68 | 1053 | Confirm X Travel (0 - 125mm)
68 | 1053 | Confirm Y Travel (0 - 125mm)
68 | 1053 | Confirm Z Travel (0 - 50mm)
68 | 1053 | Check Tilt Travel -10° to +90°
68 | 1053 | Confirm X Travel (0 - 125mm)
68 | 1053 | Confirm Y Travel (0 - 125mm)
68 | 1053 | Confirm Z Travel (0 - 50mm)
68 | 1053 | Check Tilt Travel -10° to +90°
26 | 1549 | Objective Cap - No particles visible
26 | 1549 | Lower Aperture - No large particles visible (See WI for details)
26 | 1549 | Upper Aperture - No cut-off on image and no particles visible.
66 | 1022 | Check Set Point
66 | 1023 | Check Set Point
66 | 1024 | Check Set Point
66 | 1025 | Check Set point (-10.8º)
66 | 1024 | Check Set Point X Axis Micro Switch
66 | 1024 | Check Set Point Y Axis Micro Switch
66 | 1024 | Check Set Point Z Axis Micro Switch
66 | 1024 | Check Set Point T Axis Micro Switch
66 | 1019 | Checked Carried Our By:
66 | 1026 | Check Stage Initialises Correctly
66 | 1027 | Select Stage Type and Check Travel
66 | 1028 | Select Stage Type and Check Travel
66 | 1029 | Select Stage Type and Check Travel
66 | 1030 | Check Tilt Travel -10° to +90°
66 | 1031 | Stage Axis Cycle Test with 0.5Kg Weight for 10 Mins
66 | 1032 | Run Axis Cycle Test
66 | 1038 | Stage Kit
66 | 1038 | 27mm Spacer Block Fitted
66 | 1037 | Stage Kit
66 | 1035 | Select Stage Type
66 | 1035 | Stage Serial Number:
66 | 1033 | Quality Sign Off Carried Out By:
66 | 1024 | Check Set Point X Axis Micro Switch
66 | 1024 | Check Set Point Y Axis Micro Switch
66 | 1024 | Check Set Point Z Axis Micro Switch
66 | 1024 | Check Set Point T Axis Micro Switch
66 | 1024 | Check Set Point X Axis Micro Switch
66 | 1024 | Check Set Point Y Axis Micro Switch
66 | 1024 | Check Set Point Z Axis Micro Switch
66 | 1024 | Check Set Point T Axis Micro Switch
66 | 1034 | Stage Kit
66 | 1034 | 27mm Spacer Block
66 | 1034 | 54mm Spacer Block
66 | 1034 | 81mm Spacer Block - Fitted
70 | 1095 | Checked Carried Our By:
68 | 1059 | Stage Kit
68 | 1059 | 27mm Spacer Block Fitted
66 | 1023 | Check Set Point
66 | 1024 | Check Set Point X Axis Micro Switch
66 | 1024 | Check Set Point Y Axis Micro Switch
66 | 1024 | Check Set Point Z Axis Micro Switch
66 | 1024 | Check Set Point T Axis Micro Switch
66 | 1024 | Check Set point (-10.8º)
40 | 1554 | Obtain an image using the CL sample.(See WI for detials)
82 | 1591 | Take and save images at Mag 200, 2kX and 35kX.
82 | 1591 | Mag is within +/- 1% for each.
82 | 1591 | Orthognality is within +/- 0.1 deg.
82 | 1591 | Take and save images at Mag 200, 2kX and 35kX.
82 | 1591 | Mag is within +/- 1% for each.
82 | 1591 | Orthognality is within +/- 0.1 deg.
14 | 185 | IGP Pressure (≤ 9.0e-10 mbar)
45 | 362 | EM Server.db
45 | 362 | Em Server.Log
45 | 362 | MyBackup.smartbackup
45 | 362 | L500.ini
45 | 362 | License file.xml
45 | 362 | Registry file.reg
45 | 362 | CalibSC.dat
45 | 11058 | Em Server.db
45 | 11058 | Em Server.log Text File
45 | 11058 | Mybackup.Smartbackup
45 | 11058 | Image Folder
45 | 11058 | L500.ini
45 | 11058 | Dynamic EHT.MN2
45 | 11058 | Probe Stability.MN2
45 | 11058 | License file.xml
45 | 11058 | Registry file.reg
45 | 11058 | System Checker Full Report
45 | 11058 | Export Alignment Data From Table Editor(.MERLINCDB)
32 | 311 | Actual =
32 | 311 | Store objective drift results
68 | 1054 | Confirm X Travel (0 - 80mm)
68 | 1054 | Confirm Y Travel (0 - 80mm)
68 | 1054 | Confirm Z Travel (0 - 35mm)
68 | 1054 | Check Tilt Travel -10° to +90°
71 | 1118 | Stage Kit
71 | 1118 | 27mm Spacer Block Fitted
66 | 1024 | Check Set Point X Axis Micro Switch
66 | 1024 | Check Set Point Y Axis Micro Switch
66 | 1024 | Check Set Point Z Axis Micro Switch
66 | 1024 | Check Set point T Axis (-10.8º)
71 | 1107 | Checked Carried Our By:
71 | 1109 | Check Set Point X Axis Micro Switch
71 | 1109 | Check Set Point Y Axis Micro Switch
71 | 1109 | Check Set Point Z Axis Micro Switch
71 | 1109 | Check Set point T Axis (-10.8º)
71 | 1110 | Check Stage Initialises Correctly
71 | 1111 | Select Stage Type and Check Travel
71 | 1112 | Select Stage Type and Check Travel
71 | 1113 | Select Stage Type and Check Travel
71 | 1114 | Check Tilt Travel -10° to +90°
71 | 1115 | Stage Axis Cycle Test with 0.5Kg Weight for 10 Mins
71 | 1116 | Run Axis Cycle Test
71 | 1117 | Stage Kit
71 | 1117 | 27mm Spacer Block
71 | 1117 | 54mm Spacer Block
71 | 1117 | 81mm Spacer Block - Fitted
71 | 1118 | Stage Kit
71 | 1118 | 27mm Spacer Block Fitted
32 | 311 | Actual =
32 | 311 | Store objective drift results
26 | 1549 | Objective Cap - No particles visible
26 | 1549 | Lower Aperture - No large particles visible (See WI for details)
26 | 1549 | Upper Aperture - No cut-off on image and no particles visible.
32 | 311 | Limit is 600nm .......................Actual =
32 | 311 | Store objective drift results
46 | 887 | Remove docking brackets and fit transit brackets
46 | 887 | Vent the Chamber.
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Remove On / Off switch panel
46 | 887 | Remove and pack all power and system cables (not stage cables)
46 | 887 | Remove and pack all Fibre optic cables (not stage cables)
46 | 887 | Remove and pack all cooling pipes and cooling flanges.
46 | 887 | Remove PCB frame complete
46 | 887 | Remove AC unit complete with mains cable
46 | 887 | Remove Cekon Plug From Mains Cable
46 | 887 | Remove Festo valve cluster
46 | 887 | Remove PSU unit
46 | 887 | Remove EHT Set
46 | 887 | Remove IGP PSU
46 | 887 | Remove EBBU Electonics from column and fit shorting plug
46 | 887 | Remove customer SE2 detector (460F only)
46 | 887 | Fit customer SE2 Detectors (2 off) to transport vessels (460 F only).
46 | 887 | Remove BSD4 complete from chamber and fit protective cover over Diode (460F only).
46 | 887 | Pack customer accessories as shown WI (460F Only)
46 | 887 | Remove Water flow and Temperature sensor assembly
46 | 887 | Remove column / support frame complete and attach to column storage frame
46 | 887 | Fit blank plate to column bypass pipe so that lower column can be shipped at vacuum (460F Only).
46 | 887 | Pump down lower column as shown in WI.
46 | 887 | Remove EBBU Electonics from column and fit shorting plug
46 | 887 | Chamber secured to vibration plate - Only required if test tool is being shipped to Bollore
46 | 887 | Chamber Securing Screws Tightened to 25NM - Only required if test tool is being shipped to Bollore
46 | 887 | Secure the pendulum feet using foam blocks and strap - Only required if test tool is being shipped to Bollore
46 | 887 | Print shipping stickers.
708 | 11236 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed.
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
26 | 1549 | Objective Cap - No particles visible
26 | 1549 | Lower Aperture - No large particles visible (See WI for details)
26 | 1549 | Upper Aperture - No cut-off on image and no particles visible.
71 | 1119 | Stage Kit
71 | 1121 | Select Stage Type
71 | 1121 | Stage Serial Number:
71 | 1122 | Quality Sign Off Carried Out By:
71 | 1117 | Stage Kit
71 | 1117 | 27mm Spacer Block
71 | 1117 | 54mm Spacer Block
71 | 1121 | Stage Serial Number:
71 | 1116 | Run Axis Cycle Test
71 | 1116 | Stage Axis Cycle Test with 0.5Kg Weight for 10 Mins
71 | 1116 | Run Axis Cycle Test (1Hr)
71 | 1116 | Stage Axis Cycle Test with 0.5Kg Weight for 10 Mins
71 | 1116 | Run Axis Cycle Test (1Hr)
71 | 1116 | Stage Axis Cycle Test with 0.5Kg Weight for 10 Mins
68 | 1052 | Confirm X Travel (0 - 130mm)
68 | 1052 | Confirm Y Travel (0 - 130mm)
68 | 1052 | Confirm X Travel (0 - 130mm)
68 | 1052 | Confirm Y Travel (0 - 130mm)
68 | 1052 | Confirm Z Travel (0 - 50mm)
68 | 1052 | Confirm X Travel (0 - 130mm)
68 | 1052 | Confirm Y Travel (0 - 130mm)
68 | 1052 | Confirm Z Travel (0 - 50mm)
68 | 1053 | Confirm X Travel (0 - 130mm)
68 | 1053 | Confirm Y Travel (0 - 130mm)
68 | 1053 | Confirm Z Travel (0 - 50mm)
68 | 1053 | Confirm X Travel (0 - 125mm)
68 | 1053 | Confirm Y Travel (0 - 125mm)
68 | 1053 | Confirm Z Travel (0 - 50mm)
68 | 1054 | Confirm X Travel (0 - 125mm)
68 | 1054 | Confirm Y Travel (0 - 125mm)
68 | 1054 | Confirm Z Travel (0 - 50mm)
68 | 1054 | Confirm X Travel (0 - 125mm)
68 | 1054 | Confirm Y Travel (0 - 125mm)
68 | 1054 | Confirm Z Travel (0 - 50mm)
68 | 1054 | Confirm X Travel (0 - 80mm)
68 | 1054 | Confirm Y Travel (0 - 125mm)
68 | 1054 | Confirm Z Travel (0 - 50mm)
68 | 1054 | Confirm X Travel (0 - 80mm)
68 | 1054 | Confirm Y Travel (0 - 80mm)
68 | 1054 | Confirm Z Travel (0 - 35mm)
68 | 1052 | Confirm X Travel (0 - 130mm)
68 | 1052 | Confirm Y Travel (0 - 130mm)
68 | 1052 | Confirm Z Travel (0 - 50mm)
68 | 1052 | Check Tilt Travel -10° to +90°
68 | 1056 | Run Axis Cycle Test (1Hr)
68 | 1056 | Stage Axis Cycle Test with 0.5Kg Weight for 10 Mins
68 | 1052 | Confirm X Travel (0
68 | 1053 | Confirm X Travel (0 - 125mm)
68 | 1053 | Confirm Y Travel (0 - 125mm)
68 | 1053 | Confirm Z Travel (0 - 50mm)
68 | 1053 | Check Tilt Travel -10° to +90°
68 | 1054 | Confirm X Travel (0 - 80mm)
68 | 1054 | Confirm Y Travel (0 - 80mm)
68 | 1054 | Confirm Z Travel (0 - 35mm)
68 | 1054 | Check Tilt Travel -10° to +90°
6 | 1546 | HAL SN:
6 | 1546 | HAL Calibration Date:
6 | 1546 | IGP Heaters (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
32 | 311 | Limit is 600nm .......................Actual =
32 | 311 | Store objective drift results
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
20 | 1133 | Print Shipment Labels
20 | 1133 | Label Material Boxes
20 | 1133 | Place left over stickes in Document Holder
68 | 1056 | Run Axis Cycle Test (min - 1Hr)
68 | 1056 | Stage Axis Cycle Test with 0.5Kg Weight for 10 Mins
68 | 1052 | Confirm X Travel (0 - 130mm)
71 | 1111 | Confirm X Travel (0 - 125mm)
71 | 1111 | Confirm Y Travel (0 - 125mm)
71 | 1111 | Confirm Z Travel (0 - 50mm)
71 | 1111 | Check Tilt Travel -10° to +90°
6 | 1546 | HAL SN:
6 | 1546 | HAL Calibration Date:
6 | 1546 | IGP Heaters (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
6 | 1546 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
650 | 1124 | Quality Certificate Signed By:
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
71 | 1116 | Run Axis Cycle Test (min - 1Hr)
71 | 1116 | Stage Axis Cycle Test with 0.5Kg Weight for 10 Mins
71 | 1111 | Confirm X Travel (0 - 125mm)
71 | 1111 | Confirm Y Travel (0 - 125mm)
71 | 1111 | Confirm Z Travel (0 - 50mm)
71 | 1111 | Check Tilt Travel -10° to +90°
71 | 1111 | Confirm X Travel (0 - 125mm)
71 | 1111 | Confirm Y Travel (0 - 125mm)
71 | 1111 | Confirm Z Travel (0 - 50mm)
71 | 1111 | Check Tilt Travel -10° to +90°
650 | 1124 | Quality Certificate Signed By:
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
708 | 11236 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed.
40 | 739 | Record Diode Serial Number:
40 | 739 | Mechanical Alignment
40 | 739 | Configure BSD in Administrator
40 | 739 | Input Offset Cal (Calibration Wizard)
40 | 739 | Input Offset Cal (ASAP Test Panel)
40 | 739 | Check BIAS operation
40 | 739 | Signal To Noise Test (Result >12%)
40 | 739 | Store Signal To Noise Result (Screen-shot)
40 | 739 | Set Up Diode Protection Plate
40 | 739 | Vac Degradation < 3x10-5 mbar
40 | 739 | Fit Detector Mechanics Cover
403 | 1134 | Print Shipment Labels
403 | 1134 | Label Material Boxes
403 | 1134 | Place left over stickes in Document Holder
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
708 | 11264 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be observed.
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
707 | 11166 | Measure aperture size, 55um (+/- 2um)
707 | 11166 | Anode_TS_55um - Passed inspection criteria, Image Saved
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
708 | 11236 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed.
24 | 273 | Level Frame
24 | 273 | Check pendulum height (with spacer)
24 | 273 | Set up ‘float’ height
24 | 273 | Check inlet / outlet valves
24 | 273 | Set up docking pins
82 | 1136 | (10kV, Central Aperture) - Aperture and Guns <50%
82 | 1136 | (10kV, Central Aperture) - Stigs <20%
82 | 1136 | (10kV, Outer Apertures) - Stigs <30%
82 | 1136 | (1kV, Central Aperture) - Aperture and Guns <50%
82 | 1136 | (1kV, Central Aperture) - Stigs <20%
82 | 1136 | (1kV, Outer Apertures) - Stigs <30%
82 | 1136 | <10% user align wobble required on all apertures
651 | 10598 | Check the Keyboard Language
2 | 7 | Set DIP switches
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
708 | 11264 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be observed.
708 | 11264 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be observed.
24 | 273 | Level Frame
24 | 273 | Check pendulum height (with spacer)
24 | 273 | Set up ‘float’ height
24 | 273 | Check inlet / outlet valves
24 | 273 | Set up docking pins
82 | 1593 | (10kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (10kV, Central Aperture) - Stigs <20%
82 | 1593 | (10kV, Outer Apertures) - Stigs <30%
82 | 1593 | (1kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (1kV, Central Aperture) - Stigs <20%
82 | 1593 | (1kV, Outer Apertures) - Stigs <30%
82 | 1593 | <10% user align wobble required on all apertures
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
71 | 1130 | Z Motor SeriaL Number:
71 | 1130 | T Motor Serial Number:
71 | 1130 | R Motor Serial Number:
102 | 667 | Confirm Plasma Cleaner is working by running a short sample clean
650 | 10566 | Safety Labels
650 | 10566 | Language
650 | 10566 | China RoHS Sticker
82 | 1593 | (10kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (10kV, Central Aperture) - Stigs <20%
82 | 1593 | (10kV, Outer Apertures) - Stigs <30%
82 | 1593 | (1kV, Central Aperture) - Aperture and Guns <50%
82 | 1593 | (1kV, Central Aperture) - Stigs <20%
82 | 1593 | (1kV, Outer Apertures) - Stigs <30%
82 | 1593 | <10% user align wobble required on all apertures
8 | 1645 | Serial:
46 | 887 | Remove docking brackets and fit transit brackets
46 | 887 | Vent the Chamber.
46 | 887 | Power system down.
46 | 887 | Blow water out of the cooling system.
46 | 887 | Remove On / Off switch panel
46 | 887 | Print and fit DJEL S/N Label
46 | 887 | Remove column cables only
46 | 887 | Remove column water pipes only
46 | 887 | Remove ASAP, EOS, EOD and Vacuum PCBS from plinth and pack seperately in static bag and bubble wrap.
46 | 887 | Remove AC unit complete with mains cable
46 | 887 | Remove Cekon Plug From Mains Cable
46 | 887 | Remove PSU unit
46 | 887 | Remove EHT Set
46 | 887 | Remove IGP PSU
46 | 887 | Remove EBBU Electonics from column and fit shorting plug
46 | 887 | Remove customer SE2 detector (460F only)
46 | 887 | Fit customer SE2 Detectors (2 off) to transport vessels (460 F only).
46 | 887 | Remove BSD4 complete from chamber and fit protective cover over Diode (460F only).
46 | 887 | Pack customer accessories as shown WI (460F Only)
46 | 887 | Remove Water flow and Temperature sensor assembly
46 | 887 | Remove column / support frame complete and attach to column storage frame
46 | 887 | Fit blank plate to column bypass pipe so that lower column can be shipped at vacuum (460F Only).
46 | 887 | Pump down lower column as shown in WI.
46 | 887 | Chamber secured to vibration plate
46 | 887 | Chamber Securing Screws Tightened to 25NM
46 | 887 | Secure the pendulum feet using foam blocks and strap - Only required if test tool is being shipped to Bollore
46 | 887 | Print shipping stickers.
71 | 1130 | Z Motor SeriaL Number:
71 | 1130 | T Motor Serial Number:
71 | 1130 | R Motor Serial Number:
82 | 973 | (10kV, Central Aperture) - Aperture and Guns <50%
82 | 973 | (10kV, Central Aperture) - Stigs <35%
82 | 973 | (10kV, Outer Apertures) - Stigs <45%
82 | 973 | (1kV, Central Aperture) - Aperture and Guns <50%
82 | 973 | (1kV, Central Aperture) - Stigs <20%
82 | 973 | (1kV, Outer Apertures) - Stigs <30%
82 | 973 | <10% user align wobble required on all apertures
82 | 1593 | (10kV, Central Aperture) - Aperture and Guns <60%
82 | 1593 | (10kV, Central Aperture) - Stigs <20%
82 | 1593 | (10kV, Outer Apertures) - Stigs <30%
82 | 1593 | (1kV, Central Aperture) - Aperture and Guns <60%
82 | 1593 | (1kV, Central Aperture) - Stigs <20%
82 | 1593 | (1kV, Outer Apertures) - Stigs <30%
82 | 1593 | <10% user align wobble required on all apertures
25 | 198 | Gun X
25 | 198 | Gun Y
25 | 198 | Aper X
25 | 198 | Aper Y
285 | 1135 | Print shipment lables
285 | 1135 | Lable material boxes
285 | 1135 | Place left over stickers in document holder
285 | 1135 | Print shipment labels
285 | 1135 | Label material boxes
285 | 1135 | Place left over stickers in document holder
45 | 1137 | Confirm Predictive Service is Disabled
20 | 11270 | Predictive Service S/W installed on Customer PC.
20 | 11270 | Predictive Service has been activated
20 | 972 | Service centre S/W installed on Customer PC.
22 | 1138 | The 2nd Inlens detector has been configured as an ESB
71 | 1130 | Z Motor Serial Number:
71 | 1130 | T Motor Serial Number:
71 | 1130 | R Motor Serial Number:
22 | 1138 | The 2nd Inlens detector has been configured as an ESB
70 | 1073 | Checked Carried Our By:
389 | 950 | Fit the Insitu substage
389 | 950 | Test the Insitu substage
389 | 950 | System vac <2.0e-06mbar
389 | 950 | X-ray test
389 | 950 | Repackage all hardware excluding the chamber feed through in Deben packaging
709 | 11322 | Weight
709 | 11322 | Dimensions Length
709 | 11322 | Dimensions Width
709 | 11322 | Dimensions Height
709 | 11322 | Security Seals Serial No
709 | 11322 | Shock Watch Serial Number
709 | 11322 | Tilt Watch Serial No.
709 | 11322 | Packed By 1
709 | 11322 | Packed By 2
709 | 11322 | Packed By 3
709 | 11322 | Weight
709 | 11322 | Dimensions Length
709 | 11322 | Dimensions Width
709 | 11322 | Dimensions Height
709 | 11322 | Security Seals Serial No
709 | 11322 | Shock Watch Serial Number
709 | 11322 | Tilt Watch Serial No.
709 | 11322 | Packed By 1
709 | 11322 | Packed By 2
709 | 11322 | Packed By 3
651 | 10599 | Check Keyboard Language
651 | 10599 | Check the operation of the control panel and the Keyboard
71 | 1130 | Z Motor Serial Number:
71 | 1130 | T Motor Serial Number:
71 | 1130 | R Motor Serial Number:
102 | 1183 | Check SmartStitch can be run
102 | 1183 | SmartStitch folder is present on the Desktop
43 | 346 | 1kV =
43 | 346 | Store 1KV Resolution Images and Pate image
43 | 346 | 15kV =
43 | 346 | Store 15KV Resolution Images and Pate image
28 | 291 | If applicable confirm 'Pressure Stage' is ticked in Administrator
28 | 291 | Mechanical Alignment ok
28 | 291 | Column clamping secure
70 | 1079 | Z Microswitch Position Set
70 | 1079 | T Microswitch Position Set
70 | 1079 | R Microswitch Position Set
70 | 1082 | Checked Carried Our By:
40 | 1558 | CCD
40 | 1558 | Focus
709 | 11322 | Weight
709 | 11322 | Dimensions Length
709 | 11322 | Dimensions Width
709 | 11322 | Dimensions Height
709 | 11322 | Security Seals Serial No
709 | 11322 | Shock Watch Serial Number
709 | 11322 | Tilt Watch Serial No.
709 | 11322 | Packed By 1
709 | 11322 | Packed By 2
709 | 11322 | Packed By 3
25 | 197 | Gun X
25 | 197 | Gun Y
25 | 197 | Aper X
25 | 197 | Aper Y
24 | 273 | Level Frame
24 | 273 | Check pendulum height (with spacer)
24 | 273 | Set up ‘float’ height
24 | 273 | Check inlet / outlet valves
24 | 273 | Set up docking pins
645 | 10513 | 30KV @ 5mm X- ≤35%
645 | 10513 | 30KV @ 5mm Y- ≤35%
71 | 1130 | Z Motor Serial Number:
71 | 1130 | T Motor Serial Number:
71 | 1130 | R Motor Serial Number:
71 | 1107 | Checked Carried Our By:
71 | 1109 | Check Set Point X Axis Micro Switch
71 | 1109 | Check Set Point Y Axis Micro Switch
71 | 1109 | Check Set Point Z Axis Micro Switch
71 | 1109 | Check Set point T Axis (-10.8º)
71 | 1110 | Check Stage Initialises Correctly
71 | 1111 | Confirm X Travel (0 - 125mm)
71 | 1111 | Confirm Y Travel (0 - 125mm)
71 | 1111 | Confirm Z Travel (0 - 50mm)
71 | 1111 | Check Tilt Travel -10° to +90°
71 | 1116 | Run Axis Cycle Test (min - 1Hr)
71 | 1116 | Stage Axis Cycle Test with 0.5Kg Weight for 10 Mins
71 | 1118 | Stage Kit
71 | 1118 | 27mm Spacer Block Fitted
71 | 1121 | Stage Serial Number:
71 | 1122 | Quality Sign Off Carried Out By:
389 | 11074 | Check each individual segment in normal and inverse modes
389 | 11074 | Analyze Segments
389 | 11074 | Gain Ranges
389 | 11074 | Pneumatics
389 | 11074 | Alignment
389 | 11074 | Vac Degradation < 3x10-5 mbar
389 | 11074 | CZ Detector Cal
389 | 11074 | Z Adjustment
389 | 11074 | Check each individual segment in normal and inverse modes
389 | 11074 | Analyze Segments
389 | 11074 | Gain Ranges
389 | 11074 | Pneumatics
389 | 11074 | Alignment
389 | 11074 | Vac Degradation < 3x10-5 mbar
389 | 11074 | CZ Detector Cal
389 | 11074 | Z Adjustment
709 | 11322 | Weight
709 | 11322 | Dimensions Length
709 | 11322 | Dimensions Width
709 | 11322 | Dimensions Height
709 | 11322 | Security Seals Serial No
709 | 11322 | Shock Watch Serial Number
709 | 11322 | Tilt Watch Serial No.
709 | 11322 | Packed By
70 | 1095 | Checked Carried Our By:
24 | 273 | Level Frame
24 | 273 | Check pendulum height (with spacer)
24 | 273 | Set up ‘float’ height
24 | 273 | Check inlet / outlet valves
24 | 273 | Set up docking pins
44 | 348 | Chamber Door Interlock
44 | 348 | Service Switch
44 | 348 | Lens De-activation
44 | 348 | Detector De-activation
44 | 348 | EHT + IPPSU De-activation
44 | 454 | Interlock tested and working
645 | 10524 | Pumpdown Time
44 | 348 | Chamber Door Interlock
44 | 348 | Service Switch
44 | 348 | Lens De-activation
44 | 348 | Detector De-activation
44 | 348 | EHT + IPPSU De-activation
646 | 10526 | Number of Chamberscope's on Order
646 | 10526 | Chamberscope 1 - Select Port
646 | 10526 | Check Camera for focus
646 | 10526 | Functional Check-Black/White + Colour Mode
646 | 10526 | Run wxPropview to Remove Defective Pixels
646 | 10526 | Chamberscope 2 - Select Port or Not Fitted
646 | 10526 | Chamberscope 2 - Check Camera 2 for Focus
646 | 10526 | Chamberscope 2 - Functional Check-Black/White + Colour Mode
646 | 10526 | Chamberscope 2 - Run wxPropview to Remove Defective Pixels
646 | 10526 | Chamberscope 2 - Add Icon to Toolbar
676 | 1145 | Install Airlock FW Version 2.0.0
40 | 916 | Unpack Volutome package and check it's completeness
40 | 916 | Ensure Arivis Software Licence is available (410145-1031-100)
40 | 916 | Remove transit clamps from Volutome stage
40 | 916 | Ensure hardware has been configured according to work instructions
40 | 916 | Fit Volutome stage
40 | 916 | Connect Volutome stage and controller to the SEM
40 | 916 | Install Volutome software installation package
40 | 916 | Configure Volutome stage motors
40 | 916 | Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator
40 | 916 | Set up the Volutome workbench in Zen
40 | 916 | Record VolumeBSD serial number
40 | 916 | Record Volutome stage serial number
40 | 916 | Record VolumeBSD amplifier serial number
40 | 916 | Record VolumeBSD mechanics serial number
40 | 916 | Record C-885 PI controller serial number
40 | 916 | Record E-727 PI controller serial number
40 | 916 | Function check the Volutome stage light
40 | 916 | Adjust Volutome stage centre (Crosshairs should be within 4 central squares)
40 | 916 | Install Nav cam in-line connections (If applicable)
40 | 916 | Perform the microtome and stage functionality checks
40 | 916 | Fit and align the VolumeBSD diode
40 | 916 | Check VolumeBSD alignment repeatability
40 | 916 | Check all gain ranges of the VolumeBSD (normal/inverted)
40 | 916 | Perform Signal/noise ratio check
40 | 916 | Ensure BSD diode protector is fitted and adjusted in chamber
40 | 916 | Install cable tie to loop BSD signal cable
40 | 916 | Check SCM with Volutome stage
40 | 916 | Ripple X Spec 2nm (600Kx Mag 15KV WD2)
40 | 916 | Ripple Y Spec 2nm (600Kx Mag 15KV WD2)
40 | 916 | Resolution 1KV Spec same as standard resolution
40 | 916 | Resolution 15KV Spec same as standard resolution
40 | 916 | Contamination
40 | 916 | Perform X-ray check
40 | 916 | Pump Down Time
40 | 916 | Ultimate Vacuum
40 | 916 | Focal CC fitted to Volutome stage
40 | 916 | Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar
40 | 916 | Generate live image from Axiocam
40 | 916 | Re-package Volutome and check it's completeness
40 | 916 | ManEng/ Senior Eng/ Production Manger Sign off
46 | 11123 | Remove docking brackets and fit transit brackets
46 | 11123 | Power system down.
46 | 11123 | Blow water out of the cooling system.
46 | 11123 | Put thread lock on pendulum screws and transit block Screws.
46 | 11123 | Chamber secured to vibration plate (N/A to XB 550L)
46 | 11123 | Chamber securing screws tightened to 25nm (N/A to XB 550L)
46 | 11123 | Chamber securing screws tightness checked. Type your initials (For XB 550L type N/A)
46 | 11123 | Secure Airfeet with foam blocks.
46 | 11123 | Secure Turbo pump with foam blocks.
46 | 11123 | Fit Covers to all unused connectors.
46 | 11123 | Mains and Earth cable stowed on Mains unit.
46 | 11123 | Fit shipping Eyes to transit blocks.
46 | 11123 | All doors closed and secured.
46 | 11123 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11123 | Print shipping stickers.
46 | 11123 | Fit 8x M6 x 55mm cap head screws to stage door
37 | 329 | Stig: 30μm, All EHT’s is ≤ 20%
37 | 329 | Stig: Off Axis Aper, All EHT’s is ≤ 30%
37 | 329 | EHT <1kv: 30μm Aperture Gun/Aperture Aligns ≤ 60%
37 | 329 | EHT ≥1kv: 30μm Aperture Gun/Aperture Aligns ≤ 50%
648 | 10553 | 30KV Res- ≤2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 1KV Res- ≤15nm
648 | 10553 | 200V Res- Not measured
648 | 10553 | 1KV Oscillation Check - EHT=1KV, Scan Speed=2, WD=10mmm, Mag=10Kx
648 | 10553 | Store Standard Resolution Images + Oscillation Check
646 | 10527 | Basic Movements-T/Z Joystick
646 | 10527 | Basic Movements-X/Y/R Joystick
646 | 10527 | Initialise Stage
648 | 10875 | 2KV Res - 750um Easy Align Aperture- ≤16nm
648 | 10875 | Store Easy Align Aperture (750u) Resolution Images
650 | 11097 | Check SCS Folder/Sales Order
650 | 10570 | Set Language in User Preferences
37 | 329 | Stig: 30μm, All EHT’s is ≤ 20%
37 | 329 | Stig: Off Axis Aper, All EHT’s is ≤ 30%
37 | 329 | Aper/Gun: 30μm, All EHT’s is ≤ 60%
2 | 8 | Set DIP switches
22 | 264 | Tip No.
22 | 264 | Confirm that the Tip S/N on the Denka datasheet is the same as the Tip S/N on the column sticker.
30 | 456 | Eucentricity set up
30 | 456 | Focus Deviation (should be less than 20um)
42 | 344 | Actual X-Axis
42 | 344 | Store X ripple shots and results
42 | 344 | Actual Y-Axis
42 | 344 | Store Y ripple shots and results
41 | 10964 | Column Mechanical Alignment is good
41 | 10964 | MU Metal Shield fitted
40 | 1146 | Stage mapping completed
40 | 1146 | Kompass passed - spec 0.5um
650 | 914 | Check Dongle fits in USB HUB (“ICY BOX”) and replace if needed.
26 | 1141 | ESB Serial Number.
26 | 1141 | InLens Detector Concentricity is acceptable
26 | 1141 | InLens Detector Concentricity is acceptable
646 | 10528 | Installation
646 | 10528 | select Language
646 | 10528 | Function of all knobs and buttons
646 | 10528 | Function of all keyboard keys
648 | 10553 | 30KV Res- ≤2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 1KV Res- ≤15nm
648 | 10553 | 200V Res- Not measured
648 | 10553 | 1KV Oscillation Check - EHT=1KV, Scan Speed=2, WD=10mmm, Mag=10Kx
648 | 10553 | Store Standard Resolution Images + Oscillation Check
650 | 10561 | Complete & Sign Test Certificate
650 | 10561 | Complete & Sign Test Certificate
650 | 10566 | Safety Labels
650 | 10566 | Language
650 | 10566 | China RoHS Sticker
26 | 1141 | InLens Detector Concentricity is acceptable
650 | 971 | Clean Front Panel
650 | 971 | Prepare Logo
650 | 971 | Fitting Logo
40 | 1146 | Stage mapping completed
40 | 1146 | Kompass passed - spec 0.5um
40 | 1146 | Stage mapping completed
40 | 1146 | Kompass passed - spec 0.5um
21 | 259 | Check camera Focus and Rotation
21 | 259 | Eliminate camera image tilt
21 | 259 | Ty=wrap and Label Cables
403 | 960 | Please print 'Declaration of Conformity' from URL link
50 | 480 | Mu metal shield is fitted
50 | 480 | DC Tablefiller successfully completed
665 | 10737 | Diode Holder Assembly aligned using jig.
665 | 10737 | All screw fully tightened
665 | 10737 | Check alignment is Ok again using jig
26 | 281 | Type in measured value (Limit <= 30um)
26 | 281 | Store ESB Concentricity Image
671 | 10799 | Internal Surfaces
671 | 10799 | Tapped Holes
40 | 11081 | Fit Test tool BB Electronics
40 | 11081 | Function Test
40 | 11081 | Remove Test tool BB Electronics
650 | 10571 | Set Language in User Preferences
650 | 10572 | Select Keyboard Language
650 | 10572 | Set Keyboard Language in Windows Control Panel
650 | 10572 | Destination of System
650 | 1124 | Quality Certificate Signed By:
651 | 10580 | SmartSEM Version
651 | 10587 | Initialise Stage
651 | 10587 | Stage Movements operate Correctly
651 | 10588 | Collector Volts
651 | 10588 | Scintillator
651 | 10588 | With the beam blanked, select SE Detector, set contrast and brightness to 50%, scan speed 3 and ensure that there are no white streaks across the screen
651 | 10589 | Apertures in the Correct Position
651 | 10589 | Both apertures can be zero wobbled
651 | 10599 | Check Keyboard Language
651 | 10599 | Check the operation of the control panel and the Keyboard
651 | 915 | Check that the 200V image is of a good quality.
651 | 915 | Check 1KV Resolution Image meets specification and of a good quality.
651 | 915 | Check 3KV Resolution Image meets specification and of a good quality.
651 | 915 | Check 30KV Resolution Image meets specification and of a good quality.
651 | 10603 | X Stigmator
651 | 10603 | Y Stigmator
651 | 10603 | X Gun Tilt ≤50%
651 | 10603 | Y Gun Tilt ≤50%
651 | 10603 | X Gun Shift ≤50%
651 | 10603 | Y Gun Shift ≤50%
651 | 11394 | X Stigmator
651 | 11394 | Y Stigmator
651 | 11394 | X Gun Tilt ≤50%
651 | 11394 | Y Gun Tilt ≤50%
651 | 11394 | X Gun Shift ≤50%
651 | 11394 | Y Gun Shift ≤50%
388 | 3291 | X =
388 | 3291 | Y =
388 | 11425 | 15kv 0.9nm
388 | 11425 | 15KV result
388 | 11425 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 1kv 1.3nm
388 | 11425 | 1KV result
388 | 11425 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 500V 2nm
388 | 11425 | 500V result
388 | 11425 | Store 500V Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 15kv 0.7nm
388 | 11426 | 15KV result
388 | 11426 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 1kv 1.2nm
388 | 11426 | 1KV result
388 | 11426 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 500V 1.9nm
388 | 11426 | 500V result
388 | 11426 | Store 500V Resolution Images and Pate image for 2 x consecutive images
26 | 281 | Type in measured value (Limit <= 30um)
26 | 281 | Store ESB Concentricity Image
671 | 10799 | Internal Surfaces
671 | 10799 | Tapped Holes
667 | 1147 | Stage Serial Number
667 | 1147 | Shim Stage Door
667 | 1147 | Set Stage Clamp Height
5 | 1161 | Does system have an SCS? please select
5 | 1161 | If yes, Product Manager to enter initials. If no please type N/A:
651 | 669 | The Licence file has been copied to the SmartSEM USB Stick
651 | 669 | The EM Server.db has been copied to the SmartSEM USB Stick
651 | 10608 | System Vacuum Reading (≤3.00e-06 mbar)
654 | 11197 | Remove Files
654 | 11197 | Remove Folders
654 | 11420 | PAT Tester Serial Number
654 | 11420 | PAT Tester Cal Expiry Date
654 | 11420 | Earth resistance – measured value must be < 0.1 Ω
654 | 11420 | Insulation resistance – measured value must be > 1.0M Ω
654 | 11420 | Test Certificate Completed & Signed
645 | 10880 | 30KV @ 5mm X- ≤35%
645 | 10880 | 30KV @ 5mm Y- ≤35%
645 | 10761 | Functional Check
654 | 10638 | Special Fitted
654 | 10638 | If Fitted, Checked by SCS Team Member
645 | 10762 | Easy VP Mag. Cal Complete
32 | 311 | Limit is 600nm .......................Actual =
32 | 311 | Store objective drift results
44 | 355 | Ensure the S/N & P/N match the License generator report / Sales order.
384 | 1143 | Stigmators Balanced On All Apertures (Normal Current)
384 | 1143 | Stigmators Balanced On Aperture 5 (High Current ) - X Upper Value
384 | 1143 | Stigmators Balanced On Aperture 5 (High Current ) - X Lower Value
384 | 1143 | Stigmators Balanced On Aperture 5 (High Current ) - Y Upper Value
384 | 1143 | Stigmators Balanced On Aperture 5 (High Current ) - Y Lower Value
384 | 1143 | Stigmators Balanced On Aperture 6 (High Current ) - X Upper Value
384 | 1143 | Stigmators Balanced On Aperture 6 (High Current ) - X Lower Value
384 | 1143 | Stigmators Balanced On Aperture 6 (High Current ) - Y Upper Value
384 | 1143 | Stigmators Balanced On Aperture 6 (High Current ) - Y Lower Value
384 | 1143 | Stigmators Balanced On Aperture 7 (High Current ) - X Upper Value
384 | 1143 | Stigmators Balanced On Aperture 7 (High Current ) - X Lower Value
384 | 1143 | Stigmators Balanced On Aperture 7 (High Current ) - Y Upper Value
384 | 1143 | Stigmators Balanced On Aperture 7 (High Current ) - Y Lower Value
384 | 1143 | Stigmators Balanced On The Remaining Apertures (High Current )
8 | 1646 | Serial:
671 | 10798 | Internal Surfaces
671 | 10798 | Tapped Holes
717 | 11386 | Time =
717 | 11386 | Date =
2 | 13 | Confirm EO PCB PN is '346001-9100-400'
2 | 13 | Confirm EO PCB PN is '346001-9110-315'
403 | 960 | Please print 'Declaration of Conformity' from URL link
384 | 1143 | Stigmators Balanced On All Apertures (Normal Current)
384 | 1143 | Stigmatoes Balanced On All Apertures (High Current )
384 | 1143 | Stigmators Balanced On All Apertures (Normal Current)
384 | 1143 | Stigmators Balanced On All Apertures (High Current )
384 | 1143 | Stigmators Balanced On All Apertures (Normal Current)
384 | 1143 | Stigmators Balanced On Aperture 1 (High Current ) - X Upper Value
384 | 1143 | Stigmators Balanced On Aperture 1 (High Current ) - X Lower Value
384 | 1143 | Stigmators Balanced On All Apertures (Normal Current)
384 | 1143 | Stigmators Balanced On Aperture 1 (High Current ) - X Upper Value
384 | 1143 | Stigmators Balanced On Aperture 1 (High Current ) - X Lower Value
384 | 1143 | Stigmators Balanced On Aperture 1 (High Current ) - Y Upper Value
384 | 1143 | Stigmators Balanced On Aperture 1 (High Current ) - Y Lower Value
384 | 1143 | Stigmators Balanced On Aperture 4 (High Current ) - X Upper Value
384 | 1143 | Stigmators Balanced On Aperture 4 (High Current ) - X Lower Value
384 | 1143 | Stigmators Balanced On Aperture 4 (High Current ) - Y Upper Value
384 | 1143 | Stigmators Balanced On Aperture 4 (High Current ) - Y Lower Value
384 | 1143 | Stigmators Balanced On Aperture 5 (High Current ) - X Upper Value
384 | 1143 | Stigmators Balanced On Aperture 5 (High Current ) - X Lower Value
384 | 1143 | Stigmators Balanced On Aperture 5 (High Current ) - Y Upper Value
384 | 1143 | Stigmators Balanced On Aperture 5 (High Current ) - Y Lower Value
384 | 1143 | Stigmators Balanced On The Remaining Apertures (High Current )
20 | 972 | Service centre S/W installed on Customer PC.
20 | 972 | Confirm Data Logging is running and External Logging is enabled.
26 | 10965 | ESB Serial Number.
2 | 1153 | Confirm EO PCB PN is '346001-9110-315'
2 | 12 | Serial No.
2 | 12 | Version
2 | 20 | Install Compressed Air Line
2 | 20 | Set foot height
2 | 21 | Set Valve Screw torque IN (1 turn)
2 | 21 | Set Valve Screw torque OUT (1 turn)
2 | 21 | Set foot height
2 | 21 | Install Compressed Air Line
2 | 21 | Install Control Cable
2 | 22 | Set Valve Screw torque IN (.5 turn)
2 | 22 | Set Valve Screw torque OUT(1 turn)
2 | 22 | Set foot height
2 | 22 | Install Compressed Air Line
2 | 22 | Install Control Cable
2 | 23 | Set Valve Screw torque IN (.5 turn)
2 | 23 | Set Valve Screw torque OUT(1 turn)
2 | 23 | Set foot height
2 | 23 | Install Compressed Air Line
2 | 23 | Install Control Cable
4 | 37 | Shim Stage
4 | 37 | Set Mag Clamp
4 | 37 | Fit Gun Interlock Key to Stage Door
4 | 39 | Turbo Pump S/No.
676 | 1162 | Does system have an SCS? please select
676 | 1162 | If yes, Product Manager to enter initials. If no please type N/A:
26 | 281 | Type in measured value (Limit <= 30um)
26 | 281 | Store ESB Concentricity Image
24 | 273 | Level Frame
24 | 273 | Check pendulum height (with spacer)
24 | 273 | Set up ‘float’ height
24 | 273 | Check inlet / outlet valves
24 | 273 | Set up docking pins
2 | 1153 | Confirm EO PCB PN is '346001-9110-305'
2 | 14 | Confirm EO Static PCB PN is '346001-9100-410'
2 | 14 | Fit EO Static PCB
20 | 972 | Service centre S/W installed on Customer PC.
20 | 972 | Confirm Data Logging is running and External Logging is enabled.
646 | 10546 | Chamberscope Fitted to Stage Door
646 | 10546 | SmartTouch Software Version
646 | 10546 | Install SmartTouch Software
646 | 10546 | Install Easy VP Aperture
646 | 10546 | Capture Sample Image Using the Nav Cam
646 | 10546 | Select Sample Holder and Sample Type
646 | 10546 | Set Yellow Line To 8.5mm WD
646 | 10546 | Sample Setup Complete
646 | 10546 | In the Image Tab, Check Sliders Function
646 | 10546 | Save Image - Photomode 1, Scan speed 8
646 | 10546 | Save Image - Photomode 2, Single Image
646 | 10546 | Save Image - Photomode 3, Scan Speed 2, Multiple Frames
646 | 10546 | Check Auto Saturation (W Filament Only)
646 | 10546 | Check Auto Stigmation
646 | 10546 | Check Manual Stigmation
646 | 10546 | Check Annotate
646 | 10546 | Check All Detectors Are Presaent
646 | 10546 | Select the Automate tab & Create Folder as Factory
646 | 10546 | Setup 3 Point Calibration (Nav Cam )
646 | 10546 | Automation Run Completed
646 | 10546 | Select Browse Tab and Check Functions
26 | 282 | Type in measured value (limit: <= 15um)
26 | 282 | Store Inlens to ESB image
26 | 283 | Check for ESB charging and Flashing.
26 | 281 | Type in measured value (Limit <= 30um)
26 | 281 | Store ESB Concentricity Image
23 | 269 | Run ESB conditioning MACRO prior to switching on EHT.
29 | 11381 | Run overnight ESB Run Up & Photograph Macro
29 | 11381 | Analyse images for drift, breakdown, ESB issues
40 | 1551 | Grid working
40 | 1551 | Check Cut Off
102 | 1616 | Grid Control working
102 | 1616 | ESB Concentricity in Spec
28 | 293 | ESB concentricity in spec
28 | 293 | ESB-Inlens concentricity in spec
44 | 706 | Confirm ESB SN matches
2 | 20 | Install Compressed Air Line
2 | 20 | Set foot height
666 | 1163 | Does system have an SCS? please select
666 | 1163 | If yes, Product Manager to enter initials. If no please type N/A:
384 | 1143 | Stigmators Balanced On All Apertures (Normal Current)
384 | 1143 | Stigmators Balanced On Aperture 5 (High Current ) - X Upper Value
384 | 1143 | Stigmators Balanced On Aperture 5 (High Current ) - X Lower Value
384 | 1143 | Stigmators Balanced On Aperture 5 (High Current ) - Y Upper Value
384 | 1143 | Stigmators Balanced On Aperture 5 (High Current ) - Y Lower Value
384 | 1143 | Stigmators Balanced On Aperture 6 (High Current ) - X Upper Value
384 | 1143 | Stigmators Balanced On Aperture 6 (High Current ) - X Lower Value
384 | 1143 | Stigmators Balanced On Aperture 6 (High Current ) - Y Upper Value
384 | 1143 | Stigmators Balanced On Aperture 6 (High Current ) - Y Lower Value
384 | 1143 | Stigmators Balanced On Aperture 7 (High Current ) - X Upper Value
384 | 1143 | Stigmators Balanced On Aperture 7 (High Current ) - X Lower Value
384 | 1143 | Stigmators Balanced On Aperture 7 (High Current ) - Y Upper Value
384 | 1143 | Stigmators Balanced On Aperture 7 (High Current ) - Y Lower Value
384 | 1143 | Stigmators Balanced On The Remaining Apertures (High Current )
2 | 1153 | Confirm EO PCB PN is '346001-9110-305'
2 | 13 | Confirm EO PCB PN is '346001-9110-315'
2 | 1153 | Confirm EO PCB PN is '346001-9110-305'
2 | 13 | Confirm EO Dynamic PCB PN is '346001-9110-315'
2 | 14 | Confirm EO Static PCB PN is '346001-9100-410'
2 | 1154 | Confirm EO Static PCB PN is '346001-9100-400'
2 | 13 | Confirm EO Dynamic PCB PN is '346001-9110-315'
2 | 13 | Fit EO Dynamic PCB
4 | 46 | Carboflow Holes
4 | 46 | Carboflow Chamber
4 | 46 | IPA Holes
4 | 46 | IPA Chamber
4 | 47 | Time =
4 | 47 | Date =
4 | 48 | Time =
4 | 48 | Date =
393 | 3346 | Ensure all customer licences are loaded in the Zeiss Licence Management Tool and functions correctly
5 | 58 | Configuration Check
5 | 58 | Port Check
6 | 60 | No Pressure loss in 2mins
6 | 61 | Safety Earth
6 | 61 | Compressed Air (Set to 0.5MPa)
6 | 61 | PVP Valve
6 | 61 | Power
6 | 61 | Rotary Pump
6 | 61 | Nitrogen (Set to 0.05 MPa)
6 | 61 | Cooling Water
6 | 62 | Firmware (Mini Mon)
6 | 62 | PLDs (Xlinks)
6 | 63 | System Test
6 | 63 | Ion Pump ON
6 | 64 | Start SmartSEM and confirm there are no errors in EM server.
6 | 65 | Set Stage links
6 | 65 | Initialisation Check
6 | 65 | Door alignment Check
6 | 65 | Magnet Setup Check
6 | 65 | Check all Interlock Keys are fitted
6 | 66 | Ensure Chamber Holes are Clean
6 | 68 | Full System Leak Check
6 | 69 | Foot Height
6 | 69 | Pipe Configuration
6 | 69 | Foot Nut Orientation Correct
6 | 69 | Sealing Plugs Removed
6 | 70 | Camera Cables Routing outside Plinth
6 | 71 | IPG Cables No.
6 | 1545 | Complete HiPot Result Sheet
6 | 1546 | Complete HiPot Result Sheet
6 | 1547 | Complete PAT Result Sheet
6 | 1548 | Complete PAT Result Sheet
2 | 7 | Set DIP switches
2 | 8 | Set DIP switches
2 | 7 | Set DIP switches
2 | 8 | Set DIP switches
2 | 12 | Serial No.
2 | 12 | Version
2 | 20 | Install Compressed Air Line
2 | 20 | Set foot height
2 | 21 | Set Valve Screw torque IN (1 turn)
2 | 21 | Set Valve Screw torque OUT (1 turn)
2 | 21 | Set foot height
2 | 21 | Install Compressed Air Line
2 | 21 | Install Control Cable
2 | 22 | Set Valve Screw torque IN (.5 turn)
2 | 22 | Set Valve Screw torque OUT(1 turn)
2 | 22 | Set foot height
2 | 22 | Install Compressed Air Line
2 | 22 | Install Control Cable
2 | 23 | Set Valve Screw torque IN (.5 turn)
2 | 23 | Set Valve Screw torque OUT(1 turn)
2 | 23 | Set foot height
2 | 23 | Install Compressed Air Line
2 | 23 | Install Control Cable
3 | 30 | Fit EHT Cables
3 | 30 | Torque EHT Connections to 3Nm
4 | 37 | Shim Stage
4 | 37 | Set Mag Clamp
4 | 37 | Fit Gun Interlock Key to Stage Door
4 | 39 | Turbo Pump S/No.
4 | 46 | Carboflow Holes
4 | 46 | Carboflow Chamber
4 | 46 | IPA Holes
4 | 46 | IPA Chamber
4 | 10937
6 | 60 | No Pressure loss in 2mins
6 | 61 | Safety Earth
6 | 61 | Compressed Air (Set to 0.5MPa)
6 | 61 | PVP Valve
6 | 61 | Power
6 | 61 | Rotary Pump
6 | 61 | Nitrogen (Set to 0.05 MPa)
6 | 61 | Cooling Water
6 | 62 | Firmware (Mini Mon)
6 | 62 | PLDs (Xlinks)
5 | 58 | Configuration Check
5 | 58 | Port Check
6 | 63 | System Test
6 | 63 | Ion Pump ON
6 | 64 | Start SmartSEM and confirm there are no errors in EM server.
6 | 65 | Set Stage links
6 | 65 | Initialisation Check
6 | 65 | Door alignment Check
6 | 65 | Magnet Setup Check
6 | 65 | Check all Interlock Keys are fitted
6 | 66 | Ensure Chamber Holes are Clean
6 | 1545 | Complete HiPot Result Sheet
6 | 1546 | Complete HiPot Result Sheet
6 | 1547 | Complete PAT Result Sheet
3 | 30 | Fit EHT Cables
3 | 30 | Torque EHT Connections to 3Nm
4 | 10937 | Screws fitted by
5 | 49 | Check Plinth Paperwork is complete for Takt 3
5 | 50 | Fit Camera
6 | 72 | Routed though Plinth secured on top of Plinth
702 | 11014 | Time =
702 | 11014 | Date =
702 | 11015 | Time =
702 | 11015 | Date =
4 | 10937 | Screws fitted by
703 | 11016 | Time =
703 | 11016 | Date =
703 | 11017 | Time =
703 | 11017 | Date =
6 | 65 | Set Stage links
6 | 65 | Initialisation Check
6 | 65 | Door alignment Check
6 | 65 | Magnet Setup Check
6 | 65 | Check all Interlock Keys are fitted
2 | 11253 | Check SCS Folder/Sales Order
2 | 11253 | SCS Number (If applicable)
6 | 11251 | Print Plinth Seperators
6 | 11251 | Sort Documentation with Seperators
6 | 11254 | Check SCS Folder/Sales Order
6 | 11254 | SCS Number (If applicable)
8 | 96 | Pneumatic Operation
8 | 96 | Pin Touching when Retracted
9 | 97 | Current Drawn:
2 | 11309 | Correct damping system installed
2 | 11309 | Correct version damping system supplied
6 | 1545 | Complete HiPot Result Sheet
6 | 1547 | Complete PAT Result Sheet
7 | 73 | Objective Lens Coil Type
7 | 74 | Visual Inspection
7 | 74 | Liner Fitment
7 | 74 | End Cap Fitment
7 | 75 | Serial No.
7 | 76 | Open Circuit (Coil – GND)
7 | 77 | 12 hours drying time
7 | 81 | Resistance:
7 | 82 | Pressure 15 PSI for 5 minutes
7 | 83 | Open Circuit (Coil – GND)
7 | 1638 | Serial No.
7 | 1639 | Position A:
7 | 1639 | Position B:
7 | 1639 | Position C:
7 | 1639 | Position D:
7 | 1639 | Max Difference:
7 | 1640 | Position A:
7 | 1640 | Position B:
7 | 1640 | Position C:
7 | 1640 | Position D:
7 | 1640 | Max Difference:
7 | 1641 | Stand Off 1:
7 | 1641 | Stand Off 2:
7 | 1641 | Stand Off 3:
7 | 1641 | Stand Off 4:
7 | 1641 | Max Difference:
7 | 1642 | Objective Coil
7 | 1642 | Scan Coil
8 | 91 | Housing Type
8 | 92 | Cooling Isolation Check (Open Circuit Coil – GND)
8 | 94 | Serial:
8 | 95 | Operational repeatability test x 3
10 | 106 | Position A:
10 | 106 | Position B:
10 | 106 | Position C:
10 | 106 | Position D:
10 | 106 | Max. Difference
10 | 109 | Continuity (Close Circuit)
10 | 109 | Coils to GND (Open Circuit)
10 | 110 | Continuity (Close Circuit)
10 | 110 | Coils to GND (Open Circuit)
10 | 111 | Continuity (Close Circuit)
10 | 111 | Coils to GND (Open Circuit)
10 | 112 | Continuity (Close Circuit)
10 | 112 | Coils to GND (Open Circuit)
11 | 117 | Drying time of 12hrs.
11 | 118 | C1 Coil
11 | 118 | C2 Coil
11 | 119 | C1 Coil
11 | 119 | C2 Coil
11 | 125 | Inspect
11 | 125 | Clean
11 | 128 | Continuity (Close Circuit)
11 | 128 | Coils to GND (Open Circuit)
11 | 129 | Continuity (Close Circuit)
11 | 129 | Coils to GND (Open Circuit)
11 | 130 | Continuity (Close Circuit)
11 | 130 | Coils to GND (Open Circuit)
11 | 131 | Continuity (Close Circuit)
11 | 131 | Coils to GND (Open Circuit)
11 | 133 | O-Rings TEM oiled
11 | 136 | A - B (< 100kΩ)
11 | 136 | C - D (< 10Ω)
11 | 136 | B - D (Open Circuit)
12 | 138 | Current Drawn:
12 | 144 | Continuity Check
12 | 148 | Configuration
12 | 149 | Configuration
12 | 152 | Aperture Orientation
12 | 157 | Postion A
12 | 157 | Postion B
12 | 157 | Postion C
12 | 157 | Postion D
12 | 157 | Max. Difference:
12 | 158 | Position A
12 | 158 | Position B
12 | 158 | Position C
12 | 158 | Position D
12 | 158 | Max. Difference:
12 | 159 | Position A
12 | 159 | Position B
12 | 159 | Position C
12 | 159 | Position D
12 | 159 | Max. Difference:
12 | 160 | Denka Avg. X =
12 | 160 | Extractor Avg. Y =
12 | 160 | Denka to Ext. Distance: (Ext. D. Avg.) – (Denka H. Avg.) =
12 | 160 | Shim Value: Denka to Ext. Distance – 0.75 =
12 | 161 | Position A
12 | 161 | Position B
12 | 161 | Position C
12 | 161 | Position D
12 | 161 | Max. Difference:
12 | 162 | Tip Re-measured Height Avg: Z =
12 | 162 | Actual Tip Height: Y - Z =
12 | 162 | Tip Gap: Actual Tip Height – 0.75 =
12 | 162 | Final Shim Value =
12 | 162 | Denka Torque Setting =
12 | 163 | Tip S/N
12 | 165 | A – B (0.3Ω ± 0.1Ω)
12 | 165 | A – C (Open Circuit)
12 | 165 | A – D (Open Circuit)
12 | 165 | C – D (Open Circuit)
12 | 166 | Liner In – Column GND (Open Circuit)
13 | 168 | Driver A – Liner (>1000MΩ)
13 | 168 | Driver B – Liner (>1000MΩ)
13 | 168 | Driver A – Driver B (>1000MΩ)
13 | 168 | Driver A – Liner (>1000MΩ)
13 | 171 | Driver A – Liner (>1000MΩ)
13 | 171 | Driver B – Liner (>1000MΩ)
13 | 171 | Driver A – Driver B (>1000MΩ)
13 | 171 | Driver A – Liner (>1000MΩ)
13 | 172 | Driver A – Driver A In
13 | 172 | Driver B – Driver B In
13 | 172 | Anode – Liner In
14 | 174 | Vacuum Check (-15PSI @ 5 Mins)
14 | 178 | Bake Time:
14 | 181 | Bake Time =
14 | 184 | Start Gun Vacuum Tracing
14 | 184 | Print and Save Gun Vacuum Trace
14 | 185 | IGP 1 Pressure (≤ 8.0e-10 mbar)
14 | 1649 | IGP 1 Pressure (≤ 8.0e-10 mbar)
14 | 1649 | IGP 2 Pressure
11 | 131 | Continuity (Close Circuit)
11 | 131 | Coils to GND (Open Circuit)
7 | 1650 | S/N =
8 | 1651 | S/N =
9 | 1652 | S/N =
10 | 1653 | S/N =
11 | 1654 | S/N =
12 | 1655 | S/N =
7 | 76 | Open Circuit (Coil – GND)
7 | 83 | Open Circuit (Coil – GND)
8 | 92 | Cooling Isolation Check (Open Circuit Coil – GND)
13 | 168 | Driver A – Liner (>1000MΩ)
13 | 168 | Driver B – Liner (>1000MΩ)
13 | 168 | Driver A – Driver B (>1000MΩ)
13 | 168 | Driver A – Liner (>1000MΩ)
651 | 10903 | SmartPI Software Installed
14 | 178 | Bake Time:
12 | 148 | Check Configuration is correct for $[MaxCurrent]nA Column.
12 | 149 | Confirm Configuration is correct for $[MaxCurrent]nA Column
12 | 152 | Aperture Orientation is correct for $[MaxCurrent]na Column
12 | 157 | Postion A
12 | 157 | Postion B
12 | 157 | Postion C
12 | 157 | Postion D
12 | 157 | Max. Difference:
12 | 158 | Position A
12 | 158 | Position B
12 | 158 | Position C
12 | 158 | Position D
12 | 158 | Max. Difference:
12 | 159 | Position A
12 | 159 | Position B
12 | 159 | Position C
12 | 159 | Position D
12 | 159 | Max. Difference:
12 | 160 | Denka Avg. X =
12 | 160 | Extractor Avg. Y =
12 | 160 | Denka to Ext. Distance: (Ext. D. Avg.) – (Denka H. Avg.) =
12 | 160 | Shim Value: Denka to Ext. Distance – 0.75 =
12 | 161 | Position A
12 | 161 | Position B
12 | 161 | Position C
12 | 161 | Position D
12 | 161 | Max. Difference:
12 | 162 | Tip Re-measured Height Avg: Z =
12 | 162 | Actual Tip Height: Y - Z =
12 | 162 | Tip Gap: Actual Tip Height – 0.75 =
12 | 162 | Final Shim Value =
12 | 162 | Denka Torque Setting =
14 | 174 | Vacuum Check (-15PSI @ 5 Mins)
14 | 181 | Bake Time =
14 | 1649 | IGP 1 Pressure (≤ 8.0e-10 mbar)
14 | 1649 | IGP 2 Pressure
12 | 150 | Check Configuration is correct for $[MaxCurrent]nA Column.
12 | 151 | Check Configuration is correct for $[MaxCurrent]nA Column.
12 | 1660 | Initials of 2nd engineer confirming.
9 | 1652 | S/N =
12 | 138 | Current Drawn:
14 | 184 | Start Gun Vacuum Tracing
14 | 184 | Save Gun Vacuum Trace
8 | 91 | Housing Type
8 | 94 | Serial:
8 | 3295 | Serial:
10 | 110 | Continuity (Close Circuit)
10 | 110 | Coils to GND (Open Circuit)
12 | 142 | UHV Plasma Cleaned
12 | 1660 | Initials of 2nd engineer confirming.
12 | 11076 | Assembly Engineer
12 | 11076 | Second Engineer Inspection
14 | 175 | UHV Vent cycles
14 | 11122 | UHV Vent cycles
12 | 142 | UHV Plasma Cleaned
13 | 168 | Driver A – Liner (>1000MΩ)
13 | 168 | Driver B – Liner (>1000MΩ)
13 | 168 | Driver A – Driver B (>1000MΩ)
13 | 168 | Driver A – Liner (>1000MΩ)
13 | 171 | Driver A – Liner (>1000MΩ)
13 | 171 | Driver B – Liner (>1000MΩ)
13 | 171 | Driver A – Driver B (>1000MΩ)
13 | 171 | Driver A – Liner (>1000MΩ)
13 | 172 | Driver A – Driver A In
13 | 172 | Driver B – Driver B In
13 | 172 | Anode – Liner In
10 | 10678 | Has Transzorb been fitted
14 | 175 | Number of UHV Vent cycles
14 | 175 | Number of UHV Vent cycles
14 | 11122 | UHV Vent cycles
14 | 11122 | Number of Column Interventions
12 | 11144 | Vacuum Check (-15PSI @ 5 Mins)
12 | 11145 | Number of UHV Vent cycles
12 | 138 | Current Drawn:
14 | 1649 | IGP 1 Pressure (≤ 5.0e-10 mbar)
14 | 1649 | IGP 2 Pressure
12 | 138 | Ion Pump Current draw < 1.0e-07 A
14 | 11122 | UHV Vent cycles
10 | 106 | Position A:
10 | 106 | Position B:
10 | 106 | Position C:
10 | 106 | Position D:
10 | 106 | Max. Difference
10 | 109 | Continuity (Close Circuit)
10 | 109 | Coils to GND (Open Circuit)
10 | 110 | Continuity (Close Circuit)
10 | 110 | Coils to GND (Open Circuit)
10 | 1653 | S/N =
10 | 10678 | Has Transzorb been fitted
707 | 11156 | SaveA
707 | 11156 | Save B
707 | 11156 | Save C
707 | 11156 | Images saved
7 | 1640 | Position A:
7 | 1640 | Position B:
7 | 1640 | Position C:
7 | 1640 | Position D:
7 | 1640 | Max Difference:
7 | 1641 | Stand Off 1:
7 | 1641 | Stand Off 2:
7 | 1641 | Stand Off 3:
7 | 1641 | Stand Off 4:
7 | 1641 | Max Difference:
8 | 94 | Serial:
13 | 11190 | Driver A – Liner (>1000MΩ)
13 | 11191 | Driver A In – Anode (>1000MΩ)
13 | 11191 | Driver A In – GND (>1000MΩ)
13 | 11191 | Anode - GND (>1000MΩ)
13 | 11192 | Liner In - Anode
13 | 11192 | Driver A - Driver A In
707 | 11156 | Ensure size of hole is 20um (+/- 2um)
707 | 11156 | Use LR conditions (1024x768, N=1)
707 | 11156 | Save to TIFF as Multi_TS_20um
707 | 11156 | Inspection Passed and Images Saved
707 | 11158 | Ensure size of hole is 30um (+/- 2um)
707 | 11158 | Use LR conditions (1024x768, N=1)
707 | 11158 | Save to TIFF as Multi_TS_30um
707 | 11158 | Inspection Passed and Images Saved
707 | 11159 | Ensure size of hole is 300um (+/- 2um)
707 | 11159 | Use LR conditions (1024x768, N=1)
707 | 11159 | Save to TIFF as Multi_TS_300um
707 | 11159 | Inspection Passed and Images Saved
707 | 11160 | Ensure size of hole is 120um (+/- 2um)
707 | 11160 | Use LR conditions (1024x768, N=1)
707 | 11160 | Save to TIFF as Multi_TS_120um
707 | 11160 | Inspection Passed and Images Saved
707 | 11161 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11161 | Use HR conditions (6144x4608, N=5)
707 | 11161 | Save to TIFF as Multi_US_HRQA
707 | 11161 | Inspection Passed and Images Saved
707 | 11162 | (There is no specific checks not regarding the search for particles)
707 | 11162 | Use HR conditions (6144x4608, N=5)
707 | 11162 | Save to TIFF as Anode_TS_40um
707 | 11162 | Inspection Passed and Images Saved
707 | 11163 | (There is no specific checks not regarding the search for particles)
707 | 11163 | Use LR conditions (1024x768, N=1)
707 | 11163 | Save to TIFF as Anode_TS_Bore1
707 | 11163 | Inspection Passed and Images Saved
707 | 11164 | (There is no specific checks not regarding the search for particles)
707 | 11164 | Use LR conditions (1024x768, N=1)
707 | 11164 | Save to TIFF as Anode_TS_Bore2
707 | 11164 | Inspection Passed and Images Saved
707 | 11165 | Ensure size of hole is 40um (+/- 2um)
707 | 11165 | Use LR conditions (1024x768, N=1)
707 | 11165 | Save to TIFF as Anode_TS_40um
707 | 11165 | Inspection Passed and Images Saved
707 | 11166 | Ensure size of hole is 55um (+/- 2um
707 | 11166 | Use LR conditions (1024x768, N=1)
707 | 11166 | Save to TIFF as Anode_TS_55um
707 | 11166 | Inspection Passed and Images Saved
707 | 11167 | Ensure size of hole is 90um (+/- 2um)
707 | 11167 | Use LR conditions (1024x768, N=1)
707 | 11167 | Save to TIFF as Anode_TS_90um
707 | 11167 | Inspection Passed and Images Saved
707 | 11187 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11187 | Use HR conditions (6144x4608, N=5)
707 | 11187 | Save to TIFF as Multi_TS_HRQA
707 | 11187 | Inspection Passed and Images Saved
12 | 142 | Plasma cleaning completed
707 | 11159 | Ensure size of hole is 300um (+/- 10um)
707 | 11159 | Use LR conditions (1024x768, N=1)
707 | 11159 | Save to TIFF as Multi_TS_300um
707 | 11159 | Inspection Passed and Images Saved
707 | 11160 | Ensure size of hole is 120um (+/- 3um)
707 | 11160 | Use LR conditions (1024x768, N=1)
707 | 11160 | Save to TIFF as Multi_TS_120um
707 | 11160 | Inspection Passed and Images Saved
13 | 172 | Driver A – Driver A In
13 | 172 | Driver B – Driver B In
13 | 172 | Anode – Liner In
12 | 11144 | Vacuum Check (-15PSI @ 5 Mins)
12 | 145 | Checked By Assembly Engineer:
12 | 11208 | UHV Thermocouple resistance is 110Ὡ (+/-10Ὡ)
14 | 1649 | IGP 1 Pressure (≤ 5.0e-10 mbar)
14 | 1649 | IGP 2 Pressure
14 | 11211 | IGP Pressure (≤ 8.0e-10 mbar)
12 | 11208 | UHV Thermocouple resistance is 110Ὡ (+/-10Ὡ)
14 | 185 | IGP Pressure (≤ 8.0e-10 mbar)
13 | 168 | Driver A – Liner (>1000MΩ)
13 | 168 | Driver B – Liner (>1000MΩ)
13 | 168 | Driver A – Driver B (>1000MΩ)
13 | 168 | Driver A – Liner (>1000MΩ)
13 | 171 | Driver A – Liner (>1000MΩ)
13 | 171 | Driver B – Liner (>1000MΩ)
13 | 171 | Driver A – Driver B (>1000MΩ)
13 | 171 | Driver A – Liner (>1000MΩ)
9 | 97 | Ion Pump Current draw < 1.0e-07 A
9 | 11242 | O-Rings TEM oiled
14 | 11245 | All Paperwork Sorted and Stapled Together
12 | 11208 | UHV Thermocouple resistance is 110Ὡ (+/-10Ὡ)
7 | 11258 | Continuity ok
12 | 11266 | Flange Heater Earths routed correctly (where applicable)
12 | 11266 | Thermocouple cables routed correctly
654 | 11197 | Remove Files
654 | 11197 | Remove Folders
14 | 11245 | All Paperwork Sorted and Stapled Together
13 | 172 | Driver A – Driver A In
13 | 172 | Driver B – Driver B In
13 | 172 | Anode – Liner In
645 | 10525 | Pumpdown Time
7 | 11340 | Build quality of Objective coil is satisfactory
707 | 11187 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11187 | Use HR conditions (6144x4608, N=5)
707 | 11187 | Save to TIFF as Multi_TS_HRQA
707 | 11187 | Inspection Passed and Images Saved
8 | 11409 | Serial:
651 | 10904 | Mineralogic Software Installed
707 | 11277 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11277 | Use HR conditions (6144x4608, N=5)
707 | 11277 | Save to TIFF as Multi_TS_HRQA
707 | 11277 | Inspection Passed and Images Saved
707 | 11283 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11283 | Use HR conditions (6144x4608, N=5)
707 | 11283 | Save to TIFF as Multi_TS_HRQA
707 | 11283 | Inspection Passed and Images Saved
707 | 11284 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11284 | Use HR conditions (6144x4608, N=5)
707 | 11284 | Save to TIFF as Multi_TS_HRQA
707 | 11284 | Inspection Passed and Images Saved
707 | 11156 | Ensure size of hole is 20um (+/- 2um)
707 | 11156 | Use LR conditions (1024x768, N=1)
707 | 11156 | Save to TIFF as Multi_TS_20um
707 | 11156 | Inspection Passed and Images Saved
707 | 11159 | Ensure size of hole is 300um (+/- 10um)
707 | 11159 | Use LR conditions (1024x768, N=1)
707 | 11159 | Save to TIFF as Multi_TS_300um
707 | 11159 | Inspection Passed and Images Saved
707 | 11160 | Ensure size of hole is 120um (+/- 3um)
707 | 11160 | Use LR conditions (1024x768, N=1)
707 | 11160 | Save to TIFF as Multi_TS_120um
707 | 11160 | Inspection Passed and Images Saved
651 | 10905 | Run Up & Obtain An Image
707 | 11161 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11161 | Use HR conditions (6144x4608, N=5)
707 | 11161 | Save to TIFF as Multi_US_HRQA
707 | 11161 | Inspection Passed and Images Saved
707 | 11252 | Ensure size of hole is 30um (+/- 2um)
707 | 11252 | Use LR conditions (1024x768, N=1)
707 | 11252 | Use LR conditions (1024x768, N=1)
707 | 11252 | Inspection Passed and Images Saved
707 | 11284 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11284 | Use HR conditions (6144x4608, N=5)
707 | 11284 | Save to TIFF as Multi_TS_HRQA
707 | 11284 | Inspection Passed and Images Saved
707 | 11285 | Ensure size of hole is 30um (+/- 2um)
707 | 11285 | Use LR conditions (1024x768, N=1)
707 | 11285 | Use LR conditions (1024x768, N=1)
707 | 11285 | Inspection Passed and Images Saved
707 | 11286 | Ensure size of hole is 20um (+/- 2um)
707 | 11286 | Use LR conditions (1024x768, N=1)
707 | 11286 | Save to TIFF as Multi_TS_20um
707 | 11286 | Inspection Passed and Images Saved
643 | 10488 | Run EO Test
707 | 11287 | Ensure size of hole is 120um (+/- 3um)
707 | 11287 | Use LR conditions (1024x768, N=1)
707 | 11287 | Save to TIFF as Multi_TS_120um
707 | 11287 | Inspection Passed and Images Saved
707 | 11288 | Ensure size of hole is 300um (+/- 10um)
707 | 11288 | Use LR conditions (1024x768, N=1)
707 | 11288 | Save to TIFF as Multi_TS_300um
707 | 11288 | Inspection Passed and Images Saved
707 | 11289 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11289 | Use HR conditions (6144x4608, N=5)
707 | 11289 | Save to TIFF as Multi_US_HRQA
707 | 11289 | Inspection Passed and Images Saved
707 | 11290 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11290 | Use HR conditions (6144x4608, N=5)
707 | 11290 | Save to TIFF as Multi_US_HRQA
707 | 11290 | Inspection Passed and Images Saved
707 | 11308 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11308 | Use HR conditions (6144x4608, N=5)
707 | 11308 | Save to TIFF as Multi_US_HRQA
707 | 11308 | Inspection Passed and Images Saved
403 | 10467 | Install SmartSEM + ASAP Engine
403 | 10467 | Record version including service packs and hot fixes if applicable
651 | 10906 | ZEN Software Installed
707 | 11262 | (There is no specific checks not regarding the search for particles)
707 | 11262 | Use HR conditions (6144x4608, N=5)
707 | 11262 | Save to TIFF as Anode_TS_HRQA
707 | 11262 | Inspection Passed and Images Saved
707 | 11284 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11284 | Use HR conditions (6144x4608, N=5)
707 | 11284 | Save to TIFF as Multi_TS_HRQA
707 | 11284 | Inspection Passed and Images Saved
707 | 11317 | (There is no specific checks not regarding the search for particles)
707 | 11317 | Use HR conditions (6144x4608, N=5)
707 | 11317 | Save to TIFF as Anode_TS_HRQA
707 | 11317 | Inspection Passed and Images Saved
707 | 11318 | (There is no specific checks not regarding the search for particles)
707 | 11318 | Use HR conditions (6144x4608, N=5)
707 | 11318 | Save to TIFF as Anode_TS_HRQA
707 | 11318 | Inspection Passed and Images Saved
707 | 11319 | (There is no specific checks not regarding the search for particles)
707 | 11319 | Use HR conditions (6144x4608, N=5)
707 | 11319 | Save to TIFF as Anode_TS_HRQA
707 | 11319 | Inspection Passed and Images Saved
8 | 11336 | Serial Number:
707 | 11163 | (There is no specific checks not regarding the search for particles)
707 | 11163 | Use LR conditions (1024x768, N=1)
707 | 11163 | Save to TIFF as Anode_TS_Bore1
707 | 11163 | Inspection Passed and Images Saved
707 | 11164 | (There is no specific checks not regarding the search for particles)
707 | 11164 | Use LR conditions (1024x768, N=1)
707 | 11164 | Save to TIFF as Anode_TS_Bore2
707 | 11164 | Inspection Passed and Images Saved
707 | 11165 | Ensure size of hole is 40um (+/- 2um)
707 | 11165 | Use LR conditions (1024x768, N=1)
707 | 11165 | Save to TIFF as Anode_TS_40um
707 | 11165 | Inspection Passed and Images Saved
707 | 11166 | Ensure size of hole is 55um (+/- 2um
707 | 11166 | Use LR conditions (1024x768, N=1)
707 | 11166 | Save to TIFF as Anode_TS_55um
707 | 11166 | Inspection Passed and Images Saved
707 | 11167 | Ensure size of hole is 90um (+/- 5um)
707 | 11167 | Use LR conditions (1024x768, N=1)
707 | 11167 | Save to TIFF as Anode_TS_90um
707 | 11167 | Inspection Passed and Images Saved
707 | 11168 | Ensure size of hole is 150um (+/- 5um)
707 | 11168 | Use LR conditions (1024x768, N=1)
707 | 11168 | Save to TIFF as Anode_TS_150um
707 | 11168 | Inspection Passed and Images Saved
707 | 11169 | (There is no specific checks not regarding the search for particles)
707 | 11169 | Use HR conditions (6144x4608, N=5)
707 | 11169 | Save to TIFF as Anode_US_HRQA
707 | 11169 | Inspection Passed and Images Saved
707 | 11170 | (Check surface finish on the inner bore edge is free from marks)
707 | 11170 | Use HR conditions (6144x4608, N=5)
707 | 11170 | Save to TIFF as VP_TS_HRQA
707 | 11170 | Inspection Passed and Images Saved
707 | 11171 | (Check surface finish on the inner bore edge is free from marks)
707 | 11171 | Use LR conditions (1024x768, N=1))
707 | 11171 | Save to TIFF as VP_TS_Bore
707 | 11171 | Inspection Passed and Images Saved
707 | 11172 | (Check there is no engraved ring around the aperture, this should be on the underside.
707 | 11172 | Use HR conditions (6144x4608, N=5)
707 | 11172 | Save to TIFF as SF_TS_HRQA
707 | 11172 | Inspection Passed and Images Saved
707 | 11173 | Ensure size of hole is 35um (+/- 2 um)
707 | 11173 | Use HR conditions (6144x4608, N=5)
707 | 11173 | Save to TIFF as SF_TS_35um
707 | 11173 | Inspection Passed and Images Saved
707 | 11174 | Ensure size of hole is 750um (+/- 15um)
707 | 11174 | Use HR conditions (6144x4608, N=5)
707 | 11174 | Save to TIFF as Plat_TS_750um_HRQA
707 | 11174 | Inspection Passed and Images Saved
707 | 11262 | (There is no specific checks not regarding the search for particles)
707 | 11262 | Use HR conditions (6144x4608, N=5)
707 | 11262 | Save to TIFF as Anode_TS_HRQA
707 | 11262 | Inspection Passed and Images Saved
707 | 11317 | (There is no specific checks not regarding the search for particles)
707 | 11317 | Use HR conditions (6144x4608, N=5)
707 | 11317 | Save to TIFF as Anode_TS_HRQA
707 | 11317 | Inspection Passed and Images Saved
707 | 11326 | (There is no specific checks not regarding the search for particles)
707 | 11326 | Use LR conditions (1024x768, N=1)
707 | 11326 | Save to TIFF as Anode_TS_Bore1
707 | 11326 | Inspection Passed and Images Saved
707 | 11327 | (There is no specific checks not regarding the search for particles)
707 | 11327 | Use LR conditions (1024x768, N=1)
707 | 11327 | Save to TIFF as Anode_TS_Bore1
707 | 11327 | Inspection Passed and Images Saved
642 | 8699 | Cable Up
642 | 8699 | Label Cables
642 | 8699 | Tape PC Configuration Image to PC
707 | 11328 | (There is no specific checks not regarding the search for particles)
707 | 11328 | Use LR conditions (1024x768, N=1)
707 | 11328 | Save to TIFF as Anode_TS_Bore1
707 | 11328 | Inspection Passed and Images Saved
707 | 11329 | (There is no specific checks not regarding the search for particles)
707 | 11329 | Use LR conditions (1024x768, N=1)
707 | 11329 | Save to TIFF as Anode_TS_Bore2
707 | 11329 | Inspection Passed and Images Saved
707 | 11330 | (There is no specific checks not regarding the search for particles)
707 | 11330 | Use LR conditions (1024x768, N=1)
707 | 11330 | Save to TIFF as Anode_TS_Bore2
707 | 11330 | Inspection Passed and Images Saved
707 | 11331 | (There is no specific checks not regarding the search for particles)
707 | 11331 | Use LR conditions (1024x768, N=1)
707 | 11331 | Save to TIFF as Anode_TS_Bore2
707 | 11331 | Inspection Passed and Images Saved
707 | 11332 | (There is no specific checks not regarding the search for particles)
707 | 11332 | Use HR conditions (6144x4608, N=5)
707 | 11332 | Save to TIFF as Anode_US_HRQA
707 | 11332 | Inspection Passed and Images Saved
707 | 11156 | Ensure size of hole is 20um (+/- 2um)
707 | 11156 | Use LR conditions (1024x768, N=1)
707 | 11156 | Save to TIFF as Multi_TS_20um
707 | 11156 | Inspection Passed and Images Saved
707 | 11333 | (There is no specific checks not regarding the search for particles)
707 | 11333 | Use HR conditions (6144x4608, N=5)
707 | 11333 | Save to TIFF as Anode_US_HRQA
707 | 11333 | Inspection Passed and Images Saved
707 | 11334 | (There is no specific checks not regarding the search for particles)
707 | 11334 | Use HR conditions (6144x4608, N=5)
707 | 11334 | Save to TIFF as Anode_US_HRQA
707 | 11334 | Inspection Passed and Images Saved
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
708 | 11238 | Make sure there are no cut-outs on the edge of the anode. No particles are to be observed.
707 | 11161 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11161 | Use HR conditions (6144x4608, N=5)
707 | 11161 | Save to TIFF as Multi_US_HRQA
707 | 11161 | Inspection Passed and Images Saved
707 | 11163 | (There is no specific checks not regarding the search for particles)
707 | 11163 | Use LR conditions (1024x768, N=1)
707 | 11163 | Save to TIFF as Anode_TS_Bore1
707 | 11163 | Inspection Passed and Images Saved
707 | 11284 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11284 | Use HR conditions (6144x4608, N=5)
707 | 11284 | Save to TIFF as Multi_TS_HRQA
707 | 11284 | Inspection Passed and Images Saved
707 | 11164 | (There is no specific checks not regarding the search for particles)
707 | 11164 | Use LR conditions (1024x768, N=1)
707 | 11164 | Save to TIFF as Anode_TS_Bore2
707 | 11164 | Inspection Passed and Images Saved
707 | 11167 | Ensure size of hole is 90um (+/- 5um)
707 | 11167 | Use LR conditions (1024x768, N=1)
707 | 11167 | Save to TIFF as Anode_TS_90um
707 | 11167 | Inspection Passed and Images Saved
707 | 11169 | (There is no specific checks not regarding the search for particles)
707 | 11169 | Use HR conditions (6144x4608, N=5)
707 | 11169 | Save to TIFF as Anode_US_HRQA
707 | 11169 | Inspection Passed and Images Saved
707 | 11187 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11187 | Use HR conditions (6144x4608, N=5)
707 | 11187 | Save to TIFF as Multi_TS_HRQA
707 | 11187 | Inspection Passed and Images Saved
707 | 11252 | Ensure size of hole is 30um (+/- 2um)
707 | 11252 | Use LR conditions (1024x768, N=1)
707 | 11252 | Use LR conditions (1024x768, N=1)
707 | 11252 | Inspection Passed and Images Saved
708 | 11265 | Make sure there are 3 cut-outs on the edge of the anode. No particles are to be observed.
707 | 11262 | (There is no specific checks not regarding the search for particles)
707 | 11262 | Use HR conditions (6144x4608, N=5)
707 | 11262 | Save to TIFF as Anode_TS_HRQA
707 | 11262 | Inspection Passed and Images Saved
707 | 11288 | Ensure size of hole is 300um (+/- 10um)
707 | 11288 | Use LR conditions (1024x768, N=1)
707 | 11288 | Save to TIFF as Multi_TS_300um
707 | 11288 | Inspection Passed and Images Saved
707 | 11159 | Ensure size of hole is 300um (+/- 10um)
707 | 11159 | Use LR conditions (1024x768, N=1)
707 | 11159 | Save to TIFF as Multi_TS_300um
707 | 11159 | Inspection Passed and Images Saved
707 | 11160 | Ensure size of hole is 120um (+/- 3um)
707 | 11160 | Use LR conditions (1024x768, N=1)
707 | 11160 | Save to TIFF as Multi_TS_120um
707 | 11160 | Inspection Passed and Images Saved
707 | 11161 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11161 | Use HR conditions (6144x4608, N=5)
707 | 11161 | Save to TIFF as Multi_US_HRQA
707 | 11161 | Inspection Passed and Images Saved
8 | 94 | Serial:
707 | 11185 | Hold aperture holder, look for holes on the sides.No particles are to be observed
707 | 11186 | Select PC
707 | 11186 | Hold aperture holder, look for holes on the sides.No particles are to be observed
707 | 11187 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11187 | Use HR conditions (6144x4608, N=5)
707 | 11187 | Save to TIFF as Multi_TS_HRQA
707 | 11187 | Inspection Passed and Images Saved
707 | 11252 | Ensure size of hole is 30um (+/- 2um)
707 | 11252 | Use LR conditions (1024x768, N=1)
707 | 11252 | Use LR conditions (1024x768, N=1)
707 | 11252 | Inspection Passed and Images Saved
707 | 11277 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11277 | Use HR conditions (6144x4608, N=5)
707 | 11277 | Save to TIFF as Multi_TS_HRQA
707 | 11277 | Inspection Passed and Images Saved
707 | 11283 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11283 | Use HR conditions (6144x4608, N=5)
707 | 11283 | Save to TIFF as Multi_TS_HRQA
707 | 11283 | Inspection Passed and Images Saved
707 | 11285 | Ensure size of hole is 30um (+/- 2um)
707 | 11285 | Use LR conditions (1024x768, N=1)
707 | 11285 | Use LR conditions (1024x768, N=1)
707 | 11285 | Inspection Passed and Images Saved
707 | 11286 | Ensure size of hole is 20um (+/- 2um)
707 | 11286 | Use LR conditions (1024x768, N=1)
707 | 11286 | Save to TIFF as Multi_TS_20um
707 | 11286 | Inspection Passed and Images Saved
707 | 11287 | Ensure size of hole is 120um (+/- 3um)
707 | 11287 | Use LR conditions (1024x768, N=1)
707 | 11287 | Save to TIFF as Multi_TS_120um
707 | 11287 | Inspection Passed and Images Saved
707 | 11288 | Ensure size of hole is 300um (+/- 10um)
707 | 11288 | Use LR conditions (1024x768, N=1)
707 | 11288 | Save to TIFF as Multi_TS_300um
707 | 11288 | Inspection Passed and Images Saved
707 | 11289 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11289 | Use HR conditions (6144x4608, N=5)
707 | 11289 | Save to TIFF as Multi_US_HRQA
707 | 11289 | Inspection Passed and Images Saved
707 | 11290 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11290 | Use HR conditions (6144x4608, N=5)
707 | 11290 | Save to TIFF as Multi_US_HRQA
707 | 11290 | Inspection Passed and Images Saved
707 | 11308 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11308 | Use HR conditions (6144x4608, N=5)
707 | 11308 | Save to TIFF as Multi_US_HRQA
707 | 11308 | Inspection Passed and Images Saved
7 | 11258 | Continuity ok
643 | 10489 | Scintillator
643 | 10489 | Collector Volts
708 | 11236 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed.
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
708 | 11241 | Make sure there are no cut-outs on the edge of the anode. No particles are to be observed.
708 | 11263 | Make sure there are 3 cut-outs on the edge of the anode. No particles are to be observed
708 | 11264 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be observed.
707 | 11286 | Ensure size of hole is 20um (+/- 2um)
707 | 11286 | Use LR conditions (1024x768, N=1)
707 | 11286 | Save to TIFF as Multi_TS_20um
707 | 11286 | Inspection Passed and Images Saved
707 | 11165 | Ensure size of hole is 40um (+/- 2um)
707 | 11165 | Use LR conditions (1024x768, N=1)
707 | 11165 | Save to TIFF as Anode_TS_40um
707 | 11165 | Inspection Passed and Images Saved
707 | 11185 | Hold aperture holder, look for holes on the sides.No particles are to be observed
707 | 11284 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11284 | Use HR conditions (6144x4608, N=5)
707 | 11284 | Save to TIFF as Multi_TS_HRQA
707 | 11284 | Inspection Passed and Images Saved
707 | 11327 | (There is no specific checks not regarding the search for particles)
707 | 11327 | Use LR conditions (1024x768, N=1)
707 | 11327 | Save to TIFF as Anode_TS_Bore1
707 | 11327 | Inspection Passed and Images Saved
707 | 11330 | (There is no specific checks not regarding the search for particles)
707 | 11330 | Use LR conditions (1024x768, N=1)
707 | 11330 | Save to TIFF as Anode_TS_Bore2
707 | 11330 | Inspection Passed and Images Saved
12 | 11354 | UHV Thermocouple resistance is 110Ὡ (+/-10Ὡ)
707 | 11333 | (There is no specific checks not regarding the search for particles)
707 | 11333 | Use HR conditions (6144x4608, N=5)
707 | 11333 | Save to TIFF as Anode_US_HRQA
707 | 11333 | Inspection Passed and Images Saved
708 | 11236 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed.
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
708 | 11241 | Make sure there are no cut-outs on the edge of the anode. No particles are to be observed.
708 | 11263 | Make sure there are 3 cut-outs on the edge of the anode. No particles are to be observed
12 | 11385 | Confirm correct fitting of Raith spray aperture
14 | 184 | Start Gun Vacuum Tracing
14 | 184 | Save Gun Vacuum Trace
707 | 11156 | Ensure size of hole is 20um (+/- 2um)
707 | 11156 | Use LR conditions (1024x768, N=1)
707 | 11156 | Save to TIFF as Multi_TS_20um
707 | 11156 | Inspection Passed and Images Saved
7 | 11393 | Confirm there is no Electical breakdown between Scan coils and Objective body.
707 | 11160 | Ensure size of hole is 120um (+/- 3um)
707 | 11160 | Use LR conditions (1024x768, N=1)
707 | 11160 | Save to TIFF as Multi_TS_120um
707 | 11160 | Inspection Passed and Images Saved
707 | 11165 | Ensure size of hole is 40um (+/- 2um)
707 | 11165 | Use LR conditions (1024x768, N=1)
707 | 11165 | Save to TIFF as Anode_TS_40um
707 | 11165 | Inspection Passed and Images Saved
707 | 11283 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11283 | Use HR conditions (6144x4608, N=5)
707 | 11283 | Save to TIFF as Multi_TS_HRQA
707 | 11283 | Inspection Passed and Images Saved
707 | 11287 | Ensure size of hole is 120um (+/- 3um)
707 | 11287 | Use LR conditions (1024x768, N=1)
707 | 11287 | Save to TIFF as Multi_TS_120um
707 | 11287 | Inspection Passed and Images Saved
7 | 83 | Open Circuit (Coil – GND).
403 | 6935 | Record MAC address
403 | 6935 | Network access requested
707 | 11308 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11308 | Use HR conditions (6144x4608, N=5)
707 | 11308 | Save to TIFF as Multi_US_HRQA
707 | 11308 | Inspection Passed and Images Saved
707 | 11318 | (There is no specific checks not regarding the search for particles)
707 | 11318 | Use HR conditions (6144x4608, N=5)
707 | 11318 | Save to TIFF as Anode_TS_HRQA
707 | 11318 | Inspection Passed and Images Saved
707 | 11327 | (There is no specific checks not regarding the search for particles)
707 | 11327 | Use LR conditions (1024x768, N=1)
707 | 11327 | Save to TIFF as Anode_TS_Bore1
707 | 11327 | Inspection Passed and Images Saved
707 | 11330 | (There is no specific checks not regarding the search for particles)
707 | 11330 | Use LR conditions (1024x768, N=1)
707 | 11330 | Save to TIFF as Anode_TS_Bore2
707 | 11330 | Inspection Passed and Images Saved
707 | 11333 | (There is no specific checks not regarding the search for particles)
707 | 11333 | Use HR conditions (6144x4608, N=5)
707 | 11333 | Save to TIFF as Anode_US_HRQA
707 | 11333 | Inspection Passed and Images Saved
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
708 | 11238 | Make sure there are no cut-outs on the edge of the anode. No particles are to be observed.
12 | 138 | Ion Pump Current draw < 1.0e-07 A
707 | 11186 | Select PC
707 | 11186 | Hold aperture holder, look for holes on the sides.No particles are to be observed
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
648 | 10875 | 2KV Res - 750um Easy Align Aperture- ≤16nm
708 | 11238 | Make sure there are no cut-outs on the edge of the anode. No particles are to be observed.
707 | 11166 | Ensure size of hole is 55um (+/- 2um
707 | 11166 | Use LR conditions (1024x768, N=1)
707 | 11166 | Save to TIFF as Anode_TS_55um
707 | 11166 | Inspection Passed and Images Saved
707 | 11172 | (Check there is no engraved ring around the aperture, this should be on the underside.
707 | 11172 | Use HR conditions (6144x4608, N=5)
707 | 11172 | Save to TIFF as SF_TS_HRQA
707 | 11172 | Inspection Passed and Images Saved
707 | 11173 | Ensure size of hole is 35um (+/- 2 um)
707 | 11173 | Use HR conditions (6144x4608, N=5)
707 | 11173 | Save to TIFF as SF_TS_35um
707 | 11173 | Inspection Passed and Images Saved
707 | 11319 | (There is no specific checks not regarding the search for particles)
707 | 11319 | Use HR conditions (6144x4608, N=5)
707 | 11319 | Save to TIFF as Anode_TS_HRQA
707 | 11319 | Inspection Passed and Images Saved
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
707 | 11328 | (There is no specific checks not regarding the search for particles)
707 | 11328 | Use LR conditions (1024x768, N=1)
707 | 11328 | Save to TIFF as Anode_TS_Bore1
707 | 11328 | Inspection Passed and Images Saved
707 | 11331 | (There is no specific checks not regarding the search for particles)
707 | 11331 | Use LR conditions (1024x768, N=1)
707 | 11331 | Save to TIFF as Anode_TS_Bore2
707 | 11331 | Inspection Passed and Images Saved
707 | 11334 | (There is no specific checks not regarding the search for particles)
707 | 11334 | Use HR conditions (6144x4608, N=5)
707 | 11334 | Save to TIFF as Anode_US_HRQA
707 | 11334 | Inspection Passed and Images Saved
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
708 | 11241 | Make sure there are no cut-outs on the edge of the anode. No particles are to be observed.
44 | 455 | Interlock tested and working
708 | 11264 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be observed.
708 | 11265 | Make sure there are 3 cut-outs on the edge of the anode. No particles are to be observed.
707 | 11165 | Ensure size of hole is 40um (+/- 2um)
707 | 11165 | Use LR conditions (1024x768, N=1)
707 | 11165 | Save to TIFF as Anode_TS_40um
707 | 11165 | Inspection Passed and Images Saved
707 | 11185 | Hold aperture holder, look for holes on the sides.No particles are to be observed
707 | 11284 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11284 | Use HR conditions (6144x4608, N=5)
707 | 11284 | Save to TIFF as Multi_TS_HRQA
707 | 11284 | Inspection Passed and Images Saved
403 | 10465 | Record Product ID
403 | 10465 | Disable Renesas USB Power Management
707 | 11318 | (There is no specific checks not regarding the search for particles)
707 | 11318 | Use HR conditions (6144x4608, N=5)
707 | 11318 | Save to TIFF as Anode_TS_HRQA
707 | 11318 | Inspection Passed and Images Saved
707 | 11327 | (There is no specific checks not regarding the search for particles)
707 | 11327 | Use LR conditions (1024x768, N=1)
707 | 11327 | Save to TIFF as Anode_TS_Bore1
707 | 11327 | Inspection Passed and Images Saved
707 | 11330 | (There is no specific checks not regarding the search for particles)
707 | 11330 | Use LR conditions (1024x768, N=1)
707 | 11330 | Save to TIFF as Anode_TS_Bore2
707 | 11330 | Inspection Passed and Images Saved
707 | 11333 | (There is no specific checks not regarding the search for particles)
707 | 11333 | Use HR conditions (6144x4608, N=5)
707 | 11333 | Save to TIFF as Anode_US_HRQA
707 | 11333 | Inspection Passed and Images Saved
12 | 11208 | Megger reads >1000Mohms
12 | 11354 | UHV Thermocouple resistance is 110Ὡ (+/-10Ὡ)
7 | 11258 | Continuity ok
12 | 11208 | Megger reads >1000Mohms
642 | 10476 | ASAP Mainboard Revision
642 | 10476 | ASAP Mainboard Hardware
642 | 10476 | ASAP VBE 1 Revision
642 | 10476 | ASAP VBE 1 Firmware Version
642 | 10476 | ASAP VBE 2 (if fitted) Revision
642 | 10476 | ASAP VBE 2 (if fitted) Firmware Version
642 | 10476 | ASAP DFE 1 Revision
642 | 10476 | ASAP DFE 1 Firmware Version
642 | 10476 | ASAP DFE 2 (if fitted) Revision
642 | 10476 | ASAP DFE 2 (if fitted) Firmware Version
642 | 10476 | ASAP DFE 3 (if fitted) Revision
642 | 10476 | ASAP DFE 3 (if fitted) Firmware Version
642 | 10476 | ASAP DFE 4 (if fitted) Revision
642 | 10476 | ASAP DFE 4 (if fitted) Firmware Version
642 | 10477 | Stage Version
642 | 10477 | Controller Version
642 | 10477 | Axes Version
642 | 10478 | Hardware Version
642 | 10479 | Firmware Revision
642 | 10480 | Firmware Version
642 | 10480 | Hardware Version
642 | 10481 | Fit Link on Vacuum PCB
642 | 10481 | Gun Chamber Valve Motorised=Yes
642 | 10485 | Approx. No. of Hours
643 | 10492 | C3 Reversal
643 | 10492 | 30KV-5KV
643 | 10492 | Image Shift
643 | 10494 | C1 Cal
643 | 10494 | C1 Rem
643 | 10494 | C1 Xtalk
643 | 10494 | C2 Cal
643 | 10494 | C2 Rem
643 | 10494 | C2 Xtalk
644 | 10496 | Accelerating Voltage=30KV
644 | 10496 | Specimen Current
644 | 10496 | Monitor Model
644 | 10496 | Serial Number
644 | 10496 | Cal. Expiry Date
644 | 10496 | No X-ray Leaks Detected
644 | 10496 | Complete & Sign X-Ray Certificate
644 | 10497 | Gun Chamber has been cleaned and inspected using UV light
644 | 10497 | EHT set socket well has been cleaned and inspected using UV light
644 | 10497 | Lab6 Firing Unit Fitted
644 | 10497 | Overnight Bakeout
644 | 10498 | Run up to 32KV and leave for 1 minute
644 | 10498 | Shutdown and run up to 32KV, Repeat 5 times
644 | 10498 | Run up to 32KV and leave for 15 minutes
644 | 10500 | Lab6 Firing Unit Run Up to 30KV
644 | 10502 | 30KV @ 20mm X
644 | 10502 | 30KV @ 20mm Y
644 | 10502 | 2KV @ 20mm X
644 | 10502 | 2KV @ 20mm Y
644 | 10502 | 30KV @ 5mm X
644 | 10502 | 30KV @ 5mm Y
644 | 10502 | 2KV @ 5mm X
644 | 10502 | 2KV @ 5mm Y
644 | 10504 | VP 30KV @ 5mm (100um) X
644 | 10504 | VP 30KV @ 5mm (100um) Y
644 | 10505 | VP 30KV @ 5mm (140um) X
644 | 10505 | VP 30KV @ 5mm (140um) Y
644 | 10507 | EP 500um 30KV @ 5mm X
644 | 10507 | EP 500um 30KV @ 5mm Y
644 | 10507 | EP 500um Alignment
644 | 10507 | EP 1000um 30KV @ 5mm X
644 | 10507 | EP 1000um 30KV @ 5mm Y
644 | 10507 | EP 1000um Alignment
644 | 10508 | BeamSleeve 500um 30KV @ 8mm X
644 | 10508 | BeamSleeve 500um 30KV @ 8mm Y
644 | 10508 | BeamSleeve 500um Alignment
644 | 10509 | BeamSleeve 1000um 30KV @ 8mm X
644 | 10509 | BeamSleeve 1000um 30KV @ 8mm Y
644 | 10509 | BeamSleeve 1000um Alignment
645 | 10512 | 30KV @ 5mm X
645 | 10512 | 30KV @ 5mm Y
645 | 10513 | 30KV @ 5mm X
645 | 10513 | 30KV @ 5mm Y
645 | 10516 | X Axis
645 | 10516 | Y Axis
645 | 10516 | X+Y Axis
645 | 10516 | Y+X Axis
645 | 10516 | 10mm Square
645 | 10517 | Z Axis
645 | 10518 | X Axis
645 | 10518 | Y Axis
645 | 10519 | Rotate=20 deg
645 | 10519 | Rotate=45 deg
645 | 10519 | Rotate=90 deg
645 | 10519 | Rotate=180 deg
645 | 10520 | Tilt=5 deg
645 | 10520 | Tilt=10 deg
645 | 10520 | Tilt=20 deg
645 | 10520 | Tilt=45 deg
645 | 10524 | Pumpdown Time
646 | 10526 | Select Port
646 | 10526 | Check Camera for focus
646 | 10526 | Functional Check-Black/White + Colour Mode
646 | 10526 | Run wxPropview to Remove Defective Pixels
646 | 10527 | Basic Movements-T/Z Joystick
646 | 10527 | Basic Movements-X/Y/R Joystick
646 | 10527 | Initialise Stage
646 | 10528 | Installation
646 | 10528 | select Language
646 | 10528 | Function of all knobs and buttons
646 | 10528 | Function of all keyboard keys
646 | 10529 | BSD Diode Pre-test
646 | 10529 | Select Type
646 | 10529 | Black Level Calibration
646 | 10529 | Check Diode Segment Function
646 | 10529 | Check All Gain ranges
646 | 10529 | Check Auto range
646 | 10529 | Check Auto Level
646 | 10529 | Setup and Check TOPO
646 | 10529 | Take 3 images - Atomic, 5pA, 2KV
646 | 10530 | Calibration of Bias
646 | 10530 | Functional Check
646 | 10530 | Function Image
646 | 10531 | Functional Check
646 | 10531 | Function Image
646 | 10532 | Set-Up (Mag.Cal. Wizard
646 | 10532 | Functional Check
646 | 10533 | Functional Check
646 | 10534 | Alignment of Ring (Using Jig)
646 | 10534 | Black Level Calibration
646 | 10534 | Functional Check (Dry)
646 | 10534 | Functional Check (Wet)
646 | 10535 | Black Level Calibration
646 | 10535 | Functional Check
646 | 10535 | Function Image
646 | 10536 | Check 5KV is achievable using the Bias Slider
646 | 10536 | safety Check - Remove PL1 (Vac Ready from vac switch
646 | 10536 | Safety Check - Remove PL2 (Turbo Speed signal)
646 | 10536 | Safety Check - Remove PL4 (CAN Bus signal)
646 | 10536 | Functional Check and Images (Using the BSD Detector)
646 | 10537 | Functional Check
646 | 10537 | Function Image
646 | 10539 | Probe Current Measurment Select in Administrator
646 | 10539 | Functional Check
646 | 10539 | Probe Current Monitor and Specimen Current Meter Read the same
646 | 10539 | Checked at 500pA. 1nA, 10nA
646 | 10539 | Picoammeter and Probe Current Monitor read the same
646 | 10540 | Installation
646 | 10540 | Cable Open Circuit to Ground
646 | 10540 | Blanking Tests Using 20um Aperture
646 | 10540 | Blanking Tests Using 30um Aperture
646 | 10540 | Blank/UnBlank Ratio is better than 0.1% for all settings
646 | 10540 | Low Mag Blanking Check
646 | 10541 | Installation of Quiet Mode Vessel
646 | 10541 | Quiet Mode Fitted Selected in SEM Status
646 | 10541 | Compare On/Off Times With Graph
646 | 10541 | Pump Starts When Quiet Mode is Deselected
646 | 10541 | Run Gun Monitor-Turbo Power and System Vacuum
646 | 10542 | Tested Using Video Mix
646 | 10543 | $NavCam Postion Setup
646 | 10543 | Image Navigation Setup
646 | 10543 | Functional Check
646 | 10544 | Installation
646 | 10544 | Functional Test
646 | 10545 | Install SmartSEM Licences
646 | 10545 | Install MIL Redistribution
646 | 10545 | Fit MIL Dongle
646 | 10545 | Install SmartPI Software
646 | 10545 | Calibrate I/O-SCAN card
646 | 10545 | Setup & Run Application
646 | 10545 | Check Results & Image Montage
646 | 10547 | ZEN Version
646 | 10547 | Correct Version of Windows 10 Installed
646 | 10547 | Install ZEN
646 | 10547 | Install Licence
646 | 10547 | Install Remote API Client
646 | 10547 | Functional Test
646 | 10548 | Connect Silver Box and Cables
646 | 10548 | Install ATLAS Software
646 | 10548 | Request Licence File from Fibics
646 | 10548 | Install Licence File
646 | 10548 | Calibrate SEM
646 | 10548 | Obtain Mosaic
646 | 10548 | Set ATLAS Rotation Offset
646 | 10548 | Save ATLAS Licence Document
646 | 10549 | Install YAG BSD Lightpipe
646 | 10549 | Alignment
646 | 10549 | Select YAG BSD in the Administrator
646 | 10549 | 30KV Function Image
646 | 10549 | Obtain a Image at 3KV
646 | 10645 | Filter Has Been Fitted
646 | 10650 | Basic Movements-T/Z Joystick
646 | 10650 | Basic Movements-X/Y/Z Joystick
646 | 10650 | Initialise Stage
646 | 10650 | Check Break Button
646 | 10651 | Installation
646 | 10651 | Functional Test
646 | 10652 | Type of Detector
646 | 10652 | Serial number
646 | 10652 | Fit Flange & Mount Detector
646 | 10652 | Install EDS PCB Into PC
646 | 10652 | Installation of DPP Box & Cables
646 | 10652 | Configue PC Network & Test
646 | 10652 | Disable Windows10 Fast Startup
646 | 10652 | Install Licences
646 | 10652 | Fit Dongle
646 | 10652 | APEX Software Version
646 | 10652 | Install APEX Software
646 | 10652 | Install Remoting
646 | 10652 | Setup RConfigure
646 | 10652 | Run API Test
646 | 10652 | Configure Imaging-Run SGShell
646 | 10652 | Calibrate Detector-Run eDPP Shell
646 | 10652 | Save Calibration on Exit
646 | 10652 | Run APEX
646 | 10652 | Check Spectrum
646 | 10652 | Check Mapping
646 | 10652 | Check Line Scanning
646 | 10652 | Save Aquired Spectrum, Mappimg & Line Scanning to the Image Folder
646 | 10652 | Save Calibration Data to the Licence Folder on the SmartSEM USB Stick
646 | 10652 | Save Calibration Data to the Image Folder
646 | 10653 | Manufacture of Detector
646 | 10653 | Number of Detectors
646 | 10653 | Detector Serial Number
646 | 10653 | Detector Resolution
646 | 10653 | Detector Area
646 | 10653 | Detector Installed
646 | 10653 | In/Out Stops Set
646 | 10653 | Isolation Check of Detector & Column >2MΩ
646 | 10653 | Install Remoting & Setup Network Connection
646 | 10653 | Install EDS Software
646 | 10653 | Install EDS Manager
646 | 10653 | Install EDS Licences
646 | 10653 | Setup Pulse Processor
646 | 10653 | Check Timing Cards
646 | 10654 | Install SmartSEM Licences
646 | 10654 | Instal MIL Redistribution
646 | 10654 | Fit MIL Dongle
646 | 10654 | Install Mineralogic Software
646 | 10654 | Calibrate I/O-SCAN card
646 | 10654 | Setup & Run Application
646 | 10654 | Check Results & Image Montage
646 | 10654 | 24 Hour Run
648 | 10553 | 30KV Res
648 | 10553 | 3KV Res
648 | 10553 | 1KV Res
648 | 10553 | 200V Res- Not measured
648 | 10554 | 2KV Res - 100um VP Aperture
648 | 10554 | 2KV Res - 400um Easy VP Aperture
648 | 10611 | 30KV Res
648 | 10611 | 3KV Res
648 | 10611 | 1KV Res
648 | 10611 | 200V Res- Not Measured
648 | 10612 | 2KV Res - 100um VP Apertutre
648 | 10612 | 2KV Res - 750um Easy VP Aperture
648 | 10613 | 2KV Res - 400um Easy Align Aperture =<16nm
648 | 10614 | 2KV Res - 750um Easy Align Aperture
648 | 10657 | 2KV - 100um VP Aperture
648 | 10657 | 2KV - 140um VP Aperture
648 | 10658 | 2KV - 100um VP Aperture
648 | 10658 | 2KV - 140um VP Aperture
650 | 10569 | Check thet the following coditions are set= I Beam=20uA, Fil I=1.95A, Long Fil. Life selected
649 | 10555 | BSD Resolution
649 | 10556 | 30KV Resolution - Easy VP
649 | 10557 | Resolution @ 7KV
649 | 10659 | 30KV Resolution
650 | 10561 | Complete & Sign Test Certificate
650 | 10566 | Safety Labels
650 | 10566 | Language
650 | 10566 | Serial Number
650 | 10566 | China RoHS Sticker
650 | 10567 | Install Customer Licences
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer to the Product Folder
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Add Licence to the SmartSEM USB Stick
650 | 10568 | Calibration File Added to SmartSEM USB Stick
650 | 10568 | Copy Calibration File to the Product Folder
650 | 10570 | Set Language in User Preferences
650 | 10571 | Set Language in User Preferences
650 | 10572 | Select Keyboard Language
650 | 10572 | Set Keyboard Language in Windows Control Panel
650 | 10572 | Destination of System
650 | 10574 | Easy VP/Easy Align Aperture Fitted
651 | 10576 | QA Engineer
651 | 10577 | Configuration Matches the Sales Order
651 | 10577 | Sales Order is the Latest Version
651 | 10578 | Correct Version of PC fitted
651 | 10578 | PC Configuration Image Taped to Side of PC
651 | 10580 | SmartSEM Version
651 | 10587 | Initialise Stage
651 | 10587 | Stage Movements operate Correctly
651 | 10588 | Collector Volts
651 | 10588 | Scintillator
651 | 10588 | With the beam blanked, select SE Detector, set contrast and brightness to 50%, scan speed 3 and ensure that there are no white streaks across the screen
651 | 10589 | Apertures in the Correct Position
651 | 10589 | Both apertures can be zero wobbled
651 | 10603 | X Stigmator
651 | 10603 | Y Stigmator
651 | 10607 | Check Cables
651 | 10607 | Check Electrical Connectors
651 | 10607 | Check Vacuum Connections
652 | 10619 | User manual language
652 | 10620 | Fit LH Bracket
652 | 10620 | Fit RH Bracket
652 | 10620 | Fit Rear Bracket
652 | 10620 | Fit Front Bracket
403 | 6935 | Record MAC address
403 | 6935 | Network access requested
653 | 10622 | Check LH Bracket
653 | 10622 | Check RH Bracket
653 | 10622 | Check Rear Bracklet
653 | 10622 | Check Front Bracket
654 | 10638 | Special Fitted
654 | 10638 | If Fitted, Checked by SCS Team Member
643 | 10492 | C3 Reversal-≤20um
643 | 10492 | 30KV-5KV
643 | 10492 | Image Shift
643 | 10492 | C3 Reversal-≤20µm
643 | 10492 | 30KV-5KV-≤20µm
643 | 10492 | Image Shift-≤3000nm
644 | 10508 | BeamSleeve 500um 30KV @ 8mm X- ≤35%
644 | 10508 | BeamSleeve 500um 30KV @ 8mm Y- ≤35%
644 | 10508 | BeamSleeve 500um Alignment- ≤70%
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
403 | 10465 | Record Product ID
403 | 10465 | Windows 10 Version
403 | 10465 | Disable Renesas USB Power Management
403 | 10467 | Install SmartSEM + ASAP Engine
403 | 10467 | SmartSEM Version
403 | 10467 | SmartSEM Service Pack
403 | 10467 | SmartSEM Hotfix
403 | 10467 | ASAP Engine Version
643 | 10492 | C3 Reversal-≤20µm
643 | 10492 | 30KV-5KV-≤20µm
643 | 10492 | Image Shift-≤3000nm
645 | 10519 | Rotate=20 deg
645 | 10519 | Rotate=45 deg
645 | 10519 | Rotate=90 deg
645 | 10519 | Rotate=180 deg
651 | 10897 | PCM and SCM read the same @ 500pA
644 | 10502 | 30KV @ 20mm X- ≤35%
644 | 10502 | 30KV @ 20mm Y- ≤35%
644 | 10502 | 2KV @ 20mm X- ≤35%
644 | 10502 | 2KV @ 20mm Y- ≤35%
644 | 10502 | 30KV @ 5mm X- ≤35%
644 | 10502 | 30KV @ 5mm Y- ≤35%
644 | 10502 | 2KV @ 5mm X- ≤35%
644 | 10502 | 2KV @ 5mm Y- ≤35%
644 | 10504 | VP 30KV @ 5mm (100um) X- ≤35%
644 | 10504 | VP 30KV @ 5mm (100um) Y- ≤35%
644 | 10505 | VP 30KV @ 5mm (140um) X- ≤35%
644 | 10505 | VP 30KV @ 5mm (140um) Y- ≤35%
644 | 10507 | EP 500um 30KV @ 5mm X- ≤35%
644 | 10507 | EP 500um 30KV @ 5mm Y- ≤35%
644 | 10507 | EP 500um Alignment- ≤55%
644 | 10507 | EP 1000um 30KV @ 5mm X- ≤35%
644 | 10507 | EP 1000um 30KV @ 5mm Y- ≤35%
644 | 10507 | EP 1000um Alignment- ≤55%
644 | 10509 | BeamSleeve 1000um 30KV @ 8mm X- ≤35%
644 | 10509 | BeamSleeve 1000um 30KV @ 8mm Y- ≤35%
644 | 10509 | BeamSleeve 1000um Alignment- ≤70%
645 | 10512 | 30KV @ 5mm X- ≤35%
645 | 10512 | 30KV @ 5mm Y- ≤35%
645 | 10513 | 30KV @ 5mm X- ≤35%
645 | 10513 | 30KV @ 5mm Y- ≤35%
645 | 10516 | X Axis
645 | 10516 | Y Axis
645 | 10516 | X+Y Axis
645 | 10516 | Y+X Axis
645 | 10516 | 10mm Square
645 | 10517 | Z Axis
645 | 10518 | X Axis
645 | 10518 | Y Axis
645 | 10520 | Tilt=5 deg
645 | 10520 | Tilt=10 deg
645 | 10520 | Tilt=20 deg
645 | 10520 | Tilt=45 deg
646 | 10529 | BSD Diode Pre-test
646 | 10529 | Select Type
646 | 10529 | Black Level Calibration
646 | 10529 | Check Diode Segment Function
646 | 10529 | Check All Gain ranges
646 | 10529 | Check Auto range
646 | 10529 | Check Auto Level
646 | 10529 | Setup and Check TOPO
646 | 10529 | Take 3 images - Atomic, 5pA, 2KV
647 | 10551 | Macro Passed
647 | 10551 | Gun Monitor Saved to Product Folder
647 | 10552 | Stability Run Passed
647 | 10552 | Gun Monitor Saved to Product Folder
648 | 10553 | 30KV Res- ≤2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 1KV Res- ≤20nm
648 | 10553 | 200V Res- Not measured
648 | 10554 | 2KV Res - 100um VP Aperture- ≤16nm
648 | 10554 | 2KV Res - 400um Easy VP Aperture- ≤16nm
648 | 10611 | 30KV Res- ≤2nm
648 | 10611 | 3KV Res- ≤6nm
648 | 10611 | 1KV Res- ≤9nm
648 | 10611 | 200V Res- Not Measured
648 | 10612 | 2KV Res - 100um VP Apertutre- ≤12nm
648 | 10612 | 2KV Res - 750um Easy VP Aperture- ≤12nm
648 | 10613 | 2KV Res - 400um Easy Align Aperture- ≤16nm
27 | 428 | Polarity is correct
648 | 10614 | 2KV Res - 750um Easy Align Aperture- ≤12nm
648 | 10657 | 2KV - 100um VP Aperture- ≤16nm
648 | 10657 | 2KV - 140um VP Aperture- ≤16nm
648 | 10658 | 2KV - 100um VP Aperture- ≤12nm
648 | 10658 | 2KV - 140um VP Aperture- ≤12nm
649 | 10555 | BSD Resolution- ≤4nm
649 | 10556 | 30KV Resolution - Easy VP- ≤3.4nm
649 | 10557 | Resolution @ 7KV- ≤9nm
649 | 10659 | 30KV Resolution- ≤5.5nm
645 | 10524 | Pumpdown Time
645 | 10525 | Pumpdown Time
651 | 10608 | System Vacuum Reading
651 | 10759 | Gun Vacuum Reading
651 | 10608 | System Vacuum Reading (≤3.00e-06 mbar)
651 | 10759 | Gun Vacuum Reading (≤5.00e-07 mbar)
643 | 10488 | Run EO Test
643 | 10488 | X Gun Tilt ≤50%
643 | 10488 | Y Gun Tilt ≤50%
643 | 10488 | X Gun Shift ≤50%
643 | 10488 | Y Gun Shift ≤50%
644 | 10496 | Accelerating Voltage=30KV
644 | 10496 | Specimen Current
644 | 10496 | Monitor Model
644 | 10496 | Serial Number
644 | 10496 | Cal. Expiry Date
644 | 10496 | No X-ray Leaks Detected
644 | 10496 | Complete & Sign X-Ray Certificate
644 | 10760 | Serial No. (or N/A)
644 | 10760 | No X-Ray Leakage Detected
645 | 10761 | Functional Check
645 | 10762 | Easy VP Mag. Cal Complete
646 | 10534 | Alignment of Ring (Using Jig)
646 | 10534 | Black Level Calibration
646 | 10534 | Functional Check (Dry)
646 | 10534 | Functional Check (Wet)
646 | 10537 | Functional Check
646 | 10537 | Function Image
646 | 10763 | Alignment of Ring (Using jig)
646 | 10763 | Black Level Calibration
646 | 10763 | Functional Check
646 | 10534 | Alignment of Ring (Using Jig)
646 | 10534 | Black Level Calibration
646 | 10534 | Functional Check
646 | 10652 | Type of Detector
646 | 10652 | Serial number
646 | 10652 | Fit Flange & Mount Detector
646 | 10652 | Install EDS PCB Into PC
646 | 10652 | Installation of DPP Box & Cables
646 | 10652 | Configue PC Network & Test
646 | 10652 | Disable Windows10 Fast Startup
646 | 10652 | Install Licences
646 | 10652 | Fit Dongle
646 | 10652 | APEX Software Version
646 | 10652 | Install APEX Software
646 | 10652 | Install Remoting
646 | 10652 | Setup RConfigure
646 | 10652 | Run API Test
646 | 10652 | Configure Imaging-Run SGShell
646 | 10652 | Calibrate Detector-Run eDPP Shell
646 | 10652 | Save Calibration on Exit
646 | 10652 | Run APEX
646 | 10652 | Check Spectrum
646 | 10652 | Check Mapping
646 | 10652 | Check Line Scanning
646 | 10652 | Save Aquired Spectrum, Mappimg & Line Scanning to the Image Folder
646 | 10652 | Save Calibration Data to the Licence Folder on the SmartSEM USB Stick
646 | 10652 | Save Calibration Data to the Image Folder
649 | 10555 | BSD Resolution ≤4nm
651 | 10589 | Apertures in the Correct Position
651 | 10589 | Both apertures can be zero wobbled
645 | 10512 | 30KV @ 5mm X- ≤35%
645 | 10512 | 30KV @ 5mm Y- ≤35%
645 | 10513 | 30KV @ 5mm X- ≤35%
645 | 10513 | 30KV @ 5mm Y- ≤35%
645 | 10513 | 30KV @ 5mm X- ≤35%
645 | 10513 | 30KV @ 5mm Y- ≤35%
645 | 10513 | 30KV @ 5mm X- ≤35%
645 | 10513 | 30KV @ 5mm Y- ≤35%
644 | 10496 | Accelerating Voltage=30KV
644 | 10496 | Specimen Current
644 | 10496 | Monitor Model
644 | 10496 | Serial Number
644 | 10496 | Cal. Expiry Date
644 | 10496 | No X-ray Leaks Detected
644 | 10496 | Complete & Sign X-Ray Certificate
645 | 10513 | 30KV @ 5mm X- ≤35%
645 | 10513 | 30KV @ 5mm Y- ≤35%
646 | 10548 | Connect Silver Box and Cables
646 | 10548 | Check Altas Test USB Stick Has The latest Version Of Software
646 | 10548 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
646 | 10548 | Request Licence File from Fibics
646 | 10548 | Install Licence File
646 | 10548 | Install Atlas Software
646 | 10548 | Calibrate SEM
646 | 10548 | Obtain Mosaic
646 | 10548 | Set ATLAS Rotation Offset
646 | 10548 | Save ATLAS Licence Document
646 | 10548 | Connect Silver Box and Cables
646 | 10548 | Check Altas Test USB Stick Has The latest Version Of Software
646 | 10548 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
646 | 10548 | Request Licence File from Fibics
646 | 10548 | Install Atlas Software
646 | 10548 | Install Licence File
646 | 10548 | Calibrate SEM
646 | 10548 | Obtain Mosaic
646 | 10548 | Set ATLAS Rotation Offset
646 | 10548 | Save ATLAS Licence Document
646 | 10866 | Install Hard Disk
646 | 10866 | Format Hard Disk (NTFS)
646 | 10866 | Hard Disk Is Seen In File Explorer
646 | 10866 | Copy A File To The Hard Disk, Then Delete It
646 | 10866 | Install Cloning Software
648 | 10875 | 2KV Res - 750um Easy Align Aperture- ≤16nm
644 | 10498 | Run up to 32KV and leave for 1 minute
644 | 10498 | Shutdown and run up to 32KV, Repeat 5 times
644 | 10498 | Run up to 32KV and leave for 15 minutes
651 | 10898 | Blanking Unit Installed
651 | 10898 | Blanking Cable Connected To Column Connection
646 | 10548 | Connect Silver Box and Cables
646 | 10548 | Check Altas Test USB Stick Has The latest Version Of Software
646 | 10548 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
646 | 10548 | Request Licence File from Fibics
646 | 10548 | Install Atlas Software
646 | 10548 | Install Licence File
646 | 10548 | Calibrate SEM
646 | 10548 | Obtain Mosaic
646 | 10548 | Set ATLAS Rotation Offset
646 | 10548 | Save ATLAS Licence Document
646 | 10653 | Number of Detectors
646 | 10653 | Detector Serial Number
646 | 10653 | Detector Resolution
646 | 10653 | Detector Area
646 | 10653 | Detector Installed
646 | 10653 | In/Out Stops Set
646 | 10653 | Isolation Check of Detector & Column >2MΩ
646 | 10653 | Install Remoting & Setup Network Connection
646 | 10653 | Install EDS Software
646 | 10653 | Install EDS Manager
646 | 10653 | Install EDS Licences
646 | 10653 | Setup Pulse Processor
646 | 10653 | Check Timing Cards
646 | 10869 | Camera Model
646 | 10869 | Install Camera Software
646 | 10871 | Install EMO Box
646 | 10871 | Check All Mains Outputs
646 | 10871 | Check Cut Off Switch
648 | 10614 | 2KV Res - 750um Easy Align Aperture- ≤12nm
648 | 10875 | 2KV Res - 750um Easy Align Aperture- ≤16nm
645 | 10513 | 30KV @ 5mm X- ≤35%
645 | 10513 | 30KV @ 5mm Y- ≤35%
646 | 10546 | SmartTouch Software Version
646 | 10546 | Install SmartTouch Software
646 | 10546 | Install Easy VP Aperture
646 | 10546 | Capture Sample Image Using the Nav Cam
646 | 10546 | Select Sample Holder and Sample Type
646 | 10546 | Set Yellow Line To 8.5mm WD
646 | 10546 | Sample Setup Complete
646 | 10546 | In the Image Tab, Check Sliders Function
646 | 10546 | Save Image - Photomode 1, Scan speed 8
646 | 10546 | Save Image - Photomode 2, Single Image
646 | 10546 | Save Image - Photomode 3, Scan Speed 2, Multiple Frames
646 | 10546 | Check Auto Saturation (W Filament Only)
646 | 10546 | Check Auto Stigmation
646 | 10546 | Check Manual Stigmation
646 | 10546 | Check Annotate
646 | 10546 | Check All Detectors Are Presaent
646 | 10546 | Select the Automate tab & Create Project as Zeiss
646 | 10546 | Setup 3 Point Calibration (Nav Cam )
646 | 10546 | Automation Run Completed
646 | 10546 | Select Browse Tab and Check Functions
645 | 10512 | 30KV @ 5mm X- ≤35%
645 | 10512 | 30KV @ 5mm Y- ≤35%
645 | 10513 | 30KV @ 5mm X- ≤35%
645 | 10513 | 30KV @ 5mm Y- ≤35%
645 | 10512 | 30KV @ 5mm X- ≤35%
645 | 10512 | 30KV @ 5mm Y- ≤35%
645 | 10513 | 30KV @ 5mm X- ≤35%
645 | 10513 | 30KV @ 5mm Y- ≤35%
645 | 10880 | 30KV @ 5mm X- ≤35%
645 | 10880 | 30KV @ 5mm Y- ≤35%
645 | 10512 | 30KV @ 5mm X- ≤35%
645 | 10512 | 30KV @ 5mm Y- ≤35%
645 | 10882 | 30KV @ 5mm X- ≤35%
645 | 10882 | 30KV @ 5mm Y- ≤35%
648 | 10554 | 2KV Res - 100um VP Aperture- ≤16nm
648 | 10554 | 2KV Res - 400um Easy VP Aperture- ≤16nm
648 | 10612 | 2KV Res - 100um VP Apertutre- ≤12nm
648 | 10881 | 2KV Res - 750um Easy VP Aperture- ≤12nm
645 | 10512 | 30KV @ 5mm X- ≤35%
645 | 10512 | 30KV @ 5mm Y- ≤35%
648 | 10881 | 2KV Res - 750um Easy VP Aperture- ≤12nm
648 | 10614 | 2KV Res - 750um Easy Align Aperture- ≤12nm
651 | 10899 | Video Input Kit Installed
31 | 310 | Actual
648 | 10612 | 2KV Res - 100um VP Apertutre- ≤12nm
648 | 10613 | 2KV Res - 400um Easy Align Aperture- ≤16nm
648 | 10614 | 2KV Res - 750um Easy Align Aperture- ≤12nm
648 | 10875 | 2KV Res - 750um Easy Align Aperture- ≤16nm
648 | 10881 | 2KV Res - 750um Easy VP Aperture- ≤12nm
645 | 10880 | 30KV @ 5mm X- ≤35%
645 | 10880 | 30KV @ 5mm Y- ≤35%
645 | 10880 | 30KV @ 5mm X- ≤35%
645 | 10880 | 30KV @ 5mm Y- ≤35%
645 | 10513 | 30KV @ 5mm X- ≤35%
645 | 10513 | 30KV @ 5mm Y- ≤35%
645 | 10513 | 30KV @ 5mm X- ≤35%
645 | 10513 | 30KV @ 5mm Y- ≤35%
651 | 10900 | Stage Moves To Correct Position When Arm Is Opened
651 | 10900 | Image Is Captured
651 | 10900 | Functional Check
644 | 10504 | VP 30KV @ 5mm (100um) X- ≤35%
644 | 10504 | VP 30KV @ 5mm (100um) Y- ≤35%
646 | 10543 | $NavCam Postion Setup
646 | 10543 | Image Navigation Setup
646 | 10543 | Functional Check
652 | 10621 | Check LH Bracket
652 | 10621 | Check RH Bracket
652 | 10621 | Check Rear Bracket
652 | 10621 | Check Front Bracket
652 | 10621 | Checked By (Engineer's Name)
645 | 10761 | Functional Check
651 | 10585 | C2D Is Selected When Going Into VP Mode
651 | 10585 | Image Brightness Changes When Changing Bias
651 | 10585 | Breakdown Of Bias Occurs at 100%
651 | 10590 | VPSE Is Selected When Going Into VP Mode
651 | 10590 | Image Brightness Changes When Changing Bias
651 | 10590 | Breakdown Of Bias Occurs at 100%
651 | 10591 | C2DX Is Selected When Going Into EP Mode
651 | 10591 | Image Brightness Changes When Changing Bias
651 | 10591 | Breakdown Of Bias Occurs at 100%
651 | 10592 | Check Diode Segments
651 | 10592 | Check Gain Ranges
651 | 10592 | Check Auto Range
651 | 10592 | Check TOPO
651 | 10884 | Check Aquire Spectrum
651 | 10884 | Check Line Scanning
651 | 10884 | Check Mapping
651 | 10884 | Check EDX Icon Is Present In All User Toolbars
651 | 10884 | Check Aquire Spectrum
651 | 10884 | Check Line Scanning
651 | 10884 | Check Mapping
651 | 10884 | Check EDX Icon Is Present In All User Toolbars
651 | 10884 | EDAX Calibration File Saved To SmartSEM USB Stick
651 | 10887 | Power Off Button Tested
651 | 10888 | Software Loaded
651 | 10889 | Hard Disk Is Seen In File Manager
651 | 10889 | Cloning Software Installed
651 | 10907 | Atlas Hardware Installed
651 | 10907 | Atlas Software Installed
651 | 10907 | Atlas Licence Saved To SmartSEM USB Stick
646 | 10534 | Alignment of Ring (Using Jig)
646 | 10534 | Black Level Calibration
646 | 10534 | Functional Check
646 | 10541 | Installation of Quiet Mode Vessel
646 | 10541 | Quiet Mode Fitted Selected in SEM Status
646 | 10541 | Compare On/Off Times With Graph
646 | 10541 | Pump Starts When Quiet Mode is Deselected
646 | 10541 | Run Gun Monitor-Turbo Power and System Vacuum
646 | 10544 | Installation
646 | 10544 | Functional Test
646 | 10651 | Installation
646 | 10651 | Functional Test
646 | 10652 | Type of Detector
646 | 10652 | Serial number
646 | 10652 | Fit Flange & Mount Detector
646 | 10652 | Install EDS PCB Into PC
646 | 10652 | Installation of DPP Box & Cables
646 | 10652 | Configue PC Network & Test
646 | 10652 | Disable Windows10 Fast Startup
646 | 10652 | Install Licences
646 | 10652 | Fit Dongle
646 | 10652 | APEX Software Version
646 | 10652 | Install APEX Software
646 | 10652 | Install Remoting
646 | 10652 | Setup RConfigure
646 | 10652 | Run API Test
646 | 10652 | Configure Imaging-Run SGShell
646 | 10652 | Calibrate Detector-Run eDPP Shell
646 | 10652 | Save Calibration on Exit
646 | 10652 | Run APEX
646 | 10652 | Check Spectrum
646 | 10652 | Check Mapping
646 | 10652 | Check Line Scanning
646 | 10652 | Save Aquired Spectrum, Mappimg & Line Scanning to the Image Folder
646 | 10652 | Save Calibration Data to the Licence Folder on the SmartSEM USB Stick
646 | 10652 | Save Calibration Data to the Image Folder
651 | 10903 | SmartPI Software Installed and Tested
651 | 10904 | Mineralogic Software Installed and Tested
645 | 10929 | Calibrate Stage Centre (Right Click on Stgae tab)
645 | 10929 | Tilt Centre Calibration (Located in Docking panel)
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
645 | 10513 | 30KV @ 5mm X- ≤35%
645 | 10513 | 30KV @ 5mm Y- ≤35%
650 | 10566 | Safety Labels
650 | 10566 | Language
650 | 10566 | Serial Number
650 | 10566 | China RoHS Sticker
403 | 10465 | Record Product ID
403 | 10465 | Windows 10 Version
403 | 10465 | Disable Renesas USB Power Management
652 | 10618 | Radiation Cert
652 | 10618 | Itemes removed
652 | 10618 | Items retained
652 | 10618 | Syteme and accessories packing list
654 | 11063 | Download Folder Images Deleted
654 | 11063 | Download Folder Images Deleted
654 | 11063 | Clear Internet History
403 | 10465 | Record Product ID
403 | 10465 | Windows 10 Version
403 | 10465 | Disable Renesas USB Power Management
27 | 429 | Angle =
644 | 10497 | Gun Chamber has been cleaned and inspected using UV light
644 | 10497 | EHT set socket well has been cleaned and inspected using UV light
644 | 10497 | Lab6 Firing Unit Fitted
644 | 10497 | Overnight Bakeout
644 | 10498 | Run up to 32KV and leave for 1 minute
644 | 10498 | Shutdown and run up to 32KV, Repeat 5 times
644 | 10498 | Run up to 32KV and leave for 15 minutes
648 | 10611 | 30KV Res- ≤2nm
648 | 10611 | 3KV Res- ≤6nm
648 | 10611 | 1KV Res- ≤9nm
648 | 10611 | 200V Res- Not Measured
651 | 10759 | Gun Vacuum Reading (≤5.00e-07 mbar)
403 | 11096 | Check SCS Folder/Sales Order
643 | 10495 | C3 Calibration
643 | 10495 | 30KV to 2KV Focus Check- ≤0.2mm change in WD.
644 | 10497 | Gun Chamber has been cleaned and inspected using UV light
644 | 10497 | EHT set socket well has been cleaned and inspected using UV light
644 | 10497 | Lab6 Firing Unit Fitted
644 | 10497 | Overnight Bakeout
644 | 10504 | VP 30KV @ 5mm (100um) X- ≤35%
644 | 10504 | VP 30KV @ 5mm (100um) Y- ≤35%
644 | 10507 | EP 500um 30KV @ 5mm X- ≤35%
644 | 10507 | EP 500um 30KV @ 5mm Y- ≤35%
644 | 10507 | EP 500um Alignment- ≤55%
644 | 10507 | EP 1000um 30KV @ 5mm X- ≤35%
644 | 10507 | EP 1000um 30KV @ 5mm Y- ≤35%
644 | 10507 | EP 1000um Alignment- ≤55%
645 | 10762 | Easy VP Mag. Cal Complete
648 | 10553 | 30KV Res- ≤2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 1KV Res- ≤20nm
648 | 10553 | 200V Res- Not measured
654 | 10638 | Special Fitted
654 | 10638 | If Fitted, Checked by SCS Team Member
646 | 11100 | Select Port
646 | 11100 | Check Camera for focus
646 | 11100 | Functional Check-Black/White + Colour Mode
646 | 11100 | Run wxPropview to Remove Defective Pixels
650 | 11097 | Check SCS Folder/Sales Order
403 | 11096 | Check SCS Folder/Sales Order
650 | 11097 | Check SCS Folder/Sales Order
403 | 10465 | Windows 10 Version
403 | 10465 | Disable Renesas USB Power Management
403 | 11096 | Check SCS Folder/Sales Order
403 | 11096 | SCS Number (If applicable)
652 | 10618 | Radiation Cert
652 | 10618 | Itemes removed
652 | 10618 | Items retained
652 | 10618 | System and accessories packing list
403 | 10465 | Windows 10 Version
403 | 10465 | Disable Renesas USB Power Management
652 | 10621 | Check LH Bracket
652 | 10621 | Check RH Bracket
652 | 10621 | Check Rear Bracket
652 | 10621 | Check Front Bracket
652 | 10621 | Checked By (Engineer's Name)
403 | 6935 | Record MAC address
403 | 6935 | Network access requested
646 | 10526 | Select Port
646 | 10526 | Check Camera for focus
646 | 10526 | Functional Check-Black/White + Colour Mode
646 | 10526 | Run wxPropview to Remove Defective Pixels
646 | 10543 | $NavCam Postion Setup
646 | 10543 | Image Navigation Setup
646 | 10543 | Functional Check
647 | 10551 | Macro Passed
647 | 10551 | Gun Monitor Saved to Product Folder
646 | 10537 | Functional Check
646 | 10537 | Function Image
646 | 10531 | Functional Check
646 | 10531 | Function Image
643 | 10487 | Gun Monitor Setup
643 | 10487 | Test No. 96 Selected
643 | 10487 | EHT Stable at 32KV - ≥15 minutes
643 | 10487 | 32KV Gun Monitor Traced Saved
644 | 10498 | Gun Monitor Setup
644 | 10498 | Test No. 96 Selected
644 | 10498 | Run up to 32KV and leave for 1 minute
644 | 10498 | Shutdown and run up to 32KV, Repeat 5 times
644 | 10498 | Run up to 32KV and leave for 15 minutes
644 | 10498 | 32KV Gun Monitor Traced Saved
644 | 10500 | Gun Monitor Setup
644 | 10500 | Test No. 96 Selected
644 | 10500 | Lab6 Firing Unit Run Up to 30KV
644 | 11195 | Check That 40uA is seen in the Gun Monitor
644 | 10498 | Gun Monitor Setup
644 | 10498 | Test No. 96 Selected
644 | 10498 | Run up to 32KV and leave for 1 minute
644 | 10498 | Shutdown and run up to 32KV, Repeat 5 times
644 | 10498 | Run up to 32KV and leave for 15 minutes
644 | 10498 | 32KV Gun Monitor Traced Saved
646 | 11200 | Select EVO25-L-ZTR. in Administaror as Stage Type
646 | 11200 | Remove Standard ZTR
646 | 11200 | Fit The Two Inserts to the XY Platform
646 | 11200 | Fit The Large Z Module
646 | 11200 | Remove the Two Transit Clamps and Fit the Linear Bearing Blanks
646 | 11200
646 | 11200 | In Stage Nav Settings, Set Spacer Offset and Stage Thickness to 0
646 | 11200 | Check Function of Touch Alarm
646 | 11200 | Initialise Stage
646 | 11200 | Check Full Movement of Each Axis on ZTR
646 | 11201 | X Axis
646 | 11201 | Y Axis
646 | 11201 | X+Y Axis
646 | 11201 | Y+X Axis
646 | 11201 | 10mm Square
646 | 11202 | Z Axis
646 | 11203 | Calibrate Stage Centre (Right Click on Stgae tab)
646 | 11203 | Tilt Centre Calibration (Located in Docking panel)
646 | 11204 | Rotate=20 deg
646 | 11204 | Rotate=45 deg
646 | 11204 | Rotate=90 deg
646 | 11204 | Rotate=180 deg
646 | 11205 | Tilt=5 deg
646 | 11205 | Tilt=10 deg
646 | 11205 | Tilt=20 deg
646 | 11205 | Tilt=45 deg
649 | 11206 | 30KV Res- ≤2nm
644 | 11195 | Check That 40uA is seen in the Gun Monitor
646 | 11100 | Select Port
646 | 11100 | Check Camera for focus
646 | 11100 | Functional Check-Black/White + Colour Mode
646 | 11100 | Run wxPropview to Remove Defective Pixels
646 | 11100 | Add Icon to Toolbar
646 | 11200 | Select EVO25-L-ZTR. in Administaror as Stage Type
646 | 11200 | Remove Standard ZTR
646 | 11200 | Fit The Two Inserts to the XY Platform
646 | 11200 | Fit The Large Z Module
646 | 11200 | Remove the Transit Clamp and Fit the Linear Bearing Blanks
646 | 11200
646 | 11200 | In Stage Nav Settings, Set Spacer Offset and Stage Thickness to 0
646 | 11200 | Check Function of Touch Alarm
646 | 11200 | Initialise Stage
646 | 11200 | Check Full Movement of Each Axis on ZTR
650 | 11097 | Check SCS Folder/Sales Order
654 | 10638 | Special Fitted
654 | 10638 | If Fitted, Checked by SCS Team Member
654 | 11197 | Remove Files
654 | 11197 | Remove Folders
651 | 10900 | Stage Moves To Correct Position When Arm Is Opened
651 | 10900 | Image Is Captured
651 | 10900 | Functional Check
642 | 11215 | Install EMO Box
642 | 11215 | Check All Mains Outputs
642 | 11215 | Check Cut Off Switch
646 | 10545 | Install SmartSEM Licences
646 | 10545 | Install MIL Redistribution
646 | 10545 | Fit MIL Dongle
646 | 10545 | Install SmartPI Software
646 | 10545 | I/O-SCAN Card Fitted (Bruker Only)
646 | 10545 | I/O-SCAN Card Calibrated
646 | 10545 | Setup & Run Application
646 | 10545 | Check Results & Image Montage
642 | 11215 | Install EMO Box
642 | 11215 | Check All Mains Outputs
642 | 11215 | Check Cut Off Switch
646 | 10545 | Install SmartSEM Licences
646 | 10545 | Install MIL Redistribution
646 | 10545 | Fit MIL Dongle
646 | 10545 | Install SmartPI Software
646 | 10545 | I/O-SCAN Card Fitted (Bruker Only)
646 | 10545 | I/O-SCAN Card Calibrated
646 | 10545 | Setup & Run Application
646 | 10545 | Check Results & Image Montage
646 | 10545 | Print Out SmartPI Check List (Link in Documents)
646 | 10545 | Install SmartSEM Licences
646 | 10545 | Install MIL Redistribution
646 | 10545 | Fit MIL Dongle
646 | 10545 | Install SmartPI Software
646 | 10545 | I/O-SCAN Card Fitted (Bruker Only)
646 | 10545 | I/O-SCAN Card Calibrated
646 | 10545 | Setup & Run Application
646 | 10545 | Check Results & Image Montage
646 | 10654 | Print Out Mineralogic Check List (Link in Documents)
646 | 10654 | Install SmartSEM Licences
646 | 10654 | Instal MIL Redistribution
646 | 10654 | Fit MIL Dongle
646 | 10654 | Install Mineralogic Software
646 | 10654 | Calibrate I/O-SCAN card
646 | 10654 | Setup & Run Application
646 | 10654 | Check Results & Image Montage
646 | 10654 | 24 Hour Run
644 | 10504 | VP 30KV @ 5mm (100um) X- ≤35%
644 | 10504 | VP 30KV @ 5mm (100um) Y- ≤35%
646 | 10545 | Correct PC - 347823-9024-660 Installed
646 | 10545 | Print Out SmartPI Check List (Link in Documents)
646 | 10545 | Install SmartSEM Licences
646 | 10545 | Install MIL Redistribution
646 | 10545 | Fit MIL Dongle
646 | 10545 | Install SmartPI Software
646 | 10545 | I/O-SCAN Card Fitted (Bruker Only)
646 | 10545 | I/O-SCAN Card Calibrated
646 | 10545 | Setup & Run Application
646 | 10545 | Check Results & Image Montage
646 | 10531 | Black Level Cal
646 | 10531 | C2D Bias Target Check
646 | 10531 | Gain Range Check
646 | 10531 | Auto Level Check
646 | 10531 | Auto Range Check
646 | 10531 | Function Image
646 | 10654 | Print Out Mineralogic Check List (Link in Documents)
646 | 10654 | Correct PC - Installed
646 | 10654 | Install SmartSEM Licences
646 | 10654 | Instal MIL Redistribution
646 | 10654 | Fit MIL Dongle
646 | 10654 | Install Mineralogic Software
646 | 10654 | Calibrate I/O-SCAN card
646 | 10654 | Setup & Run Application
646 | 10654 | Check Results & Image Montage
646 | 10654 | 24 Hour Run
644 | 11195 | Check That 40uA (Spec: ±3) is seen in the Gun Monitor (Beam I Monitor)
646 | 10546 | Chamberscope Fitted to Stage Door
646 | 10546 | SmartTouch Software Version
646 | 10546 | Install SmartTouch Software
646 | 10546 | Install Easy VP Aperture
646 | 10546 | Capture Sample Image Using the Nav Cam
646 | 10546 | Select Sample Holder and Sample Type
646 | 10546 | Set Yellow Line To 8.5mm WD
646 | 10546 | Sample Setup Complete
646 | 10546 | In the Image Tab, Check Sliders Function
646 | 10546 | Save Image - Photomode 1, Scan speed 8
646 | 10546 | Save Image - Photomode 2, Single Image
646 | 10546 | Save Image - Photomode 3, Scan Speed 2, Multiple Frames
646 | 10546 | Check Auto Saturation (W Filament Only)
646 | 10546 | Check Auto Stigmation
646 | 10546 | Check Manual Stigmation
646 | 10546 | Check Annotate
646 | 10546 | Check All Detectors Are Presaent
646 | 10546 | Select the Automate tab & Create Project as Zeiss
646 | 10546 | Setup 3 Point Calibration (Nav Cam )
646 | 10546 | Automation Run Completed
646 | 10546 | Select Browse Tab and Check Functions
644 | 10498 | Gun Monitor Setup
644 | 10498 | Test No. 96 Selected
644 | 10498 | Run up to 32KV and leave for 1 minute
644 | 10498 | Shutdown and run up to 32KV, Repeat 5 times
644 | 10498 | Run up to 32KV and leave for 15 minutes
644 | 10498 | 32KV Gun Monitor Traced Saved
644 | 11195 | Check That 40uA (Spec: ±3) is seen in the Gun Monitor (Beam I Monitor)
645 | 10512 | 30KV @ 5mm X- ≤35%
645 | 10512 | 30KV @ 5mm Y- ≤35%
645 | 10513 | 30KV @ 5mm X- ≤35%
645 | 10513 | 30KV @ 5mm Y- ≤35%
645 | 10519 | Rotate=20 deg
645 | 10519 | Rotate=45 deg
645 | 10519 | Rotate=90 deg
645 | 10519 | Rotate=180 deg
645 | 10520 | Tilt=5 deg
645 | 10520 | Tilt=10 deg
645 | 10520 | Tilt=20 deg
645 | 10520 | Tilt=45 deg
645 | 10761 | Functional Check
645 | 10762 | Easy VP Mag. Cal Complete
645 | 10880 | 30KV @ 5mm X- ≤35%
645 | 10880 | 30KV @ 5mm Y- ≤35%
645 | 10929 | Calibrate Stage Centre (Right Click on Stgae tab)
645 | 10929 | Tilt Centre Calibration (Located in Docking panel)
648 | 10875 | 2KV Res - 750um Easy Align Aperture- ≤16nm
648 | 10881 | 2KV Res - 750um Easy VP Aperture- ≤12nm
651 | 11269 | Check 30KV to 5KV image Shift is ≤20µm
644 | 10500 | Gun Monitor Setup
644 | 10500 | Test No. 96 Selected
644 | 10500 | Lab6 Firing Unit Run Up to 30KV
644 | 10500 | X Gun Tilt ≤50%
644 | 10500 | Y Gun Tilt ≤50%
644 | 10500 | X Gun Shift ≤50%
644 | 10500 | Y Gun Shift ≤50%
647 | 10552 | Conditions set - 30KV, 1nA, Long Filament Life
647 | 10552 | Gun Monitor Setup and Running
647 | 10552 | Stability Run Passed
647 | 10552 | Gun Monitor Saved to Product Folder
651 | 10603 | X Stigmator
651 | 10603 | Y Stigmator
651 | 10603 | X Gun Tilt ≤50%
651 | 10603 | Y Gun Tilt ≤50%
651 | 10603 | X Gun Shift ≤50%
651 | 10603 | Y Gun Shift ≤50%
645 | 10524 | Pumpdown Time
645 | 10525 | Pumpdown Time
40 | 451 | Calibrations done.
651 | 10608 | System Vacuum Reading (≤3.00e-06 mbar)
651 | 10759 | Gun Vacuum Reading (≤5.00e-07 mbar)
403 | 11272 | Predictive Service S/W installed on Customer PC.
646 | 10652 | Type of Detector
646 | 10652 | Serial number
646 | 10652 | Fit Flange & Mount Detector
646 | 10652 | Install EDS PCB Into PC
646 | 10652 | Installation of DPP Box & Cables
646 | 10652 | Configue PC Network & Test
646 | 10652 | Disable Windows10 Fast Startup
646 | 10652 | Install Licences
646 | 10652 | Fit Dongle
646 | 10652 | APEX Software Version
646 | 10652 | Install APEX Software
646 | 10652 | Install Remoting
646 | 10652 | Setup RConfigure
646 | 10652 | Run API Test
646 | 10652 | Configure Imaging-Run SGShell
646 | 10652 | Calibrate Detector-Run eDPP Shell
646 | 10652 | Save Calibration Report
646 | 10652 | Detector Resolution - ≥ 129 eV
646 | 10652 | Save Calibration on Exit
646 | 10652 | Run APEX
646 | 10652 | Check Spectrum
646 | 10652 | Check Mapping
646 | 10652 | Check Line Scanning
646 | 10652 | Save Aquired Spectrum, Mappimg & Line Scanning to the Image Folder
646 | 10652 | Save Calibration Data to the Licence Folder on the SmartSEM USB Stick
646 | 10652 | Save Calibration Data to the Image Folder
646 | 10652 | Type of Detector
646 | 10652 | Detector Serial number
646 | 10652 | DDP Module Serial Number
646 | 10652 | Fit Flange & Mount Detector
646 | 10652 | Install EDS PCB Into PC
646 | 10652 | Installation of DPP Box & Cables
646 | 10652 | Configue PC Network & Test
646 | 10652 | Disable Windows10 Fast Startup
646 | 10652 | Install Licences
646 | 10652 | Fit Dongle
646 | 10652 | APEX Software Version
646 | 10652 | Install APEX Software
646 | 10652 | Install Remoting
646 | 10652 | Setup RConfigure
646 | 10652 | Run API Test
646 | 10652 | Configure Imaging-Run SGShell
646 | 10652 | Calibrate Detector-Run eDPP Shell
646 | 10652 | Save Calibration Report
646 | 10652 | Detector Resolution - ≥ 129 eV
646 | 10652 | Save Calibration on Exit
646 | 10652 | Run APEX
646 | 10652 | Check Spectrum
646 | 10652 | Check Mapping
646 | 10652 | Check Line Scanning
646 | 10652 | Save Aquired Spectrum, Mappimg & Line Scanning to the Image Folder
646 | 10652 | Save Calibration Data to the Licence Folder on the SmartSEM USB Stick
646 | 10652 | Save Calibration Data to the Image Folder
651 | 10585 | C2D Is Selected When Going Into VP Mode
651 | 10585 | Image Brightness Changes When Changing Bias
651 | 10585 | Breakdown Of Bias Occurs at 100%
646 | 10652 | Type of Detector
646 | 10652 | Detector Serial number
646 | 10652 | DDP Module Serial Number
646 | 10652 | Fit Flange & Mount Detector
646 | 10652 | Install EDS PCB's Into PC
646 | 10652 | Installation of DPP Module & Cables
646 | 10652 | Configure PC Network & Test
646 | 10652 | Disable Windows10 Fast Startup
646 | 10652 | Install Licences
646 | 10652 | Fit Dongle
646 | 10652 | APEX Software Version
646 | 10652 | Install APEX Software
646 | 10652 | Install Remoting
646 | 10652 | Setup RConfigure
646 | 10652 | Run API Test
646 | 10652 | Configure Imaging-Run SGShell
646 | 10652 | Calibrate Detector-Run eDPP Shell
646 | 10652 | Save Calibration Report
646 | 10652 | Detector Resolution - ≥ 129 eV
646 | 10652 | Save Calibration on Exit
646 | 10652 | Run APEX
646 | 10652 | Check Spectrum
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected
646 | 10652 | Check Mapping
646 | 10652 | Check Line Scanning
646 | 10652 | Save Aquired Spectrum, Mappimg & Line Scanning to the Image Folder
646 | 10652 | Save Calibration Data to the Licence Folder on the SmartSEM USB Stick
646 | 10652 | Save Calibration Data to the Image Folder
644 | 10500 | Gun Monitor Setup
644 | 10500 | Test No. 96 Selected
644 | 10500 | Lab6 Firing Unit Run Up to 30KV
644 | 10500 | X Gun Tilt ≤50%
644 | 10500 | Y Gun Tilt ≤50%
644 | 10500 | X Gun Shift ≤50%
644 | 10500 | Y Gun Shift ≤50%
21 | 387 | Order revision in use or n/a
644 | 10500 | Gun Monitor Setup
644 | 10500 | Test No. 96 Selected
644 | 10500 | New Filament Selected
644 | 10500 | Long Fil. Life Selected
644 | 10500 | Lab6 Firing Unit Run Up to 30KV
644 | 10500 | X Gun Tilt ≤50%
644 | 10500 | Y Gun Tilt ≤50%
644 | 10500 | X Gun Shift ≤50%
644 | 10500 | Y Gun Shift ≤50%
28 | 292 | Cut off ok
28 | 292 | No Charging particles
644 | 11195 | Check That 20uA (Spec: ±3) is seen in the Gun Monitor (Beam I Monitor)
646 | 10653 | Number of Detectors
646 | 10653 | Detector Serial Number
646 | 10653 | Detector Resolution
646 | 10653 | Detector Area
646 | 10653 | Detector Installed
646 | 10653 | In/Out Stops Set
646 | 10653 | Isolation Check of Detector & Column >2MΩ
646 | 10653 | Install Remoting & Setup Network Connection
646 | 10653 | Install EDS Software
646 | 10653 | Install EDS Manager
646 | 10653 | Install Bruker Licences
646 | 10653 | Setup Pulse Processor
646 | 10653 | Check Timing Cards
646 | 10653 | Install Detector Data
646 | 10653 | Check Microscope driver
646 | 10653 | Setup SEM Conditions
646 | 10653 | Determine Analytical Working Distance
646 | 10653 | Calibrate Detector
646 | 10653 | Calibrate Image in Espirit
647 | 10552 | Conditions set - 30KV, 1nA, Long Filament Life
647 | 10552 | Gun Monitor Setup and Running
647 | 10552 | Stability Run Passed
647 | 10552 | Gun Monitor Saved to Product Folder
646 | 10652 | Type of Detector
646 | 10652 | Detector Serial number
646 | 10652 | DDP Module Serial Number
646 | 10652 | Fit Flange & Mount Detector
646 | 10652 | Install EDS PCB's Into PC
646 | 10652 | Installation of DPP Module & Cables
646 | 10652 | Configure PC Network & Test
646 | 10652 | Disable Windows10 Fast Startup
646 | 10652 | Install Licences
646 | 10652 | Fit Dongle
646 | 10652 | APEX Software Version
646 | 10652 | Install APEX Software
646 | 10652 | Install Remoting
646 | 10652 | Setup RConfigure
646 | 10652 | Run API Test
646 | 10652 | Configure Imaging-Run SGShell
646 | 10652 | Calibrate Detector-Run eDPP Shell
646 | 10652 | Save Calibration Report
646 | 10652 | Detector Resolution (measured using Cu) ≈ 129 eV
646 | 10652 | Save Calibration on Exit
646 | 10652 | Run APEX
646 | 10652 | Check Spectrum
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected
646 | 10652 | Check Mapping
646 | 10652 | Check Line Scanning
646 | 10652 | Save Aquired Spectrum, Mappimg & Line Scanning to the Image Folder
646 | 10652 | Save Calibration Data to the Licence Folder on the SmartSEM USB Stick
646 | 10652 | Save Calibration Data to the Image Folder
644 | 11195 | Check That 20uA (Spec: ±3) is seen in the Gun Monitor (Beam I Monitor)
643 | 10489 | Scintillator
643 | 10489 | Collector Volts
643 | 10489 | Check that the Signal Level in Line Mode is zero with Brightness=50% and Contrast =0%
645 | 10512 | 30KV @ 5mm X- ≤35%
645 | 10512 | 30KV @ 5mm Y- ≤35%
651 | 10579 | Rear Plinth Fan Running in Standby (yellow button)
651 | 10579 | Rear Plinth Fan Running in Operate (green button)
651 | 10579 | EO Fans Running in Operate (green button)
651 | 10579 | PC Runs Up
643 | 10489 | Scintillator
643 | 10489 | Collector Volts
643 | 10489 | Check that the Signal Level in Line Mode is at Black Level with Brightness=50% and Contrast =0%
643 | 10489 | Check that the Signal Level in Line Mode is at White Level with Brigfhtness =62% and Contrast =0%
647 | 10552 | Conditions set - 20KV, 1nA, Long Filament Life
647 | 10552 | Gun Monitor Setup and Running
647 | 10552 | Stability Run Passed
647 | 10552 | Gun Monitor Saved to Product Folder
654 | 11063 | Download Folder Images Deleted
654 | 11063 | Clear Internet History
646 | 10866 | Install Hard Disk
646 | 10866 | Format Hard Disk (NTFS)
646 | 10866 | Hard Disk Is Seen In File Explorer
646 | 10526 | Number of Chamberscope's on Order
646 | 10526 | Chamberscope 1 - Select Port
646 | 10526 | Check Camera for focus
646 | 10526 | Functional Check-Black/White + Colour Mode
646 | 10526 | Run wxPropview to Remove Defective Pixels
646 | 10526 | Chamberscope 2 - Select Port or Not Fitted
646 | 10526 | Chamberscope 2 - Check Camera 2 for Focus
646 | 10526 | Chamberscope 2 - Functional Check-Black/White + Colour Mode
646 | 10526 | Chamberscope 2 - Run wxPropview to Remove Defective Pixels
646 | 10526 | Chamberscope 2 - Add Icon to Toolbar
643 | 10486 | Jacking Feet Set
643 | 10486 | Air Feet Inflated
643 | 10486 | Rear Plinth Fan is Running in Operate and Standby
21 | 389 | Tick option 'Low KV Objective Cap'
646 | 11100 | Select Port
646 | 11100 | Check Camera for focus
646 | 11100 | Functional Check-Black/White + Colour Mode
646 | 11100 | Run wxPropview to Remove Defective Pixels
646 | 11100 | Add Icon to Toolbar
651 | 11343 | Fan Running in Standby (yellow button)
651 | 11343 | Fan Running in Operate (green button)
650 | 10574 | Easy VP/Easy Align Aperture Fitted
403 | 10465 | Windows 10 Version
403 | 10465 | Disable Renesas USB Power Management
651 | 10578 | Correct Version of PC fitted
651 | 11394 | X Stigmator
651 | 11394 | Y Stigmator
651 | 11394 | X Gun Tilt ≤50%
651 | 11394 | Y Gun Tilt ≤50%
651 | 11394 | X Gun Shift ≤50%
651 | 11394 | Y Gun Shift ≤50%
651 | 10608 | System Vacuum Reading (≤3.00e-06 mbar)
651 | 10759 | Gun Vacuum Reading (≤5.00e-07 mbar)
651 | 11397 | Change in Working Distance ≤0.2mm
22 | 260 | Column traveller completed?
22 | 260 | Plinth traveller completed?
22 | 260 | System Configuration correct?
22 | 260 | Build quality satisfactory?
646 | 10547 | ZEN Version
646 | 10547 | Correct Version of Windows 10 Installed
646 | 10547 | Install ZEN
646 | 10547 | Install Licence
646 | 10547 | Install Remote API Client
646 | 10547 | Functional Test
646 | 10543 | $NavCam Postion Setup
646 | 10543 | Image Navigation Setup
646 | 10543 | Functional Check
646 | 10545 | Correct PC - 347823-9024-660 Installed
646 | 10545 | Print Out SmartPI Check List (Link in Documents)
646 | 10545 | Install SmartSEM Licences
646 | 10545 | Install MIL Redistribution
646 | 10545 | Fit MIL Dongle
646 | 10545 | SmartPI Software Version
646 | 10545 | Install SmartPI Software
646 | 10545 | I/O-SCAN Card Fitted (Bruker Only)
646 | 10545 | I/O-SCAN Card Calibrated
646 | 10545 | Setup & Run Application
646 | 10545 | Check Results & Image Montage
646 | 10653 | Number of Detectors
646 | 10653 | Detector Serial Number
646 | 10653 | Detector Resolution
646 | 10653 | Detector Area
646 | 10653 | Detector Installed
646 | 10653 | In/Out Stops Set
646 | 10653 | Isolation Check of Detector & Column >2MΩ
646 | 10653 | Install Remoting & Setup Network Connection
646 | 10653 | Bruker Esprit Software Version
646 | 10653 | Install Bruker Esprit Software
646 | 10653 | Bruker EDS Manager Version
646 | 10653 | Install EDS Manager
646 | 10653 | Install Bruker Licences
646 | 10653 | Setup Pulse Processor
646 | 10653 | Check Timing Cards
646 | 10653 | Install Detector Data
646 | 10653 | Check Microscope driver
646 | 10653 | Setup SEM Conditions
646 | 10653 | Determine Analytical Working Distance
646 | 10653 | Calibrate Detector
646 | 10653 | Calibrate Image in Espirit
650 | 10566 | Safety Labels
650 | 10566 | Language
650 | 10566 | China RoHS Sticker
654 | 11420 | Earth resistance – measured value must be < 0.1 Ω
654 | 11420 | Insulation resistance – measured value must be > 1.0M Ω
654 | 11420 | Leakage – measured value must be below specified on PAT test print out conducted in assembly.
654 | 11420 | Test Certificate Completed & Signed
22 | 261 | Carry out a 3hrs Plasma clean (Remove all samples from chamber!)
40 | 452 | Calibrations done.
654 | 11420 | Earth resistance – measured value must be < 0.1 Ω
654 | 11420 | Insulation resistance – measured value must be > 1.0M Ω
654 | 11420 | Leakage – measured value must be below specified on PAT test print out conducted in assembly.
654 | 11420 | Test Certificate Completed & Signed
642 | 10481 | Fit Link on Vacuum PCB
642 | 10481 | Gun Chamber Valve Motorised=Yes
642 | 10481 | Check that when Ion Pump is Pumping, the Vacuum is Improving.
642 | 10481 | Gun/ Chamber Valve Closes ≈ 2.0e-06
643 | 11424 | Remove EO/Stage PSU Spacers
643 | 11424 | Remove Vac/ASAP PSU Spacers
652 | 11421 | Fit EO/Stage PSU Spacers
652 | 11421 | Fit Vac/ASAP PSU Spacers
21 | 387 | Order revision in use or 'n/a'
24 | 273 | Level Frame
24 | 273 | Check pendulum height (with spacer)
24 | 273 | Set up ‘float’ height
24 | 273 | Check inlet / outlet valves
24 | 273 | Set up docking pins
21 | 389 | Tick option 'Low KV Objective Cap'
20 | 252 | PC S/N
20 | 252 | Hard disk Image (Version)
20 | 252 | PC/Model (SAP No.)
20 | 252 | Record Operating System
20 | 252 | Activate Windows
20 | 252 | Record Windows Product Key
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them.
20 | 252 | Check for available SmartSEM hotfixes and install them.
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Approved Datazone Loaded
20 | 254 | Language
20 | 425 | Record MAC adress
20 | 425 | Network access requested
21 | 255 | Air (Pressure = 6 Bar)
21 | 255 | Water
21 | 255 | Mains
21 | 255 | Safety Earth
21 | 255 | N2
21 | 256 | Save Full System Checker report
21 | 257 | In Factory Tab: Omit Dongle Check
21 | 257 | In Factory Tab: Enter S/N
21 | 257 | In License Tab: Set All Licenses
21 | 258 | Confirm there are no Errors in EMserver (Apart from Dongle check and Stage Init)
21 | 259 | Check camera Focus and Rotation
21 | 259 | Eliminate camera image tilt
21 | 385 | Select correct System type (HV / VP)
21 | 385 | If Pressure Stage is fitted Tick box
22 | 262 | Carry out overnight ‘Full service Bakeout’ and note actual duration.
22 | 262 | Check thermocouple readings and iron pump plates heat up.
22 | 264 | Tip No.
22 | 265 | Set +/-21V Supplies to +/- 22.5V
22 | 266 | Use the 35kV External supply for 30mins.
22 | 267 | Use Autokathode tool if available.
22 | 268 | Use Autokathode tool if available
22 | 421 | Select "Tools and User Preferences": Set "Expert Gun Mode = Yes"
23 | 269 | Run macro 'ESB Runup 3h pumpen & 9h hochfahren.MLF'
28 | 291 | Mechanical Alignment ok
28 | 291 | Column clamping secure
25 | 275 | Are Gun Head clamping screws secure?
25 | 275 | Are Condensor clamping screws secure?
25 | 275 | Save Emission image @1Kv (cond off)
25 | 275 | Save Emission image @10Kv (cond off)
25 | 276 | Actual =
25 | 277 | Actual =
25 | 279 | Fil I
25 | 279 | Ext V
25 | 426 | Water pipes / connectors clear.
25 | 427 | Mech. align the column using the Merlin Alignment Wizard
26 | 281 | Type in measured value (Limit <=30um)
26 | 282 | Type in measured value (limit: <=15um)
26 | 283 | Check for ESB charging and Flashing.
26 | 396 | Type in measured value (Limit: <= 50um)
27 | 284 | 20nA-Emmision rotates CW with I increase
27 | 285 | No image rotation with WD change
27 | 286 | Angle =
27 | 289 | C2 Calibration Value
27 | 289 | C3 Calibration Value
27 | 404 | 100nA-Emmision rotates ACW with I increase
27 | 406 | Alignment done and confirmed.
27 | 430 | C1 Cal =
27 | 430 | C2 Cal =
27 | 430 | C3 Cal =
27 | 431 | Anode Diameter =
27 | 434 | Shift is < 10um over the whole IProbe range
27 | 435 | Focus deviation within 30um
27 | 436 | Focus deviation within 30um
28 | 290 | Gun running to Datasheet
28 | 290 | Gun Vacuum acceptable (EHT on/off)
28 | 290 | Nominal Ext V acceptable
28 | 293 | ESB concentricity in spec
28 | 293 | ESB-Inlens concenrticity in spec
28 | 294 | Run short image jump macro and analyse images
28 | 409 | Gun X
28 | 409 | Gun Y
28 | 409 | Aper X
28 | 409 | Aper Y
29 | 295 | Marco successfully completed.
29 | 295 | Check for Drift, Charging, ESB issues
30 | 296 | Calibration wizard completed.
30 | 297 | Check Mag cal errors are <1% in X & Y axes in all 3 Mag Ranges
30 | 298 | Macro (runs 30 mins.) completed.
30 | 299 | Macro successfully completed.
30 | 299 | Save Image to Image directory.
30 | 300 | X =
30 | 300 | Y =
30 | 302 | Wizard successfully completed.
30 | 303 | Actual
30 | 305 | Holder Height
30 | 305 | Holder Rotation
30 | 305 | Select Customer Sample Holders
30 | 305 | Confirm Holder touch alarm works (Alarm tone and error message)
31 | 306 | Actual Gun Axis X =
31 | 306 | Actual Gun Axis Y =
31 | 307 | Actual Align 1 X =
31 | 307 | Actual Align 1 Y =
31 | 308 | Actual 1kV X Axis =
31 | 308 | Actual 1kV Y Axis =
31 | 308 | Actual 20kV X Axis =
31 | 308 | Actual 20kV Y Axis =
31 | 309 | Difference between Eht on and off < 1pA.
32 | 311 | Actual =
32 | 312 | Actual @ 1kV WD5 =
32 | 312 | Actual @ 20kV WD2 =
32 | 314 | Normal mode
32 | 314 | High Current mode
32 | 444 | Actual @ 1kV WD5 =
32 | 444 | Actual @ 20kV WD2 =
32 | 445 | Actual @ Normal Mode =
32 | 445 | Actual @ High Current Mode =
33 | 317 | No excessive stigs etc.
33 | 446 | No excessive stigs etc.
34 | 318 | Actual =
36 | 321 | Analytic mode, 10kV, WD2.Max dev= Measured value +3pa of slider
36 | 322 | Analytic mode, 10kV, WD2. Max dev= Measured value +/-15% of slider.Note: Use gun aligns to maximise probe
36 | 323 | Analytic mode, 10kV, WD2. Max dev= Measured value +/-15% of slider
36 | 324 | High res mode, 10kV, WD2. Max dev= Measured value +/-15% of slid-er.
36 | 326 | <7pa +/-2pa
37 | 328 | Tablefiller successfully completed.
37 | 328 | Manual alignment completed.
37 | 329 | Stig: 30μm, All EHT’s is ≤ 20%
37 | 329 | Stig: Off Axis Aper, All EHT’s is ≤ 30%
37 | 329 | Aper/Gun 30μm, All EHT’s is ≤ 60%
37 | 330 | Stig: 20μm, 10kV is ≤ 35%
37 | 330 | Stig: Off Axis Aper, 10kV is ≤ 45%
37 | 330 | Stig: 20μm 1kV is ≤ 20%
37 | 330 | Stig: Off Axis Aper, 1kV is ≤ 30%
37 | 330 | Aper/Gun: 20μm, All EHT’s is ≤ 60%
37 | 415 | Operation 5310 submitted
37 | 415 | Operation 5350 submitted
38 | 333 | 30kV Max Ext Voltage =
38 | 333 | Complete X-Ray certificate
38 | 334 | Fil I:
38 | 334 | Ext V:
38 | 335 | Note specimen current at 20kV
38 | 336 | Alignment done.
38 | 337 | Actual (Limit <= 300nm) =
38 | 422 | Note specimen current at 20kV
23 | 269 | Run macro 'ESB Runup 3h pumpen & 9h hochfahren.MLF'
39 | 339 | Actual =
39 | 339 | After conditioning confirm there are no EMServer error messages
39 | 339 | Save Gun Monitor as Probestability.MN2
40 | 340 | 5 – 30Pa Variable Pressure
40 | 340 | VP image taken (30kV, 200kX)
40 | 340 | Ensure Penning fitted is Type ‘PKR’
40 | 340 | No arcing of collector up to 290v @ 30Pa
40 | 341 | Test all accessory options as per customer order and record under options fitted
40 | 342 | Carry out any relevant additional tests – Refer to Man Eng if unsure.
41 | 343 | Before Cladding the system check that all column connections aresecurely fastened.
41 | 343 | Strip down Aluminium support frame and clean tapped holes. Re-assemble frame and secure screws with Loctite 222.
41 | 343 | Ensure Cladding geometry is correct and all parts aligned.
42 | 344 | Actual X-Axis
42 | 344 | Actual Y-Axis
43 | 345 | For Gemini 300: Aper 1 Normal current; for Gemini 500: Aper 1 High current
43 | 346 | 1kV
43 | 346 | 15kV
43 | 347 | 1kV
43 | 347 | 15kV
43 | 410 | 1kV ESB
43 | 417 | Operation 7000 sumitted
43 | 417 | Operation 8000 submitted
44 | 348 | Chamber Door Interlock
44 | 348 | Service Switch
44 | 348 | Lens De-activation
44 | 348 | Detector De-activation
44 | 348 | EHT + IPPSU De-activation
44 | 348 | A/L I/L test (if Fitted)
44 | 348 | EMO I/L test (if Fitted)
44 | 349 | Carry out safety tests using the Secutest S111 pat Tester.
44 | 349 | Fill out completely the Product Safety Tests form in Section 13.
44 | 350 | Actual
44 | 351 | Actual
44 | 353 | EMserver.mdb
44 | 353 | Customer Images
44 | 353 | License File
44 | 354 | Check Installed Licences Against Sales Order
44 | 355 | Attach label to chamber
44 | 356 | Check System And Dongle S/No. Match In Service Panel
44 | 356 | Un-tick Omit Dongle In Factory Panel
44 | 357 | Macro successfully completed.
44 | 357 | Check for Drift, Charging, ESB issues
44 | 358 | All verification & Safety tests done and within threshold limits defined
44 | 358 | Sales Order is latest version
44 | 358 | All items on Sales order accounted for
44 | 359 | Remove all Test Macro Shortcuts
44 | 359 | Remove all Test Program Shortcuts
44 | 359 | Set up Desktop icons (both accounts)
44 | 359 | Remove Proxy Server
44 | 359 | Reset WSUS Settings
45 | 362 | Em Server.mdb
45 | 362 | Em Server.txt Log
45 | 362 | Optiprobe.txt Log
45 | 362 | Smart Backup.log
45 | 362 | L500.ini
45 | 362 | License file.xml
45 | 362 | Registry file.reg
45 | 362 | CalibSc.dat
45 | 363 | Em Server.mdb
45 | 363 | License File
45 | 363 | Saved Registry
45 | 363 | L500.ini
45 | 364 | Copy all sub directories and files from the Image directory
45 | 365 | Language:
45 | 366 | Nominal Fil I:
45 | 366 | Nominal extractor V:
45 | 367 | High Res Fil I:
45 | 367 | High Res Extractor V:
45 | 368 | Analytic Fil I
45 | 368 | Analytic Extractor V:
45 | 411 | Images and files from the user directory deleted.
45 | 411 | Recycle bin cleared.
45 | 413 | Go to 'Select Tools and Preferences' and set 'Expert Gun Mode = No'
46 | 369 | Create Directory C:\Support\Manuals
46 | 369 | Create Shortcut to above Directory on Desktops, called ‘Carl Zeiss User Documentation’.
46 | 370 | Completed as per packing procedure.
46 | 371 | X-Ray Certificate
46 | 371 | Quality Certificate
46 | 371 | Quality Feedback Form
46 | 371 | DENKA Data Sheet
46 | 371 | Options Removed packing List
46 | 371 | Options Retained packing List
46 | 371 | Standard Packing List
46 | 371 | PC Inferface assignment
46 | 371 | Declaration of conformity
46 | 371 | Installation report
46 | 371 | Acceptance test record
46 | 371 | Licence Generator report
46 | 373 | Confirmed.
46 | 374 | Completed.
46 | 419 | C:\Support\Manuals
46 | 419 | Customer USB documentation folder
46 | 420 | Copy SmartSEM instruction manual from Customer USB to C:\Support\Manuals
40 | 447 | Calibrations done.
40 | 448 | Calibrations done.
40 | 449 | Calibrations done.
40 | 450 | Calibrations done.
40 | 453 | Calibrations done.
20 | 253 | Enter the FIB Extension number using EEPromUpdater.exe
25 | 279 | Fil I
25 | 279 | Ext V
26 | 281 | Type in measured value (Limit <= 30um)
26 | 282 | Type in measured value (limit: <= 15um)
30 | 456 | Eucentricity set up
36 | 323 | Analytic mode, 10kV, WD2. Max dev= Measured value +/-15% of slider
36 | 324 | Max dev =
36 | 326 | Actual < 7pA +/- 2pA
38 | 334 | Fil I =
38 | 334 | Ext V =
38 | 335 | Specimen current @ 20kV =
40 | 447 | Leack check
40 | 447 | X-Ray check
40 | 447 | Analyze Segments
40 | 447 | Gain Ranges
40 | 447 | CZ Detector Cal
40 | 448 | Leack check
40 | 448 | X-Ray check
40 | 448 | Analyze Segments
40 | 448 | Gain Ranges
40 | 448 | CZ Detector Cal
40 | 449 | Leack check
40 | 449 | X-Ray check
40 | 449 | Analyze Segments
40 | 449 | Gain Ranges
40 | 449 | Pneumatics
40 | 449 | Alignment
40 | 449 | Vac Degradation < 3x10^-15 mbar
40 | 449 | CZ Detector Cal
40 | 449 | Z Adjustment
40 | 450 | Analyze Segments
40 | 450 | Gain Ranges
40 | 450 | Alignment
40 | 450 | CZ Detector Cal
40 | 451 | Leack check
40 | 451 | X-Ray check
40 | 451 | Analyze Segments
40 | 451 | Gain Ranges
40 | 451 | Pneumatics
40 | 451 | Alignment
40 | 451 | Vac Degradation < 3x10^-15 mbar
40 | 451 | CZ Detector Cal
40 | 451 | Z Adjustment
40 | 452 | Leack check
40 | 452 | X-Ray check
40 | 452 | Analyze Segments
40 | 452 | Gain Ranges
40 | 452 | Pneumatics
40 | 452 | Alignment
40 | 452 | Stage Safe Position
40 | 452 | Vac Degradation < 3x10^-15 mbar
40 | 452 | CZ Detector Cal
40 | 452 | Store Holder Position 1
44 | 454 | Interlock tested and working
40 | 453 | Leack check
40 | 453 | X-Ray check
40 | 453 | Analyze Segments
40 | 453 | Gain Ranges
40 | 453 | Pneumatics
40 | 453 | Alignment
40 | 453 | Stage Safe Position
40 | 453 | Vac Degradation < 3x10^-15 mbar
40 | 453 | CZ Detector Cal
40 | 453 | Store Holder Position 1
40 | 1550 | Leack check
40 | 1550 | X-Ray check
40 | 1550 | Function check
40 | 1551 | Leack check
40 | 1551 | X-Ray check
40 | 1551 | Grid
40 | 1551 | Check Cut Off
40 | 1552 | Leack check
40 | 1552 | X-Ray check
40 | 1552 | Collector
40 | 1552 | Scintillator
40 | 1553 | Leack check
40 | 1553 | X-Ray check
40 | 1553 | Collector
40 | 1554 | Leack check
40 | 1554 | X-Ray check
40 | 1554 | Function Check
40 | 1555 | Leack check
40 | 1555 | X-Ray check
40 | 1555 | Pneumatics
40 | 1555 | Alignment
40 | 1555 | Repeatability
40 | 1555 | Bias Voltage
40 | 1555 | Vac Degradation < 3x10^-15 mbar
40 | 1555 | NanoVP Operation
40 | 1555 | XVP Operation
40 | 1556 | Functionality
40 | 1556 | Language
40 | 1557 | Leack check
40 | 1557 | X-Ray check
40 | 1557 | CCD
40 | 1557 | Focus
40 | 1558 | Leack check
40 | 1558 | X-Ray check
40 | 1558 | CCD
40 | 1558 | Focus
40 | 1559 | Leack check
40 | 1559 | X-Ray check
40 | 1559 | Macros loaded/tested
40 | 1559 | $exchange
40 | 1559 | Rod Position
40 | 1559 | Front Panel Control
40 | 1560 | Leack check
40 | 1560 | Disable/Enable
40 | 1560 | Save Vac Trace
40 | 1561 | Leack check
40 | 1561 | X-Ray check
40 | 1561 | EHT Cal
40 | 1561 | Spark Alarm
40 | 1561 | Res Prep
40 | 1561 | Resolution
40 | 1562 | Leack check
40 | 1562 | X-Ray check
40 | 1562 | Run Clean Cycle
40 | 1563 | All Axes Funtion
40 | 1563 | Break Button Function
40 | 1564 | Leack check
40 | 1564 | X-Ray check
40 | 1564 | Align Needle
40 | 1564 | Pneumatics
40 | 1564 | Pressure Test
40 | 1565 | Sample Cleaning (O2)
43 | 346 | 1kV =
43 | 346 | 15kV =
43 | 347 | 1kV
43 | 347 | 15kV
43 | 417 | Operation 7000 sumitted
43 | 417 | Operation 8000 submitted
44 | 348 | Chamber Door Interlock
44 | 348 | Service Switch
44 | 348 | Lens De-activation
44 | 348 | Detector De-activation
44 | 348 | EHT + IPPSU De-activation
45 | 360 | Factory acceptance test completed
45 | 365 | Language
45 | 457 | Factory acceptance test completed
46 | 371 | X-Ray Certificate
46 | 371 | Quality Certificate
46 | 371 | Quality Feedback Form
46 | 371 | DENKA Data Sheet
46 | 371 | Options Removed packing List
46 | 371 | Options Retained packing List
46 | 371 | Standard Packing List
46 | 371 | PC Inferface assignment
46 | 371 | Declaration of conformity
46 | 371 | Installation report
46 | 371 | Acceptance test record
46 | 371 | Licence Generator report
40 | 1551 | Leack check
40 | 1551 | X-Ray check
40 | 1551 | Grid
40 | 1551 | Check Cut Off
32 | 314 | Normal mode
32 | 314 | High Current mode
37 | 329 | Stig: 30μm, All EHT’s is ≤ 20%
37 | 329 | Stig: Off Axis Aper, All EHT’s is ≤ 30%
37 | 329 | Aper/Gun 30μm, All EHT’s is ≤ 60%
37 | 330 | Stig: 20μm, 10kV is ≤ 35%
37 | 330 | Stig: Off Axis Aper, 10kV is ≤ 45%
37 | 330 | Stig: 20μm 1kV is ≤ 20%
37 | 330 | Stig: Off Axis Aper, 1kV is ≤ 30%
37 | 330 | Aper/Gun: 20μm, All EHT’s is ≤ 60%
46 | 1567 | Remove the sample holder from the system.
46 | 1567 | Remove Docking Brackets and fit Transit clamps.
46 | 1567 | Airfeet deflated and free to move.
46 | 1567 | Backing lines re-configured so that PVSV valve is at-tached to flexi hose. Secure valve to frame.
46 | 1567 | If applicable, ensure that the A/L push rod is removed and blank fitted in its place. Ensure A/L is pumped down before powering down system.
46 | 1567 | Select 'Partial vent on standby' in Smartsem.
46 | 1568 | Power system down.
46 | 1568 | Blow water out of the cooling system.
46 | 1568 | Chamber secured to vibration plate
46 | 1568 | Secure Airfeet with foam blocks.
46 | 1568 | Fit foam blocks around column cladding.
46 | 1568 | Secure Turbo pump with foam blocks.
46 | 1568 | Fit Covers to all unused connectors.
46 | 1568 | Mains and Earth cable stowed on Mains unit.
46 | 1568 | Fit shipping Eyes to transit blocks.
46 | 1568 | Confirm all column cladding is secure and screws Loctited in place.
46 | 1568 | Put thread lock on pendulum screws and transit block Screws.
46 | 1568 | All doors closed and secured.
46 | 1568 | All plinth panels removed and stored in transit trolley.
46 | 1568 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 1568 | Print shipping stickers.
46 | 1568 | Fit M8 screw to secure stage door to chamber.
80 | 1570 | Reading =
80 | 1586 | Reading =
81 | 1571 | Measured value =
81 | 1587 | Measured value =
81 | 1588 | Measured value =
81 | 1589 | Measured value =
81 | 1590 | Filament Age
82 | 1591 | Take and save images at Mag 200, 2kX and 35kX.
82 | 1591 | Mag is within +/- 1% for each.
82 | 1591 | Orthognality is within +/- 1 deg for each.
82 | 1592 | Take and save images at Mag 200, 2kX and 35kX.
82 | 1592 | Mag is within +/- 2% for each.
82 | 1592 | Orthognality is within +/- 1 deg for each.
84 | 1574 | 5 and 30 Pa without beamsleeve.
84 | 1574 | 5 and 150 Pa in NanoVP Mode
84 | 1574 | 5 and 500 Pa in XVP Mode
84 | 1601 | Ventilation of chamber possible. Confirm Gun does not switch off.
84 | 1601 | While system is vented check touch alarm works.
102 | 1576 | Function checked in VP mode, signal intensity changes with VP pressure.
102 | 1576 | Signal intensity changes with collector voltage.
102 | 1616 | Grid Control working
102 | 1616 | ESB Concentricity in Spec
102 | 1617 | Pneumatics functional
102 | 1617 | SmartSEM Control tested
102 | 1617 | Images saved
102 | 1618 | Pneumatics functional
102 | 1618 | SmartSEM Control tested
102 | 1618 | Images saved
102 | 1619 | Pneumatics functional
102 | 1619 | SmartSEM Control tested
102 | 1619 | Images saved
102 | 1620 | Pneumatics functional
102 | 1620 | SmartSEM Control tested
102 | 1620 | Images saved
102 | 1621 | Pneumatics functional
102 | 1621 | SmartSEM Control tested
102 | 1621 | Images saved
102 | 1622 | SmartSEM Control tested
102 | 1622 | Images saved
102 | 1623 | SmartSEM Control tested
102 | 1623 | Images saved
102 | 1625 | Function of CCD cameras/ focus.
102 | 1625 | Confirm that camera toggle function works.
102 | 1626 | Function of Internal Specimen Current Monitor checked.
102 | 1626 | Confirm current is zero when beam blanked (no offset).
397 | 10924 | SmartPI Software Installed and Tested
46 | 1567 | Remove the sample holder from the system.
46 | 1567 | Remove Docking Brackets and fit Transit clamps.
46 | 1567 | Airfeet deflated and free to move.
46 | 1567 | Backing lines re-configured so that PVSV valve is at-tached to flexi hose. Secure valve to frame.
46 | 1567 | If applicable, ensure that the A/L push rod is removed and blank fitted in its place. Ensure A/L is pumped down before powering down system.
46 | 1567 | Select 'Partial vent on standby' in Smartsem.
46 | 1568 | Power system down.
46 | 1568 | Blow water out of the cooling system.
46 | 1568 | Chamber secured to vibration plate
46 | 1568 | Secure Airfeet with foam blocks.
46 | 1568 | Fit foam blocks around column cladding.
46 | 1568 | Secure Turbo pump with foam blocks.
46 | 1568 | Fit Covers to all unused connectors.
46 | 1568 | Mains and Earth cable stowed on Mains unit.
46 | 1568 | Fit shipping Eyes to transit blocks.
46 | 1568 | Confirm all column cladding is secure and screws Loctited in place.
46 | 1568 | Put thread lock on pendulum screws and transit block Screws.
46 | 1568 | All doors closed and secured.
46 | 1568 | All plinth panels removed and stored in transit trolley.
46 | 1568 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 1568 | Print shipping stickers.
46 | 1568 | Fit M8 screw to secure stage door to chamber.
20 | 425 | Record MAC adress
20 | 425 | Network access requested
21 | 385 | Select correct System type (HV / VP)
21 | 385 | If Pressure Stage is fitted Tick Box In Administrator
25 | 276 | Ext V =
25 | 276 | Ext I =
25 | 276 | Measured Speciment Current =
25 | 277 | Ext V =
25 | 277 | Ext I =
25 | 277 | Measured Speciment Current =
38 | 335 | Specimen current @ 20kV =
44 | 350 | Actual
44 | 351 | Actual
46 | 1567 | Remove the sample holder from the system.
46 | 1567 | Remove Docking Brackets and fit Transit clamps.
46 | 1567 | Airfeet deflated and free to move.
46 | 1567 | Backing lines re-configured so that PVSV valve is at-tached to flexi hose. Secure valve to frame.
46 | 1567 | If applicable, ensure that the A/L push rod is removed and blank fitted in its place. Ensure A/L is pumped down before powering down system.
46 | 1567 | Select 'Partial vent on standby' in Smartsem.
38 | 335 | HV Mode: Measured Specimen Current =
38 | 335 | VP Mode: Measured Specimen Current (0 if not applicable) =
38 | 1631 | HV Mode: Measured Specimen Current =
38 | 1631 | HV Mode: Measured Specimen Current =
38 | 1631 | VP Mode: Measured Specimen Current (0 if not applicable) =
20 | 425 | Record MAC adress
20 | 425 | Network access requested
25 | 276 | Ext V =
25 | 276 | Measured Speciment Current =
25 | 277 | Ext V =
25 | 277 | Measured Speciment Current =
38 | 335 | Measured Specimen Current @ 20kV =
80 | 1570 | Reading =
80 | 1586 | Reading =
20 | 252 | PC S/N
20 | 252 | Hard disk Image (Version)
20 | 252 | PC/Model (SAP No.)
20 | 252 | Record Operating System
20 | 252 | Activate Windows
20 | 252 | Record Windows Product Key
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them.
20 | 252 | Check for available SmartSEM hotfixes and install them.
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Approved Datazone Loaded
21 | 255 | Air (Pressure = 6 Bar)
21 | 255 | Water
21 | 255 | Mains
21 | 255 | Safety Earth
21 | 255 | N2
21 | 256 | Save Full System Checker report
22 | 261 | Carry out a 3hrs Plasma clean (Remove all samples from chamber!)
22 | 262 | Carry out overnight ‘Full service Bakeout’ and note actual duration.
22 | 262 | Check thermocouple readings and iron pump plates heat up.
22 | 266 | Use the 35kV External supply for 30mins.
22 | 267 | Use Autokathode tool if available.
31 | 11042 | Confirm that when the beam is blanked the measured I Probe = 0+/- 5pA
22 | 268 | Use Autokathode tool if available
23 | 269 | Run macro 'ESB Runup 3h pumpen & 9h hochfahren.MLF'
24 | 273 | Level Frame
24 | 273 | Check pendulum height (with spacer)
24 | 273 | Set up ‘float’ height
24 | 273 | Check inlet / outlet valves
24 | 273 | Set up docking pins
25 | 275 | Are Gun Head clamping screws secure?
25 | 275 | Are Condensor clamping screws secure?
25 | 275 | Save Emission image @1Kv (cond off)
25 | 275 | Save Emission image @10Kv (cond off)
25 | 279 | Fil I
25 | 279 | Ext V
26 | 283 | Check for ESB charging and Flashing.
27 | 284 | 20nA-Emmision rotates CW with I increase
27 | 285 | No image rotation with WD change
27 | 286 | Angle =
27 | 289 | C2 Calibration Value
27 | 289 | C3 Calibration Value
102 | 11056 | Touch Alarm Functions OK
27 | 404 | 100nA-Emmision rotates ACW with I increase
40 | 1557 | Leack check
40 | 1557 | CCD
40 | 1557 | Focus
29 | 295 | Marco successfully completed.
29 | 295 | Check for Drift, Charging, ESB issues
30 | 296 | Calibration wizard completed.
30 | 297 | Check Mag cal errors are <1% in X & Y axes in all 3 Mag Ranges
30 | 298 | Macro (runs 30 mins.) completed.
30 | 299 | Macro successfully completed.
30 | 299 | Save Image to Image directory.
30 | 300 | X =
30 | 300 | Y =
30 | 302 | Wizard successfully completed.
30 | 303 | Actual
30 | 305 | Holder Height
30 | 305 | Holder Rotation
30 | 305 | Select Customer Sample Holders
30 | 305 | Confirm Holder touch alarm works (Alarm tone and error message)
32 | 311 | Actual =
32 | 314 | Normal mode
32 | 314 | High Current mode
32 | 444 | Actual @ 1kV WD5 =
32 | 444 | Actual @ 20kV WD2 =
22 | 11029 | Clear EPROM on VAC Extension
44 | 454 | Interlock tested and working
33 | 317 | No excessive stigs etc.
34 | 318 | Actual =
38 | 333 | 30kV Max Ext Voltage =
38 | 333 | Complete X-Ray certificate
38 | 335 | Measured Specimen Current @ 20kV =
38 | 337 | Actual (Limit <= 300nm) =
39 | 339 | Actual =
39 | 339 | After conditioning confirm there are no EMServer error messages
39 | 339 | Save Gun Monitor as Probestability.MN2
40 | 340 | 5 – 30Pa Variable Pressure
40 | 340 | VP image taken (30kV, 200kX)
40 | 340 | Ensure Penning fitted is Type ‘PKR’
40 | 340 | No arcing of collector up to 290v @ 30Pa
41 | 343 | Before Cladding the system check that all column connections aresecurely fastened.
41 | 343 | Strip down Aluminium support frame and clean tapped holes. Re-assemble frame and secure screws with Loctite 222.
41 | 343 | Ensure Cladding geometry is correct and all parts aligned.
80 | 1586 | Reading (≤5.00e-09 mbar) =
42 | 344 | Actual X-Axis
42 | 344 | Actual Y-Axis
43 | 345 | For Gemini 300: Aper 1 Normal current; for Gemini 500: Aper 1 High current
43 | 346 | 1kV =
43 | 346 | 15kV =
43 | 347 | 1kV
43 | 347 | 15kV
44 | 348 | Chamber Door Interlock
44 | 348 | Service Switch
44 | 348 | Lens De-activation
44 | 348 | Detector De-activation
44 | 348 | EHT + IPPSU De-activation
44 | 349 | Carry out safety tests using the Secutest S111 pat Tester.
44 | 349 | Fill out completely the Product Safety Tests form in Section 13.
106 | 1637 | X-Ray Certificate
106 | 1637 | Quality Certificate
106 | 1637 | Quality Feedback Form
106 | 1637 | Denka Data Sheet
106 | 1637 | Options Removed Packing list
106 | 1637 | Options Retained Packing list
106 | 1637 | Standard Packing List
106 | 1637 | PC Interface assignment
106 | 1637 | Declaration of conformity
106 | 1637 | Installation report
106 | 1637 | Acceptance test record
106 | 1637 | License Generator Report
106 | 1637 | Rapid Quality Feedback
27 | 428 | Polarity is correct
40 | 447 | Leack check
40 | 447 | X-Ray check
40 | 447 | Analyze Segments
40 | 447 | Gain Ranges
40 | 447 | CZ Detector Cal
40 | 1554 | Leack check
40 | 1554 | Function Check
40 | 448 | Leack check
40 | 448 | X-Ray check
40 | 448 | Analyze Segments
40 | 448 | Gain Ranges
40 | 448 | CZ Detector Cal
40 | 1554 | Leack check
40 | 1554 | X-Ray check
40 | 1554 | Function Check
40 | 1555 | Leack check
40 | 1555 | X-Ray check
40 | 1555 | Pneumatics
40 | 1555 | Alignment
40 | 1555 | Repeatability
40 | 1555 | Bias Voltage
40 | 1555 | Vac Degradation < 3x10^-15 mbar
40 | 1555 | NanoVP Operation
40 | 1555 | XVP Operation
195 | 1657 | Measured probe current
20 | 425 | Record MAC adress
20 | 425 | Network access requested
40 | 447 | Leack check
40 | 447 | Analyze Segments
40 | 447 | Gain Ranges
40 | 447 | CZ Detector Cal
40 | 448 | Leack check
40 | 448 | Analyze Segments
40 | 448 | Gain Ranges
40 | 448 | CZ Detector Cal
40 | 449 | Leack check
40 | 449 | Analyze Segments
40 | 449 | Gain Ranges
40 | 449 | Pneumatics
40 | 449 | Alignment
40 | 449 | Vac Degradation < 3x10^-15 mbar
40 | 449 | CZ Detector Cal
40 | 449 | Z Adjustment
40 | 451 | Leack check
40 | 451 | Analyze Segments
40 | 451 | Gain Ranges
40 | 451 | Pneumatics
40 | 451 | Alignment
40 | 451 | Vac Degradation < 3x10^-15 mbar
40 | 451 | CZ Detector Cal
40 | 451 | Z Adjustment
40 | 452 | Leack check
40 | 452 | Analyze Segments
40 | 452 | Gain Ranges
40 | 452 | Pneumatics
40 | 452 | Alignment
40 | 452 | Stage Safe Position
40 | 452 | Vac Degradation < 3x10^-15 mbar
40 | 452 | CZ Detector Cal
40 | 452 | Store Holder Position 1
40 | 453 | Leack check
40 | 453 | Analyze Segments
40 | 453 | Gain Ranges
40 | 453 | Pneumatics
40 | 453 | Alignment
40 | 453 | Stage Safe Position
40 | 453 | Vac Degradation < 3x10^-15 mbar
40 | 453 | CZ Detector Cal
40 | 453 | Store Holder Position 1
40 | 1550 | Leack check
40 | 1550 | Function check
40 | 1551 | Leack check
40 | 1551 | Grid
40 | 1551 | Check Cut Off
40 | 1552 | Leack check
40 | 1552 | Collector
40 | 1552 | Scintillator
40 | 1553 | Leack check
40 | 1553 | Collector
40 | 1555 | Leack check
40 | 1555 | Pneumatics
40 | 1555 | Alignment
40 | 1555 | Repeatability
40 | 1555 | Bias Voltage
40 | 1555 | Vac Degradation < 3x10^-15 mbar
40 | 1555 | NanoVP Operation
40 | 1555 | XVP Operation
40 | 1558 | Leack check
40 | 1558 | CCD
40 | 1558 | Focus
40 | 1559 | Leack check
40 | 1559 | Macros loaded/tested
40 | 1559 | $exchange
40 | 1559 | Rod Position
40 | 1559 | Front Panel Control
40 | 1560 | Leack check
40 | 1560 | Disable/Enable
40 | 1560 | Save Vac Trace
40 | 1561 | Leack check
40 | 1561 | EHT Cal
40 | 1561 | Spark Alarm
40 | 1561 | Res Prep
40 | 1561 | Resolution
40 | 1562 | Leack check
40 | 1562 | Run Clean Cycle
40 | 1564 | Leack check
40 | 1564 | Align Needle
40 | 1564 | Pneumatics
40 | 1564 | Pressure Test
195 | 1658 | Serial No. (or or N/A)
195 | 1658 | No X-Ray Leackage detected.
28 | 10647 | Confirm SEM Status that C2 & C3 Cal and System Align Offset Angle are sensible
29 | 295 | Marco successfully completed.
29 | 295 | Check for Drift, Charging, ESB issues
38 | 333 | 30kV Max Ext Voltage =
40 | 1553 | Collector
195 | 1658 | Serial No. (or or N/A)
195 | 1658 | No X-Ray Leakage detected.
38 | 335 | Measured Specimen Current @ 20kV =
40 | 1552 | Leak check
40 | 1552 | Collector
40 | 1552 | Scintillator
22 | 11030 | Configue Raith Open X Interface
195 | 10648 | Set EHT to 30KV
195 | 10649 | Complete X-Ray certificate
20 | 252 | PC S/N
20 | 252 | Hard disk Image (Version)
20 | 252 | PC/Model (SAP No.)
20 | 252 | Record Operating System
20 | 252 | Activate Windows
20 | 252 | Record Windows Product Key
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Approved Datazone Loaded
20 | 254 | Language
27 | 429 | Angle =
30 | 298 | Macro (runs 30 mins.) completed.
30 | 299 | Macro successfully completed.
30 | 299 | Save Image to Image directory.
32 | 311 | Actual =
34 | 318 | Actual =
37 | 328 | Tablefiller successfully completed.
37 | 328 | Manual alignment completed.
40 | 447 | Analyze Segments
40 | 447 | Gain Ranges
40 | 447 | CZ Detector Cal
40 | 448 | Analyze Segments
40 | 448 | Gain Ranges
40 | 448 | CZ Detector Cal
40 | 449 | Analyze Segments
40 | 449 | Gain Ranges
40 | 449 | Pneumatics
40 | 449 | Alignment
40 | 449 | Vac Degradation < 3x10^-15 mbar
40 | 449 | CZ Detector Cal
40 | 449 | Z Adjustment
40 | 451 | Analyze Segments
40 | 451 | Gain Ranges
40 | 451 | Pneumatics
40 | 451 | Alignment
40 | 451 | Vac Degradation < 3x10^-15 mbar
40 | 451 | CZ Detector Cal
40 | 451 | Z Adjustment
40 | 452 | Analyze Segments
40 | 452 | Gain Ranges
40 | 452 | Pneumatics
40 | 452 | Alignment
40 | 452 | Stage Safe Position
40 | 452 | Vac Degradation < 3x10^-15 mbar
40 | 452 | CZ Detector Cal
40 | 452 | Store Holder Position 1
40 | 453 | Analyze Segments
40 | 453 | Gain Ranges
40 | 453 | Pneumatics
40 | 453 | Alignment
40 | 453 | Stage Safe Position
40 | 453 | Vac Degradation < 3x10^-15 mbar
40 | 453 | CZ Detector Cal
40 | 453 | Store Holder Position 1
40 | 1550 | Function check
40 | 1551 | Grid
40 | 1551 | Check Cut Off
40 | 1552 | Collector
40 | 1552 | Scintillator
40 | 1554 | Function Check
40 | 1557 | CCD
40 | 1557 | Focus
40 | 1558 | CCD
40 | 1558 | Focus
40 | 1559 | Macros loaded/tested
40 | 1559 | $exchange
40 | 1559 | Rod Position
40 | 1559 | Front Panel Control
40 | 1560 | Disable/Enable
40 | 1560 | Save Vac Trace
40 | 1561 | EHT Cal
40 | 1561 | Spark Alarm
40 | 1561 | Res Prep
40 | 1561 | Resolution
40 | 1562 | Run Clean Cycle
40 | 1564 | Align Needle
40 | 1564 | Pneumatics
40 | 1564 | Pressure Test
44 | 349 | Carry out safety tests using the Secutest S111 pat Tester.
44 | 349 | Fill out completely the Product Safety Tests form in Section 13.
44 | 359 | Remove all Test Macro Shortcuts
44 | 359 | Remove all Test Program Shortcuts
44 | 359 | Set up Desktop icons (both accounts)
44 | 359 | Reset WSUS Settings
46 | 1568 | Power system down.
46 | 1568 | Blow water out of the cooling system.
46 | 1568 | Chamber secured to vibration plate
46 | 1568 | Secure Airfeet with foam blocks.
46 | 1568 | Fit foam blocks around column cladding.
46 | 1568 | Secure Turbo pump with foam blocks.
46 | 1568 | Fit Covers to all unused connectors.
46 | 1568 | Remove Cekon Plug From Mains Cable
46 | 1568 | Fit Live and Neutral Labels to Mains Cable
46 | 1568 | Mains and Earth cable stowed on Mains unit.
46 | 1568 | Fit shipping Eyes to transit blocks.
46 | 1568 | Confirm all column cladding is secure and screws Loctited in place.
46 | 1568 | Put thread lock on pendulum screws and transit block Screws.
46 | 1568 | All doors closed and secured.
46 | 1568 | All plinth panels removed and stored in transit trolley.
46 | 1568 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 1568 | Print shipping stickers.
46 | 1568 | Fit M8 screw to secure stage door to chamber.
81 | 1571 | Measured value =
82 | 1591 | Take and save images at Mag 200, 2kX and 35kX.
82 | 1591 | Mag is within +/- 1% for each.
82 | 1591 | Orthognality is within +/- 1 deg for each.
82 | 1592 | Take and save images at Mag 200, 2kX and 35kX.
82 | 1592 | Mag is within +/- 2% for each.
82 | 1592 | Orthognality is within +/- 1 deg for each.
80 | 1570 | Reading (≤2.00e-06) =
45 | 10779 | Special Fitted
45 | 10779 | If Fitted, Checked by SCS Team Member
40 | 452 | Analyze Segments
40 | 452 | Gain Ranges
40 | 452 | Pneumatics
40 | 452 | Alignment
40 | 452 | Stage Safe Position
40 | 452 | Vac Degradation < 3x10-5 mbar
40 | 452 | CZ Detector Cal
40 | 452 | Store Holder Position 1
40 | 10868 | Connect Silver Box and Cables
40 | 10868 | Check Altas Test USB Stick Has The latest Version Of Software
40 | 10868 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
40 | 10868 | Request Licence File from Fibics
40 | 10868 | Install Atlas Software
40 | 10868 | Install Licence File
40 | 10868 | Calibrate SEM
40 | 10868 | Obtain Mosaic
40 | 10868 | Set ATLAS Rotation Offset
40 | 10868 | Save ATLAS Licence Document
195 | 10648 | Set EHT to 30KV
195 | 10649 | Complete X-Ray certificate
20 | 10930 | Run the ‘Disable FIB–all GIS.reg’ file
25 | 278 | HV = Condenser Off 220pA
25 | 278 | VP = Condenser On 300Pa
25 | 278 | Actual Spec I
27 | 285 | No image rotation with WD change
27 | 10941 | 10KV Optimum Mode 0% DOF
27 | 10941 | 10KV Analytic Mode 100% DOF 100pA
27 | 10941 | 10KV Analytic Mode 100% DOF 5nA
27 | 10941 | 10KV Analytic Mode 0% DOF 100pA
27 | 10941 | 2KV optimum Mode 0% DOF
32 | 445 | Actual @ Optimum Mode =
36 | 321 | Analytic mode, 10kV, WD2.Max dev= Measured value +3pa of slider
36 | 322 | Max Dev
36 | 323 | Max dev =
36 | 325 | Configuration dependent:
36 | 325 | 40 – 48nA
36 | 325 | 100 – 120nA
36 | 325 | 300 – 360nA
36 | 325 | Configuration dependent:
36 | 325 | Actual Value
40 | 453 | Analyze Segments
40 | 453 | Gain Ranges
40 | 453 | Pneumatics
40 | 453 | Alignment
40 | 453 | Stage Safe Position
40 | 453 | Vac Degradation < 3x10-5 mbar
40 | 453 | CZ Detector Cal
40 | 453 | Store Holder Position 1
39 | 339 | Actual =
39 | 339 | After conditioning confirm there are no EMServer error messages
39 | 339 | Save Gun Monitor as Probestability.MN2
26 | 280 | Use Insight Column Tool. Check for cut off and charging particles.
26 | 10965 | ESB Serial Number.
41 | 10964 | Fit Column MU Sheild.
42 | 344 | Check Probe Current
42 | 344 | Actual X-Axis
42 | 344 | Actual Y-Axis
43 | 345 | For Gemini 300: Aper 1 Normal current; for Gemini 500: Aper 1 High current
43 | 346 | 1kV =
43 | 346 | 15kV =
43 | 347 | 1kV
43 | 347 | 15kV
43 | 410 | 1kV ESB
43 | 10971 | Use Probe I Cal
44 | 355 | Ensure the S/n and Part No. match the sales order
40 | 449 | Analyze Segments
40 | 449 | Gain Ranges
40 | 449 | Pneumatics
40 | 449 | Alignment
40 | 449 | Vac Degradation < 3x10-5 mbar
40 | 449 | CZ Detector Cal
40 | 449 | Z Adjustment
40 | 450 | Analyze Segments
40 | 450 | Gain Ranges
40 | 450 | Alignment
40 | 450 | Vac Degradation < 3x10-5 mbar
40 | 450 | CZ Detector Cal
40 | 451 | Analyze Segments
40 | 451 | Gain Ranges
40 | 451 | Pneumatics
40 | 451 | Alignment
40 | 451 | Vac Degradation < 3x10-5 mbar
40 | 451 | CZ Detector Cal
40 | 451 | Z Adjustment
82 | 10982 | Stigmator Value X
82 | 10982 | Stigmator Value Y
82 | 10983 | Position Measurement
82 | 10984 | Focus Deviation
82 | 10985 | Focus Deviation
82 | 10987 | Actual
700 | 10988 | Confirm “EHT ON“ Mode (No Interlock)
700 | 10988 | Runup Gun to correct conditions
42 | 344 | Check Probe Current
42 | 344 | Actual X-Axis
42 | 344 | Actual Y-Axis
42 | 11013 | Check Probe Current
42 | 11013 | Actual X-Axis
42 | 11013 | Actual Y-Axis
40 | 1556 | Functionality
40 | 1556 | Language
40 | 1563 | All Axes Funtion
40 | 1563 | Break Button Function
40 | 1556 | Functionality
40 | 1556 | Language
40 | 10868 | Connect Silver Box and Cables
40 | 10868 | Check Altas Test USB Stick Has The latest Version Of Software
40 | 10868 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
40 | 10868 | Request Licence File from Fibics
40 | 10868 | Install Atlas Software
40 | 10868 | Install Licence File
40 | 10868 | Calibrate SEM
40 | 10868 | Obtain Mosaic
40 | 10868 | Set ATLAS Rotation Offset
40 | 10868 | Save ATLAS Licence Document
41 | 343 | Before Cladding the system check that all column connections aresecurely fastened.
41 | 343 | Strip down Aluminium support frame and clean tapped holes. Re-assemble frame and secure screws with Loctite 222.
41 | 343 | Ensure Cladding geometry is correct and all parts aligned.
20 | 11021 | PC Serial Number:
20 | 11021 | Connect Up Raith Test Tool PC
20 | 11021 | Unistall SmartSEM
20 | 11021 | Install SmartSEM Version
20 | 11021 | Install Latest Service Pack
20 | 11021 | Install Gemini Support Tools-Version
44 | 349 | Carry out safety tests using the Secutest S111 pat Tester.
44 | 349 | Fill out completely the Product Safety Tests form in Section 13.
44 | 355 | Ensure the S/n and Part No. match the sales order
85 | 11019 | Store Images in QA folder
85 | 11020 | XB Tilt Test carried out
40 | 11023 | ZEN Version
40 | 11023 | Correct Version of Windows 10 Installed
40 | 11023 | Install ZEN
40 | 11023 | Install Licence
40 | 11023 | Install Remote API Client
40 | 11023 | Functional Test
20 | 252 | PC S/N
20 | 252 | Hard disk Image (Version)
20 | 252 | PC/Model (SAP No.)
20 | 252 | Record Operating System
20 | 252 | Activate Windows
20 | 252 | Record Windows Product Key
20 | 252 | Disable Renesas Pawer Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
20 | 254 | Language
20 | 425 | Record MAC adress
20 | 425 | Network access requested
20 | 10930 | Run the ‘Disable FIB–all GIS.reg’ file
22 | 266 | Test Power Supply Serial Number
22 | 266 | Use the 35kV External supply for 30mins.
22 | 266 | Save Gun Moitor Log as GunConditioning_TIP SN
195 | 1658 | Monitor Model
195 | 1658 | Serial Number
195 | 1658 | Cal. Expiry Date
195 | 1658 | Serial No. (or or N/A)
195 | 1658 | No X-Ray Leakage detected.
22 | 267 | Use Autokathode tool if available.
22 | 268 | Use Autokathode tool if available
22 | 268 | Save Gun Moitor Log as Gun Run Up_TIP SN
22 | 421 | Select "Tools and User Preferences": Set "Expert Gun Mode = Yes"
20 | 11038 | PC S/N
20 | 11038 | Hard disk Image (Version)
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Record Operating System
20 | 11038 | Activate Windows
20 | 11038 | Record Windows Product Key
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
21 | 255 | Air (Pressure = 6 Bar)
21 | 255 | Water
21 | 255 | Mains
21 | 255 | Safety Earth
21 | 255 | N2
21 | 256 | Save Full System Checker report
21 | 257 | In Factory Tab: Omit Dongle Check
21 | 257 | In Factory Tab: Enter S/N
21 | 257 | In License Tab: Set All Licenses
21 | 258 | Confirm there are no Errors in EMserver (Apart from Dongle check and Stage Init)
21 | 259 | Check camera Focus and Rotation
21 | 259 | Eliminate camera image tilt
21 | 385 | Select correct System type (HV / VP)
21 | 385 | If Pressure Stage is fitted Tick Box In Administrator
21 | 387 | Order revision in use or 'n/a'
21 | 11025 | Fit 2 Off Ferrites to Liner Cable
21 | 11026 | Plasma Clean Chamber for 3 Hours at 70W
21 | 11028 | Switc Panel Fitted
22 | 261 | Carry out a 3hrs Plasma clean (Remove all samples from chamber!)
22 | 262 | Carry out overnight ‘Full service Bakeout’ and note actual duration.
22 | 262 | Check thermocouple readings and iron pump plates heat up.
82 | 1592 | Take and save images at Mag 200, 3kX and 35kX.
82 | 1592 | Mag is within +/- 2% for each.
82 | 1592 | Orthognality is within +/- 1 deg for each.
22 | 11031 | Joystick Angle Set
22 | 11033 | Check Function of the EMO Button
25 | 276 | Ext V =
25 | 276 | Measured Speciment Current =
25 | 11035 | With Cond. Off Max Aper. X & Y < 45
25 | 11035 | Gun Head Clamping Screws Secure
25 | 11035 | Condenser Clamping Screws Secure
25 | 11036 | Large Aperture in Top Left Quadrant
25 | 11036 | Save Emission Image with Condenser On
25 | 11036 | Save Emission Image with Condenser Off
27 | 284 | 20nA-Emmision rotates CW with I increase
27 | 289 | C2 Calibration Value
27 | 289 | C3 Calibration Value
27 | 406 | Alignment done and confirmed.
27 | 429 | Angle =
27 | 11037 | Stage Angle is 133Degs
30 | 298 | Macro (runs 30 mins.) completed.
30 | 299 | Macro successfully completed.
30 | 299 | Save Image to Image directory.
30 | 300 | X =
30 | 300 | Y =
30 | 302 | Wizard successfully completed.
30 | 303 | Actual
30 | 305 | Holder Height
30 | 305 | Holder Rotation
30 | 305 | Select Customer Sample Holders
30 | 305 | Confirm Holder touch alarm works (Alarm tone and error message)
32 | 311 | Actual =
37 | 328 | Tablefiller successfully completed.
37 | 328 | Manual alignment completed.
44 | 359 | Remove all Test Macro Shortcuts
44 | 359 | Remove all Test Program Shortcuts
44 | 359 | Set up Desktop icons (both accounts)
44 | 359 | Reset WSUS Settings
44 | 359 | Remove ASAP Screening Software
45 | 362 | Em Server.mdb
45 | 362 | Em Server.txt Log
45 | 362 | Optiprobe.txt Log
45 | 362 | Smart Backup.log
45 | 362 | L500.ini
45 | 362 | License file.xml
45 | 362 | Registry file.reg
45 | 362 | CalibSc.dat
45 | 11024 | Em Server.mdb
45 | 11024 | Em Server.txt Log
45 | 11024 | Smart Backup.log
45 | 11024 | L500.ini
45 | 11024 | License file.xml
45 | 11024 | Registry file.reg
45 | 11024 | CalibDC.dat
45 | 11024 | Export Tablefiller
81 | 1588 | Measured value =
81 | 1589 | Measured value =
32 | 314 | Normal mode
32 | 314 | High Current mode
32 | 444 | Actual @ 1kV WD5 =
32 | 444 | Actual @ 20kV WD2 =
34 | 318 | Actual =
38 | 333 | 30kV Max Ext Voltage =
38 | 335 | Measured Specimen Current @ 20kV =
40 | 1552 | Collector
40 | 1552 | Scintillator
42 | 11046 | 30um, All EHT= ≤20%
42 | 11046 | Off Axis Aper-All EHT= ≤30%
42 | 11046 | 30um, All EHT= ≤50%
42 | 11047 | Actual X-Axis
42 | 11047 | Actual Y-Axis
43 | 11048 | 200V-Inlens, 1 to 2mm WD,!00Kx, Pixel Avg, SS10.
43 | 11048 | 1KV-Inlens, 2mm WD, 200Kx Line Avg 60, SS4.
43 | 11048 | 15KV-Inlens, 2mm WD, 400Kx, Line Avg 60, SS4.
43 | 11048 | 30KV-SE2, 5mm WD, 400Kx, Pixel Avg, SS10.
43 | 11049 | No EHT Trips or Errors During Test
43 | 11049 | Save As- Dynamic EHT.MN2
44 | 353 | EMserver.mdb
44 | 353 | Customer Images
44 | 353 | License File
44 | 355 | Ensure the S/n and Part No. match the sales order
21 | 258 | Confirm there are no Errors in EMserver (Apart from Dongle check and Stage Init)
44 | 359 | Remove all Test Macro Shortcuts
44 | 359 | Remove all Test Program Shortcuts
44 | 359 | Set up Desktop icons (both accounts)
44 | 359 | Reset WSUS Settings
44 | 359 | Remove ASAP Screening Software
45 | 362 | Em Server.accdb
45 | 362 | Em Server.txt Log
45 | 362 | Optiprobe.txt Log
45 | 362 | Smart Backup.log
45 | 362 | L500.ini
45 | 362 | License file.xml
45 | 362 | Registry file.reg
45 | 362 | CalibSc.dat
45 | 10779 | Special Fitted
45 | 10779 | If Fitted, Checked by SCS Team Member Name
45 | 11024 | Em Server.accdb
45 | 11024 | Em Server.txt Log
45 | 11024 | Smart Backup.log
45 | 11024 | L500.ini
45 | 11024 | License file.xml
45 | 11024 | Registry file.reg
45 | 11024 | CalibDC.dat
45 | 11024 | Export Tablefiller
45 | 11058 | Em Server.accdb
45 | 11058 | Em Server.txt Log
45 | 11058 | Optiprobe.txt Log
45 | 11058 | Smart Backup.log
45 | 11058 | Image Folder
45 | 11058 | L500.ini
45 | 11058 | Dynamic EHT.MN2
45 | 11058 | Probe Stability.MN2
45 | 11058 | License file.xml
45 | 11058 | Registry file.reg
45 | 11058 | System Checker Full Report
45 | 11058 | Export Alignment Data(.SEMcdb)
81 | 1571 | Measured value =
81 | 1587 | Measured value =
82 | 11052 | Image with Condenser OFF
82 | 11052 | Image with Condenser ON
102 | 11054 | Function check
102 | 11055 | Mix SE2 Signal with Inlens Signal
700 | 10988 | Confirm “EHT ON“ Mode (No Interlock)
700 | 10988 | Runup Gun to correct conditions
704 | 11040 | Beamshift Cal. Done On All Entries In Table
704 | 11040 | Shift Between Apertures is <10um
704 | 11041 | Save An Entry For Each In Table Editor (Six entries)
704 | 11043 | 10KV HC - Typical=434pA
704 | 11043 | 10KV 7um HC - Typical=26pA
704 | 11043 | 10KV 10um HC - Typical=55pA
704 | 11043 | 10KV 15um HC - Typical=118pA
704 | 11043 | 10KV 20um HC - Typical=200pA
704 | 11043 | 10KV 60um HC - Typical=1840pA
704 | 11043 | 10KV 1200um HC - Typical=7195pA
704 | 11043 | 1KV - Typical=222pA
704 | 11043 | 12KV - Typical=300pA
704 | 11043 | 20KV - Typical=418pA
704 | 11043 | 20KV HC - Typical=680pA
704 | 11043 | 20.01KV - Typical=300pA
704 | 11043 | 20.01KV HC - Typical=483pA
704 | 11043 | 30KV - Typical=447pA
704 | 11045 | Touch Alarm Function Ok
705 | 11057 | UPS Open X Interface (X121) is Functioning
705 | 11057 | Function Checked Twice
705 | 11057 | Operation of UPS switch , V3 shuts, System Vents and System Shuts Down to Red Button
38 | 334 | Fil I =
38 | 334 | Ext V =
195 | 1658 | Monitor Model
195 | 1658 | Serial Number
195 | 1658 | Cal. Expiry Date
195 | 1658 | Serial No. (or or N/A)
195 | 1658 | No X-Ray Leakage detected.
195 | 1658 | Serial No. (or N/A)
195 | 1658 | No X-Ray Leakage detected.
195 | 1659 | Accelerating Voltage=30KV
195 | 1659 | Specimen Current
195 | 1659 | Monitor Model
195 | 1659 | Serial Number
195 | 1659 | Cal. Expiry Date
195 | 1659 | No X-Ray Leaks Detected
43 | 346 | 1kV =
43 | 346 | 15kV =
21 | 259 | Check camera Focus and Rotation
21 | 259 | Eliminate camera image tilt
21 | 259 | Ty=wrap and Label Cables
21 | 385 | Select correct System type (HV / VP)
21 | 385 | If Pressure Stage is fitted Tick Box In Administrator
45 | 11059 | Copy- 1KV.Tif (Tiff customer image)
45 | 11059 | Copy- 15KV.Tif (Tiff customer image)
45 | 11059 | Copy- 30KV.Tif (Tiff customer image)
45 | 11059 | Copy- 200V.Tif (Tiff customer image)
45 | 11059 | Copy- 9876543.xml (XML Smartsem Licence)
45 | 11059 | Copy- Merlin Compact XX-XX S/O (Licence Rep File)
45 | 11059 | Copy- EmServer.accdb (Database)
45 | 11059 | Copy- EmServer.txt (Server Log File)
45 | 11059 | Copy- L500.ini (SmartSEM Config)
45 | 11059 | Copy- xx-xx.reg (Export of Registry File)
45 | 11059 | Copy- Mybackup.smartbackup (SmartSEM Backup)
45 | 11059 | Copy- Logbook (scanned)-PDF
45 | 11059 | Copy- DynamicEHT (mn2 logfile)
45 | 11059 | Copy- Probe Conditioning Stability (mn2 logfile)
45 | 11059 | Copy- 1KV Res.jpg (PATE Result)
25 | 276 | Ext V =
25 | 276 | Measured Speciment Current =
25 | 276 | Ext V =
25 | 276 | Measured Speciment Current =
27 | 284 | 20nA-Emmision rotates CW with I increase
45 | 11059 | Copy- 1KV.Tif (Tiff customer image)
45 | 11059 | Copy- 15KV.Tif (Tiff customer image)
45 | 11059 | Copy- 30KV.Tif (Tiff customer image)
45 | 11059 | Copy- 200V.Tif (Tiff customer image)
45 | 11059 | Copy- 9876543.xml (XML Smartsem Licence)
45 | 11059 | Copy- Merlin Compact XX-XX S/O (Licence Rep File)
45 | 11059 | Copy- EmServer.accdb (Database)
45 | 11059 | Copy- EmServer.txt (Server Log File)
45 | 11059 | Copy- L500.ini (SmartSEM Config)
45 | 11059 | Copy- xx-xx.reg (Export of Registry File)
45 | 11059 | Copy- Mybackup.smartbackup (SmartSEM Backup)
45 | 11059 | Copy- Logbook (scanned)-PDF
45 | 11059 | Copy- DynamicEHT (mn2 logfile)
45 | 11059 | Copy- Probe Conditioning Stability (mn2 logfile)
45 | 11059 | Copy- 1KV Res.jpg (PATE Result)
45 | 11059 | Copy- 1KV Res.Tif (Tiff Res Image)
45 | 11059 | Copy- 15KV.jpg (PATE Results)
45 | 11059 | Copy- 15KV.Tif (Tiff Res Image)
45 | 11059 | Emission-Condenser OFF (Tiff Image)
45 | 11059 | Emission-Condenser ON (Tiff Image)
45 | 11059 | Merlin Compact User Manual (English) (PDF Manual)
45 | 11059 | Merlin Compact User Manual (German) (PDF Manual)
45 | 11059 | Merlin Compact Assembly (PDF Manual)
20 | 252 | PC S/N
20 | 252 | Hard disk Image (Version)
20 | 252 | PC/Model (SAP No.)
20 | 252 | Record Operating System
20 | 252 | Activate Windows
20 | 252 | Record Windows Product Key
20 | 252 | Disable Renesas Pawer Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
20 | 254 | Language
20 | 425 | Record MAC adress
20 | 425 | Network access requested
21 | 257 | In Factory Tab: Omit Dongle Check
21 | 257 | In Factory Tab: Enter S/N
21 | 257 | In License Tab: Set All Licenses
21 | 387 | Order revision in use or 'n/a'
22 | 261 | Carry out a 3hrs Plasma clean (Remove all samples from chamber!)
22 | 262 | Carry out overnight ‘Full service Bakeout’ and note actual duration.
22 | 262 | Check thermocouple readings and iron pump plates heat up.
22 | 421 | Select "Tools and User Preferences": Set "Expert Gun Mode = Yes"
37 | 331 | Stig: 30μm, All EHT’s ≤ 20%
37 | 331 | Stig: Off Axis Aper, All EHT’s ≤ 30%
37 | 331 | Aper/Gun: 30μm, All EHT’s ≤ 50%
27 | 429 | Angle =
30 | 298 | Macro (runs 30 mins.) completed.
30 | 299 | Macro successfully completed.
30 | 299 | Save Image to Image directory.
30 | 300 | X =
30 | 300 | Y =
30 | 302 | Wizard successfully completed.
30 | 303 | Actual
30 | 305 | Holder Height
30 | 305 | Holder Rotation
30 | 305 | Select Customer Sample Holders
30 | 305 | Confirm Holder touch alarm works (Alarm tone and error message)
32 | 311 | Actual =
34 | 318 | Actual =
41 | 343 | Before Cladding the system check that all column connections aresecurely fastened.
41 | 343 | Strip down Aluminium support frame and clean tapped holes. Re-assemble frame and secure screws with Loctite 222.
41 | 343 | Ensure Cladding geometry is correct and all parts aligned.
44 | 353 | EMserver.mdb
44 | 353 | Customer Images
44 | 353 | License File
44 | 355 | Ensure the S/n and Part No. match the sales order
44 | 359 | Remove all Test Macro Shortcuts
44 | 359 | Remove all Test Program Shortcuts
44 | 359 | Set up Desktop icons (both accounts)
44 | 359 | Reset WSUS Settings
44 | 359 | Remove ASAP Screening Software
45 | 363 | Em Server.mdb
45 | 363 | License File
45 | 363 | Saved Registry
45 | 363 | L500.ini
81 | 1571 | Measured value =
81 | 1587 | Measured value =
21 | 385 | Select correct System type (HV / VP)
21 | 385 | If Pressure Stage is fitted Tick Box In Administrator
21 | 11026 | Plasma Clean Chamber
21 | 11028 | Curved Switch Panel Fitted
22 | 268 | Carry Out Gun Run UP
22 | 268 | Save Gun Monitor Log as Gun Run Up_TIP SN
27 | 11037 | Stage Angle is close to 133Degs
31 | 309 | Confirm that when the beam is blanked the measured I Probe = 0+/- 5pA
31 | 309 | Difference between Eht on and off < 1pA.
30 | 297 | Check Mag cal errors are <1% in X & Y axes in all 3 Mag Ranges. Orthog. ± 1deg
30 | 297 | Save 4 Images ( 3 Mag Ranges + Ortho.)
31 | 11042 | Confirm that when the beam is blanked the measured I Probe = 0+/- 5pA
32 | 312 | Actual @ 1kV WD5mm =
32 | 312 | Actual @ 20kV WD2mm =
384 | 1143 | Stigmators Balanced On All Apertures (Normal Current)
384 | 1143 | Stigmators Balanced On Aperture 5 (High Current ) - X Upper Value
384 | 1143 | Stigmators Balanced On Aperture 5 (High Current ) - X Lower Value
384 | 1143 | Stigmators Balanced On Aperture 5 (High Current ) - Y Upper Value
384 | 1143 | Stigmators Balanced On Aperture 5 (High Current ) - Y Lower Value
384 | 1143 | Stigmators Balanced On Aperture 6 (High Current ) - X Upper Value
384 | 1143 | Stigmators Balanced On Aperture 6 (High Current ) - X Lower Value
384 | 1143 | Stigmators Balanced On Aperture 6 (High Current ) - Y Upper Value
384 | 1143 | Stigmators Balanced On Aperture 6 (High Current ) - Y Lower Value
384 | 1143 | Stigmators Balanced On Aperture 7 (High Current ) - X Upper Value
384 | 1143 | Stigmators Balanced On Aperture 7 (High Current ) - X Lower Value
384 | 1143 | Stigmators Balanced On Aperture 7 (High Current ) - Y Upper Value
384 | 1143 | Stigmators Balanced On Aperture 7 (High Current ) - Y Lower Value
384 | 1143 | Stigmators Balanced On The Remaining Apertures (High Current )
33 | 11062 | 1KV Res-Spec. 1.6nm. WD 1.5-2mm, 300Kx, Line Avg 60, SS4
33 | 11062 | 15KV Res-Spec. 0.8nm. WD 0.5-2mm, 600Kx, Line Avg 60, SS4
44 | 353 | CalibSC.dat
44 | 353 | Customer Images
44 | 353 | License File
44 | 11064 | CalibDC.dat
44 | 11064 | Customer Images
44 | 11064 | License File
45 | 411 | Images and files from the user directory deleted.
45 | 411 | Download Folder Imagesd Delected
45 | 411 | Recycle bin cleared.
46 | 369 | Create Directory C:\Support\Manuals
46 | 369 | Create Shortcut to above Directory on Desktops, called ‘Carl Zeiss User Documentation’.
46 | 419 | C:\Support\Manuals
46 | 419 | Customer USB documentation folder
46 | 420 | Copy SmartSEM instruction manual from Customer USB to C:\Support\Manuals
46 | 11060 | X-Ray Certificate
46 | 11060 | Quality Certificate
46 | 11060 | Quality Feedback Form
46 | 11060 | DENKA Data Sheet
46 | 11060 | Options Retained packing List
46 | 11060 | Standard Packing List
46 | 11060 | Declaration of Conformity
46 | 11060 | Raith Installation Report
46 | 11060 | Software Licence Agreement
704 | 11061 | Confirm that when the beam is blanked the measured I Probe = 0+/- 5pA
21 | 385 | Select correct System type (HV / VP)
21 | 385 | If Pressure Stage is fitted Tick Box In Administrator
106 | 1637 | X-Ray Certificate
106 | 1637 | Quality Certificate
106 | 1637 | Quality Feedback Form
106 | 1637 | Denka Data Sheet
106 | 1637 | Options Removed Packing list
106 | 1637 | Options Retained Packing list
106 | 1637 | Standard Packing List
106 | 1637 | PC Interface assignment
106 | 1637 | Declaration of conformity
106 | 1637 | Installation report
106 | 1637 | Acceptance test record
106 | 1637 | License Generator Report
106 | 1637 | Rapid Quality Feedback
106 | 11065 | X-Ray Certificate
106 | 11065 | Quality Certificate
106 | 11065 | Quality Feedback Form
106 | 11065 | DENKA Data Sheet
106 | 11065 | Options Retained packing List
106 | 11065 | Standard Packing List
106 | 11065 | Declaration of Conformity
106 | 11065 | Raith Installation Report
106 | 11065 | Software Licence Agreement
32 | 311 | Actual =
25 | 11067 | Confirm that when the beam is blanked the measured I Probe = 0+/- 5pA
36 | 321 | Analytic mode, 10kV, WD2.Max dev= Measured value +3pa of slider
36 | 322 | Max Dev
45 | 11068 | Remove all Test Macro Shortcuts
45 | 11068 | Remove all Test Program Shortcuts
45 | 11068 | Set up Desktop icons (both accounts)
45 | 11068 | Reset WSUS Settings
45 | 11068 | Remove ASAP Screening Software
45 | 11068 | Clear Internet History
45 | 11068 | Delete file in download folder
30 | 456 | Eucentricity set up
30 | 456 | Focus Deviation
30 | 11069 | Dynamic focus fuction check - Whole image in focus
30 | 11069 | Tilt Compensation
30 | 11069 | Orthogonality
40 | 447 | Check each individual segment in normal and inverse modes
40 | 447 | Analyze Segments
40 | 447 | Gain Ranges
40 | 447 | CZ Detector Cal
40 | 448 | Check each Channel Functions
40 | 448 | Check each individual segment in normal and inverse modes
40 | 448 | Analyze Segments
40 | 448 | Gain Ranges
40 | 448 | CZ Detector Cal
40 | 449 | Check each individual segment in normal and inverse modes
40 | 449 | Analyze Segments
40 | 449 | Gain Ranges
40 | 449 | Pneumatics
40 | 449 | Alignment
40 | 449 | Vac Degradation < 3x10-5 mbar
40 | 449 | CZ Detector Cal
40 | 449 | Z Adjustment
40 | 450 | Check each individual segment in normal and inverse modes
40 | 450 | Analyze Segments
40 | 450 | Gain Ranges
40 | 450 | Alignment
40 | 450 | Vac Degradation < 3x10-5 mbar
40 | 450 | CZ Detector Cal
102 | 11078 | Stage Moves To Correct Position When Arm Is Opened
102 | 11078 | Image Is Captured
102 | 11078 | Functional Check
676 | 11406 | Select Number of Ways
40 | 451 | Check each Channel Functions
40 | 451 | Check each individual segment in normal and inverse modes
40 | 451 | Analyze Segments
40 | 451 | Gain Ranges
40 | 451 | Pneumatics
40 | 451 | Alignment
40 | 451 | Vac Degradation < 3x10-5 mbar
40 | 451 | CZ Detector Cal
40 | 451 | Z Adjustment
40 | 452 | Check each individual segment in normal and inverse modes
40 | 452 | Analyze Segments
40 | 452 | Gain Ranges
40 | 452 | Pneumatics
40 | 452 | Alignment
40 | 452 | Stage Safe Position
40 | 452 | Vac Degradation < 3x10-5 mbar
40 | 452 | CZ Detector Cal
40 | 452 | Store Holder Position 1
40 | 453 | Check each Channel Functions
40 | 453 | Check each individual segment in normal and inverse modes
40 | 453 | Analyze Segments
40 | 453 | Gain Ranges
40 | 453 | Pneumatics
40 | 453 | Alignment
40 | 453 | Stage Safe Position
40 | 453 | Vac Degradation < 3x10-5 mbar
40 | 453 | CZ Detector Cal
40 | 453 | Store Holder Position 1
30 | 456 | Eucentricity set up
30 | 456 | Focus Deviation
30 | 11069 | Dynamic focus fuction check - Whole image in focus
30 | 11069 | Tilt Compensation
30 | 11069 | Orthogonality
43 | 346 | 1kV =
43 | 346 | 15kV =
43 | 11048 | 200V-Inlens, 1 to 2mm WD,100Kx, Pixel Avg, SS10.
43 | 11048 | 1KV-Inlens, 2mm WD, 200Kx Line Avg 60, SS4.
43 | 11048 | 15KV-Inlens, 2mm WD, 400Kx, Line Avg 60, SS4.
43 | 11048 | 30KV-SE2, 5mm WD, 400Kx, Pixel Avg, SS10.
27 | 286 | Angle =
27 | 11072 | Low mag maode off Angle =
27 | 11072 | Low mag maode on Angle =
30 | 297 | Check Mag cal errors are <1% in X & Y axes in all 3 Mag Ranges. Orthog. ± 1deg
30 | 297 | Save 4 Images ( 3 Mag Ranges + Ortho.)
102 | 11382 | Image Is Captured
102 | 11382 | Functional Check
30 | 11071 | Check Mag cal errors are <2% in X & Y axes in all 3 Mag Ranges. Orthog. ± 1deg
30 | 11071 | Save 4 Images ( 3 Mag Ranges + Ortho.)
27 | 11072 | Low mag mode off Angle =
27 | 11072 | Low mag mode on Angle =
34 | 318 | Actual =
44 | 353 | CalibSC.dat
44 | 353 | Customer Images
44 | 353 | License File
43 | 346 | 1kV =
43 | 346 | 15kV =
44 | 354 | Check Installed Licences Against Sales Order
44 | 355 | Ensure the S/n and Part No. match the sales order
44 | 356 | Check System And Dongle S/No. Match In Service Panel
44 | 356 | Un-tick Omit Dongle In Factory Panel
44 | 358 | All verification & Safety tests done and within threshold limits defined
44 | 358 | Sales Order is latest version
44 | 358 | All items on Sales order accounted for
44 | 454 | Interlock tested and working
44 | 455 | Interlock tested and working
44 | 11064 | CalibDC.dat
44 | 11064 | Customer Images
44 | 11064 | License File
42 | 11046 | 30um, All EHT= ≤20%
42 | 11046 | Off Axis Aper-All EHT= ≤30%
42 | 11046 | 30um, All EHT= ≤50%
102 | 11077 | ZEN Software Installed
102 | 11079 | Atlas Hardware Installed
102 | 11079 | Atlas Software Installed
102 | 11079 | Atlas Licence Saved To SmartSEM USB Stick
40 | 11080 | Function Test
40 | 11081 | Fit Test tool BB Electronis
40 | 11081 | Function Test
40 | 11081 | Remove Test tool BB Electronics
27 | 284 | 20nA-Emmision rotates CW with I increase
27 | 429 | Angle =
45 | 10779 | Special Fitted
45 | 10779 | If Fitted, Checked by SCS Team Member Name
40 | 1553 | Collector
40 | 11084 | Functional Check
40 | 11084 | Function Image
40 | 11085 | Black Level Calibration
40 | 11085 | Functional Check
40 | 11085 | Function Image
102 | 11086 | Function checked in VP mode, signal intensity changes with VP pressure.
102 | 11086 | Signal intensity changes with collector voltage.
45 | 11068 | Remove all Test Macro Shortcuts
45 | 11068 | Remove all Test Program Shortcuts
45 | 11068 | Set up Desktop icons (both accounts)
45 | 11068 | Reset WSUS Settings
45 | 11068 | Remove ASAP Screening Software
45 | 11068 | Clear Internet History
45 | 11068 | Delete file in download folder
45 | 11098 | Set up Desktop icons (both accounts)
45 | 11098 | Remove ASAP Screening Software
45 | 11098 | Clear Internet History
45 | 11098 | Delete file in download folder
20 | 252 | PC S/N
20 | 252 | Hard disk Image (Version)
20 | 252 | PC/Model (SAP No.)
20 | 252 | Record Operating System
20 | 252 | Activate Windows
20 | 252 | Record Windows Product Key
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
195 | 1658 | Serial No. (or N/A)
195 | 1658 | No X-Ray Leakage detected.
20 | 11038 | PC S/N
20 | 11038 | Hard disk Image (Version)
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Record Operating System
20 | 11038 | Activate Windows
20 | 11038 | Record Windows Product Key
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
27 | 429 | Angle =
34 | 318 | Actual =
44 | 11090 | Fit System Warning Label, Check Language, For Canadian systems fit English and French
45 | 411 | Images and files from the user directory deleted.
45 | 411 | Download Folder Images Deleted
45 | 411 | Empty Recycle bin.
82 | 1591 | Take and save images at Mag 200, 2kX and 35kX.
82 | 1591 | Mag is within +/- 1% for each.
82 | 1591 | Orthognality is within +/- 1 deg for each.
44 | 11090 | Fit System Warning Label, Check Language, For Canadian systems fit English and French
45 | 411 | Images and files from the user directory deleted.
45 | 411 | Download Folder Images Deleted
45 | 411 | Empty Recycle bin.
376 | 3252 | F/W Version
376 | 3252 | Hardware Version
46 | 371 | X-Ray Certificate
46 | 371 | Quality Certificate
46 | 371 | Quality Feedback Form
46 | 371 | DENKA Data Sheet
46 | 371 | Options Removed packing List
46 | 371 | Options Retained packing List
46 | 371 | Standard Packing List
46 | 371 | PC Inferface assignment
46 | 371 | Declaration of conformity
46 | 371 | Installation report
46 | 371 | Acceptance test record
46 | 371 | Licence Generator report
46 | 11099 | X-Ray Certificate
46 | 11099 | Quality Certificate
46 | 11099 | SVSEM FIB EXT Datasheet
46 | 11099 | DENKA Data Sheet
46 | 11099 | Options Removed packing List
46 | 11099 | Options Retained packing List
46 | 11099 | Standard Packing List
46 | 11099 | Licence Generator report
46 | 11099 | Print Plinth Logbook
46 | 11099 | Print Column Logbook
46 | 11099 | Print Final Test Logbook
40 | 1559 | Macros loaded/tested
40 | 1559 | $exchange
40 | 1559 | Rod Position
40 | 1559 | Front Panel Control
37 | 328 | Tablefiller successfully completed.
37 | 328 | Manual alignment completed.
38 | 337 | Actual (Limit <= 300nm) =
42 | 344 | Check Probe Current
42 | 344 | Actual X-Axis
42 | 344 | Actual Y-Axis
37 | 328 | Tablefiller successfully completed.
37 | 329 | Stig: 30μm, All EHT’s is ≤ 20%
37 | 329 | Stig: Off Axis Aper, All EHT’s is ≤ 30%
37 | 329 | Aper/Gun 30μm, All EHT’s is ≤ 60%
37 | 11110 | 99v
37 | 11110 | 22kv
37 | 11110 | 10kv vp
37 | 11110 | 10kv Widefield
37 | 11110 | 30kv
37 | 11110 | 1 & 10kv - 100na system only
37 | 11111 | Operation 5310 submitted
37 | 11111 | Operation 5350 submitted
376 | 3253 | F/W Version
376 | 3253 | Hardware Version
30 | 298 | Macro (runs 30 mins.) completed.
34 | 318 | Actual =
40 | 11085 | Black Level Calibration
40 | 11085 | Functional Check
40 | 11085 | Take Image
37 | 329 | Stig: 30μm, All EHT’s is ≤ 20%
37 | 329 | Stig: Off Axis Aper, All EHT’s is ≤ 30%
37 | 329 | Aper/Gun 30μm, All EHT’s is ≤ 60%
37 | 330 | Stig: 20μm, 10kV is ≤ 35%
37 | 330 | Stig: Off Axis Aper, 10kV is ≤ 45%
37 | 330 | Stig: 20μm 1kV is ≤ 20%
37 | 330 | Stig: Off Axis Aper, 1kV is ≤ 30%
37 | 330 | Aper/Gun: 20μm, All EHT’s is ≤ 60%
37 | 331 | Stig: 30μm, All EHT’s ≤ 20%
37 | 331 | Stig: Off Axis Aper, All EHT’s ≤ 30%
37 | 331 | Aper/Gun: 30μm, All EHT’s ≤ 60%
81 | 11112 | Measured value =
37 | 11114 | Stig: 20μm, 10kV is ≤ 35%
37 | 11114 | Stig: Off Axis Aper, 10kV is ≤ 45%
37 | 11114 | Stig: 20μm 1kV is ≤ 20%
37 | 11114 | Stig: Off Axis Aper, 1kV is ≤ 30%
37 | 11114 | Aper/Gun: 20μm, All EHT’s is ≤ 60%
37 | 11114 | Stig: 30μm, All EHT’s is ≤ 20%
37 | 11114 | Stig: Off Axis Aper, All EHT’s is ≤ 30%
37 | 11114 | Aper/Gun 30μm, All EHT’s is ≤ 60%
24 | 273 | Level Frame
24 | 273 | Check pendulum height (with spacer)
24 | 273 | Set up ‘float’ height
24 | 273 | Check inlet / outlet valves
24 | 273 | Set up docking pins
102 | 11077 | ZEN Software Installed
102 | 11079 | Atlas Hardware Installed
102 | 11079 | Atlas Software Installed
102 | 11079 | Atlas Licence Saved To SmartSEM USB Stick
40 | 10868 | Connect Silver Box and Cables
40 | 10868 | Check Altas Test USB Stick Has The latest Version Of Software
40 | 10868 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
40 | 10868 | Request Licence File from Fibics
40 | 10868 | Install Atlas Software
40 | 10868 | Install Licence File
40 | 10868 | Calibrate SEM
40 | 10868 | Obtain Mosaic
40 | 10868 | Set ATLAS Rotation Offset
40 | 10868 | Save ATLAS Licence Document
40 | 11023 | ZEN Version
40 | 11023 | Correct Version of Windows 10 Installed
40 | 11023 | Install ZEN
40 | 11023 | Install Licence
40 | 11023 | Install Remote API Client
40 | 11023 | Functional Test
37 | 330 | Stig: 20μm, 10kV is ≤ 35%
37 | 330 | Stig: Off Axis Aper, 10kV is ≤ 45%
37 | 330 | Stig: 20μm 1kV is ≤ 20%
37 | 330 | Stig: Off Axis Aper, 1kV is ≤ 30%
37 | 330 | Aper/Gun: 20μm, All EHT’s is ≤ 60%
37 | 11111 | Operation 5310 submitted
37 | 11111 | Operation 5350 submitted
195 | 1659 | Accelerating Voltage=30KV
195 | 1659 | Specimen Current
195 | 1659 | Monitor Model
195 | 1659 | Serial Number
195 | 1659 | Cal. Expiry Date
195 | 1659 | No X-Ray Leaks Detected
195 | 10649 | Complete X-Ray certificate
195 | 11119 | Accelerating Voltage=30KV
195 | 11119 | Monitor Model
195 | 11119 | Serial Number
195 | 11119 | Cal. Expiry Date
195 | 11119 | UHV Chamber No X-Ray Leaks Detected
195 | 11119 | Specimen Chamber No X-Ray Leaks Detected
195 | 11119 | Condenser Housing No X-Ray Leaks Detected
195 | 11119 | VP section No X-Ray Leaks Detected
25 | 276 | Ext V =
25 | 276 | Measured Speciment Current =
25 | 277 | Ext V =
25 | 277 | Measured Speciment Current =
25 | 278 | HV = Condenser Off 220pA
25 | 278 | VP = Condenser On 300Pa
25 | 278 | Actual Spec I
195 | 11120 | Accelerating Voltage=30KV
195 | 11120 | Monitor Model
195 | 11120 | Serial Number
195 | 11120 | Cal. Expiry Date
195 | 11120 | UHV Chamber No X-Ray Leaks Detected
195 | 11120 | Specimen Chamber No X-Ray Leaks Detected
195 | 11120 | Condenser Housing No X-Ray Leaks Detected
195 | 11120 | VP section No X-Ray Leaks Detected
195 | 11121 | Serial No. (or N/A)
195 | 11121 | No X-Ray Leakage detected.
376 | 3254 | F/W Version
376 | 3254 | Hardware Version
38 | 337 | Actual (Limit <= 300nm) =
20 | 254 | Language
40 | 11135 | Unpack Gatan + check against packing list (Photograph each layer before removing)
40 | 11135 | Set up stage disable Macro in SmartSEM + assign icon on toolbar
40 | 11135 | Remove transit clamps from Gatan stage and fit tie rods
40 | 11135 | Adjust stage rails so both Klein and Gatan stages can pump without assistance
40 | 11135 | If Nav Cam fitted please contact Peter Lewis to have in – line connection made up
46 | 11123 | Power system down.
46 | 11123 | Blow water out of the cooling system.
46 | 11123 | Chamber secured to vibration plate
46 | 11123 | Secure Airfeet with foam blocks.
46 | 11123 | Fit foam blocks around column cladding.
46 | 11123 | Secure Turbo pump with foam blocks.
46 | 11123 | Fit Covers to all unused connectors.
46 | 11123 | Mains and Earth cable stowed on Mains unit.
46 | 11123 | Fit shipping Eyes to transit blocks.
46 | 11123 | Confirm all column cladding is secure and screws Loctited in place.
46 | 11123 | Put thread lock on pendulum screws and transit block Screws.
46 | 11123 | All doors closed and secured.
46 | 11123 | All plinth panels removed and stored in transit trolley.
46 | 11123 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11123 | Print shipping stickers.
46 | 11123 | Fit M8 screw to secure stage door to chamber.
43 | 347 | 1kV Inlens
43 | 347 | 15kV Inlens
43 | 347 | 1kV ESB
20 | 253 | Enter the FIB Extension number using EEPromUpdater.exe
30 | 456 | Eucentricity set up
30 | 456 | Focus Deviation
30 | 11069 | Dynamic focus fuction check - Whole image in focus
30 | 11069 | Tilt Compensation
30 | 11069 | Orthogonality
36 | 327 | Extractor current should not change by > 4ua between 10kV & 30kV.
46 | 11060 | X-Ray Certificate ----------------------------------------------------------------------------------
46 | 11060 | Quality Certificate ---------------------------------------------------------------------------------
46 | 11060 | Quality Feedback Form
46 | 11060 | DENKA Data Sheet
46 | 11060 | Options Retained packing List
46 | 11060 | Standard Packing List
46 | 11060 | Declaration of Conformity
46 | 11060 | Raith Installation Report
46 | 11060 | Software Licence Agreement
45 | 11058 | Em Server.accdb
45 | 11058 | Em Server.log Text File
45 | 11058 | Optiprobe Calibration.Log Text File
45 | 11058 | Mybackup.Smartbackup
45 | 11058 | Image Folder
45 | 11058 | L500.ini
45 | 11058 | Dynamic EHT.MN2
45 | 11058 | Probe Stability.MN2
45 | 11058 | License file.xml
45 | 11058 | Registry file.reg
45 | 11058 | System Checker Full Report
45 | 11058 | Export Alignment Data From Table Editor(.MERLINCDB)
40 | 11132 | Plasma clean Gatan stage (20W 2Hrs)
40 | 11132 | Fit On Point detector and diode
40 | 11132 | Remove Mc Affee software from Gatan PC
40 | 11132 | Disable Firewalls on Gatan and Zeiss PC
40 | 11132 | Install and configure remote client on Gatan PC
40 | 11132 | Configure ‘R configure‘ server on Zeiss PC
40 | 11132 | Ping IP adresses
40 | 11132 | Configure Gatan settings
31 | 310 | Actual (Limit <= 300nm) =
40 | 11133 | Align BSD and obtain image
40 | 11133 | Check SCM with Gatan Stage
40 | 11133 | Ripple X Spec 3nm (600Kx Mag 15KV WD2)
40 | 11133 | Ripple Y Spec 3nm (600Kx Mag 15KV WD2)
40 | 11133 | Resolution 1KV Spec 2.1nm
40 | 11133 | Resolution 15KV Spec 1.2nm
40 | 11133 | Contamination
40 | 11133 | Pump Down Time
40 | 11133 | Ultimate Vacuum
40 | 11133 | Focal CC fitted to 3View stage
40 | 11133 | Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar
40 | 11133 | Charging paper test carried out and images saved.
40 | 11134 | Assemble Gatan vessel pumping accessories.
40 | 11134 | Re-package system
40 | 11134 | ManEng/ Senior Eng/ Production Manger Sign off
40 | 11135 | Unpack Gatan + check against packing list (Photograph each layer before removing)
40 | 11135 | Set up stage disable Macro in SmartSEM + assign icon on toolbar
40 | 11135 | Remove transit clamps from Gatan stage and fit tie rods
40 | 11135 | Adjust stage rails so both Klein and Gatan stages can pump without assistance
40 | 11135 | If Nav Cam fitted please contact Peter Lewis to have in – line connection made up
195 | 11121 | Serial No. (or N/A)
195 | 11121 | No X-Ray Leakage detected.
40 | 11132 | Plasma clean Gatan stage (20W 2Hrs)
40 | 11132 | Fit On Point detector and diode
40 | 11132 | Remove Mc Affee software from Gatan PC
40 | 11132 | Disable Firewalls on Gatan and Zeiss PC
40 | 11132 | Install and configure remote client on Gatan PC
40 | 11132 | Configure ‘R configure‘ server on Zeiss PC
40 | 11132 | Ping IP adresses
40 | 11132 | Configure Gatan settings
40 | 11133 | Align BSD and obtain image
40 | 11133 | Check SCM with Gatan Stage
40 | 11133 | Ripple X Spec 3nm (600Kx Mag 15KV WD2)
40 | 11133 | Ripple Y Spec 3nm (600Kx Mag 15KV WD2)
40 | 11133 | Resolution 1KV Spec 2.1nm
40 | 11133 | Resolution 15KV Spec 1.2nm
40 | 11133 | Contamination
40 | 11133 | Pump Down Time
40 | 11133 | Ultimate Vacuum
40 | 11133 | Focal CC fitted to 3View stage
40 | 11133 | Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar
40 | 11133 | Charging paper test carried out and images saved.
40 | 11134 | Assemble Gatan vessel pumping accessories.
40 | 11134 | Re-package system
40 | 11134 | ManEng/ Senior Eng/ Production Manger Sign off
27 | 434 | Shift is < 10um over the whole IProbe range
376 | 3255 | Controller Version
376 | 3255 | Axes Version
376 | 3255 | Hardware Version
25 | 278 | Set condenser off. Spec is 220pA
25 | 278 | Actual Spec I
37 | 328 | Tablefiller successfully completed.
699 | 10966 | DC Tablefiller successfully completed.
699 | 11188 | VP Tablefiller Entries successfully completed.
20 | 253 | Enter the FIB Extension number using EEPromUpdater.exe
22 | 261 | Carry out a 3hrs Plasma clean (Remove all samples from chamber!)
34 | 318 | Actual =
21 | 11026 | Carry out a 3hrs Plasma clean (Remove all samples from chamber!)
27 | 436 | Focus deviation within 30um
27 | 435 | Focus deviation within 30um
27 | 436 | Focus deviation within 30um
26 | 280 | Use Insight Column Tool to check for cutoff.
40 | 449 | Check each individual segment in normal and inverse modes
40 | 449 | Analyze Segments
40 | 449 | Gain Ranges
40 | 449 | Pneumatics
40 | 449 | Alignment
40 | 449 | Vac Degradation < 3x10-5 mbar
40 | 449 | CZ Detector Cal
40 | 449 | Z Adjustment
40 | 450 | Check each individual segment in normal and inverse modes
40 | 450 | Analyze Segments
40 | 450 | Gain Ranges
40 | 450 | Alignment
40 | 450 | Vac Degradation < 3x10-5 mbar
40 | 450 | CZ Detector Cal
40 | 451 | Check each Channel Functions
40 | 451 | Check each individual segment in normal and inverse modes
40 | 451 | Analyze Segments
40 | 451 | Gain Ranges
40 | 451 | Pneumatics
40 | 451 | Alignment
40 | 451 | Vac Degradation < 3x10-5 mbar
40 | 451 | CZ Detector Cal
40 | 451 | Z Adjustment
40 | 452 | Check each individual segment in normal and inverse modes
40 | 452 | Analyze Segments
40 | 452 | Gain Ranges
40 | 452 | Pneumatics
40 | 452 | Alignment
40 | 452 | Stage Safe Position
40 | 452 | Vac Degradation < 3x10-5 mbar
40 | 452 | CZ Detector Cal
40 | 452 | Store Holder Position 1
40 | 453 | Check each Channel Functions
40 | 453 | Check each individual segment in normal and inverse modes
40 | 453 | Analyze Segments
40 | 453 | Gain Ranges
40 | 453 | Pneumatics
40 | 453 | Alignment
40 | 453 | Stage Safe Position
40 | 453 | Vac Degradation < 3x10-5 mbar
40 | 453 | CZ Detector Cal
40 | 453 | Store Holder Position 1
36 | 321 | Max dev = Measured value +3pa of slider
36 | 322 | Max dev = Measured value +/-15% of slider
22 | 267 | Use Autokathode tool.
22 | 268 | Carry Out Gun Run UP
22 | 268 | Save Gun Monitor Log as Gun Run Up_TIP SN
44 | 353 | CalibSC.dat
44 | 353 | Customer Images
44 | 353 | License File
44 | 455 | Interlock tested and working
27 | 284 | 20nA-Emmission rotates CW with I increase
27 | 430 | C1 Cal =
27 | 430 | C2 Cal =
27 | 430 | C3 Cal =
27 | 431 | Anode Diameter =
27 | 10941 | 10KV Optimum Mode 0% DOF
27 | 10941 | 10KV Analytic Mode 100% DOF 100pA
27 | 10941 | 10KV Analytic Mode 100% DOF 5nA
27 | 10941 | 10KV Analytic Mode 0% DOF 100pA
27 | 10941 | 2KV optimum Mode 0% DOF
33 | 316 | Actual for 1kV =
33 | 316 | Actual for 15kV =
30 | 296 | Calibration wizard completed.
30 | 297 | Check Mag cal errors are <1% in X & Y axes in all 3 Mag Ranges. Orthog. ± 1deg
30 | 297 | Save 4 Images ( 3 Mag Ranges + Ortho.)
30 | 298 | Macro (runs 30 mins.) completed.
656 | 10663 | Damper Plate is Correct Version
656 | 10663 | Damper Plate is Located Correctlu
30 | 299 | Macro successfully completed.
30 | 299 | Save Image to Image directory.
30 | 300 | X =
30 | 300 | Y =
30 | 302 | Wizard successfully completed.
30 | 303 | Actual
30 | 305 | Holder Height
30 | 305 | Holder Rotation
30 | 305 | Select Customer Sample Holders
30 | 305 | Confirm Holder touch alarm works (Alarm tone and error message)
30 | 11071 | Check Mag cal errors are <2% in X & Y axes in all 3 Mag Ranges. Orthog. ± 1deg
30 | 11071 | Save 4 Images ( 3 Mag Ranges + Ortho.)
30 | 11212 | Holder Height
30 | 11212 | Holder Rotation
30 | 11212 | Select Customer Sample Holders
30 | 11212 | Confirm Holder touch alarm works (Alarm tone and error message)
30 | 11213 | Calibration wizard completed.
31 | 308 | Actual 1kV X Axis =
31 | 308 | Actual 1kV Y Axis =
31 | 308 | Actual 20kV X Axis =
31 | 308 | Actual 20kV Y Axis =
42 | 344 | Actual X-Axis
42 | 344 | Actual Y-Axis
658 | 10682 | Check All Crimps Are Correctly Located And Secure.
658 | 10682 | Test PSU Ensure All LED’s Light Up Correctly
658 | 10682 | To be Checked By Another Engineer - Name:
42 | 11013 | Actual X-Axis
42 | 11013 | Actual Y-Axis
43 | 345 | Refer to Work Instruction for column conditions
43 | 10971 | Use Probe I Cal Graph
26 | 281 | Type in measured value (Limit <= 30um)
26 | 396 | Type in measured value (Limit: <= 50um)
40 | 1553 | Collector
40 | 1561 | EHT Cal
40 | 1561 | Spark Alarm
40 | 1561 | Res Prep
40 | 1561 | Resolution
43 | 346 | 1kV =
43 | 346 | 15kV =
22 | 268 | Carry Out Gun Run UP
22 | 268 | Save Gun Monitor Log as Gun Run Up_TIP SN
25 | 11036 | Large Aperture in Top Left Quadrant
25 | 11036 | Save Emission Image with Condenser On
25 | 11036 | Save Emission Image with Condenser Off
32 | 444 | Actual @ 1kV WD5 =
32 | 444 | Actual @ 20kV WD2 =
42 | 11046 | Stigs: 30um Aperture, All EHT ≤20%
42 | 11046 | Stigs: Off Axis Apertures, All EHT ≤30%
42 | 11046 | Aper & Gun: 30um Aperture, All EHT ≤50%
43 | 347 | 1kV Inlens
43 | 347 | 15kV Inlens
43 | 347 | 1kV ESB
704 | 11040 | Beamshift Cal. Done On All Entries In Table
25 | 426 | Water pipes / connectors clear.
25 | 427 | Mechanically align the column using the Merlin Alignment Wizard
26 | 281 | Type in measured value (Limit <= 30um)
26 | 282 | Type in measured value (limit: <= 15um)
26 | 283 | Check for ESB charging and Flashing.
26 | 396 | Type in measured value (Limit: <= 50um)
33 | 11062 | 1KV Res-Spec. 1.6nm. WD 1.5-2mm, 300Kx, Line Avg 60, SS4
33 | 11062 | 15KV Res-Spec. 0.8nm. WD 0.5-2mm, 600Kx, Line Avg 60, SS4
32 | 314 | Normal mode
32 | 314 | High Current mode
32 | 445 | Actual @ Optimum Mode =
37 | 11110 | 99v
37 | 11110 | 22kv
37 | 11110 | 10kv vp
37 | 11110 | 10kv Widefield
37 | 11110 | 30kv
33 | 316 | Actual for 1kV =
33 | 316 | Actual for 15kV =
33 | 317 | Image quality good and Stigs in spec.
34 | 318 | Actual =
36 | 321 | Max dev = Measured value +3pa of slider
36 | 322 | Max dev = Measured value +/-15% of slider (Use Gun Aligns to maximise Probe I)
36 | 323 | Max dev= Measured value +/-15% of slider.
36 | 324 | Max Dev= Measured value +/-15% of slider.
36 | 325 | Configuration dependent:
36 | 325 | Actual Value
36 | 326 | Actual < 7pA +/- 2pA
37 | 328 | Tablefiller successfully completed.
37 | 330 | Stig: 20μm Aper, 10kV is ≤ 35%
37 | 330 | Stig: Off Axis Aper, 10kV is ≤ 45%
37 | 330 | Stig: 20μm Aper, 1kV is ≤ 20%
37 | 330 | Stig: Off Axis Aper, 1kV is ≤ 30%
37 | 330 | Aper/Gun: 20μm, All EHT’s is ≤ 60%
40 | 11084 | Functional Check
40 | 11084 | Take Image
38 | 333 | 30kV Max Ext Voltage =
38 | 335 | Measured Specimen Current @ 20kV =
46 | 1568 | Power system down.
46 | 1568 | Blow water out of the cooling system.
46 | 1568 | Chamber secured to vibration plate
46 | 1568 | Secure Airfeet with foam blocks.
46 | 1568 | Fit foam blocks around column cladding.
46 | 1568 | Secure Turbo pump with foam blocks.
46 | 1568 | Fit Covers to all unused connectors.
46 | 1568 | Remove Cekon Plug From Mains Cable
46 | 1568 | Fit Live and Neutral Labels to Mains Cable
46 | 1568 | Mains and Earth cable stowed on Mains unit.
46 | 1568 | Fit shipping Eyes to transit blocks.
46 | 1568 | Confirm all column cladding is secure and screws Loctited in place.
46 | 1568 | Put thread lock on pendulum screws and transit block Screws.
46 | 1568 | All doors closed and secured.
46 | 1568 | All plinth panels removed and stored in transit trolley.
46 | 1568 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 1568 | Print shipping stickers.
46 | 1568 | Fit M8 screw to secure stage door to chamber.
195 | 10648 | Set EHT to 30KV
42 | 344 | Actual X-Axis
42 | 344 | Actual Y-Axis
27 | 406 | Alignment done and confirmed.
30 | 11069 | Dynamic focus fuction check - Whole image in focus
30 | 11069 | Y - Axis measurement (400μm +2%)
30 | 11069 | Orthogonality
33 | 316 | Actual for 1kV =
33 | 316 | Actual for 15kV =
39 | 339 | Actual =
39 | 339 | After conditioning confirm there are no EMServer error messages
39 | 339 | Save Gun Monitor as Probestability.MN2
40 | 340 | 5 – 30Pa Variable Pressure
40 | 340 | VP image taken (30kV, 200kX)
40 | 340 | Ensure Penning fitted is Type ‘PKR’
40 | 340 | No arcing of collector up to 290v @ 30Pa
40 | 447 | Configure ASB in Administrator
40 | 447 | Input Offset Cal
40 | 447 | Black Level Cal
40 | 447 | Diode Segment Test
40 | 447 | Diode Mirror Test
40 | 447 | All Gain Ranges working
40 | 447 | Check all channels in Invert mode
40 | 448 | Configure ASB in Administrator
40 | 448 | Input Offset Cal
40 | 448 | Black Level Cal
40 | 448 | Diode Segment Test
40 | 448 | Diode Mirror Test
40 | 448 | All Gain Ranges working
40 | 448 | Check all channels in Invert mode
40 | 449 | Configure BSD in Administrator
40 | 449 | Input Offset Cal
40 | 449 | Black Level Cal
40 | 449 | Mechanical Alignment
40 | 449 | Diode Segment Test
40 | 449 | Diode Mirror Test
40 | 449 | Vac Degradation < 3x10-5 mbar
40 | 449 | All Gain Ranges working
40 | 449 | Check all channels in Invert mode
40 | 450 | Configure ASB in Administrator
40 | 450 | Input Offset Cal
40 | 450 | Black Level Cal
40 | 450 | Diode Segment Test
40 | 450 | Diode Mirror Test
40 | 450 | All Gain Ranges working
40 | 450 | Check all channels in Invert mode
40 | 451 | Configure BSD in Administrator
40 | 451 | Input Offset Cal
40 | 451 | Black Level Cal
40 | 451 | Mechanical Alignment
40 | 451 | Diode Segment Test
40 | 451 | Diode Mirror Test
40 | 451 | Vac Degradation < 3x10-5 mbar
40 | 451 | All Gain Ranges working
40 | 451 | Check all channels in Invert mode
40 | 1551 | Grid working
40 | 1551 | Check Cut Off
40 | 1552 | Collector Working
40 | 1552 | Scintillator Working
40 | 1553 | Collector Working
40 | 1555 | Leak check
40 | 1555 | Pneumatics OK
40 | 1555 | Diode Alignment OK
40 | 1555 | Good Repeatability
40 | 1555 | Bias Voltage Working
40 | 1555 | Vac Degradation < 3x10^-15 mbar
40 | 1555 | NanoVP Operation Working over full range
40 | 1555 | XVP Operation working over full range
40 | 11080 | Functional Test
40 | 11081 | Fit Test tool BB Electronics
40 | 11081 | Function Test
40 | 11081 | Remove Test tool BB Electronics
40 | 11132 | Plasma clean Gatan stage (20W 2Hrs)
40 | 11132 | Fit On Point detector and diode
40 | 11132 | Remove Mc Affee software from Gatan PC
40 | 11132 | Disable Firewalls on Gatan and Zeiss PC
40 | 11132 | Install and configure remote client on Gatan PC
40 | 11132 | Configure ‘R configure‘ server on Zeiss PC
40 | 11132 | Ping IP adresses
40 | 11132 | Configure Gatan settings
40 | 11135 | Unpack Gatan + check against packing list (Photograph each layer before removing)
40 | 11135 | Set up stage disable Macro in SmartSEM + assign icon on toolbar
40 | 11135 | Remove transit clamps from Gatan stage and fit tie rods
40 | 11135 | Adjust stage rails so both Klein and Gatan stages can pump without assistance
40 | 11135 | If Nav Cam fitted please contact Peter Lewis to have in – line connection made up
195 | 1659 | Accelerating Voltage=30KV
195 | 1659 | Specimen Current
195 | 1659 | Monitor Model
195 | 1659 | Serial Number
195 | 1659 | Cal. Expiry Date
195 | 1659 | No X-Ray Leaks Detected
195 | 10649 | Complete X-Ray certificate
195 | 11120 | Accelerating Voltage=30KV
195 | 11120 | Monitor Model
195 | 11120 | Serial Number
195 | 11120 | Cal. Expiry Date
195 | 11120 | UHV Chamber No X-Ray Leaks Detected
195 | 11120 | Specimen Chamber No X-Ray Leaks Detected
195 | 11120 | Condenser Housing No X-Ray Leaks Detected
195 | 11120 | VP section No X-Ray Leaks Detected
35 | 319 | Aperture Finder successfully run
35 | 320 | All 100na Apertures Aligned
37 | 11110 | 99v
37 | 11110 | 22kv
37 | 11110 | 30kv
37 | 11233 | 10kv vp
37 | 11234 | 10kv Widefield
390 | 3319 | Time (<5 min)
31 | 310 | Blanking Drift Distance Median (Limit <= 300nm) =
31 | 310 | Blanking Drift Time Median (Limit <= 3.5 Sec) =
31 | 310 | Corner Test Max Drift (Limit <= 300nm) =
40 | 452 | Check each individual segment in normal and inverse modes
40 | 452 | Analyze Segments
40 | 452 | Gain Ranges
40 | 452 | Pneumatics
40 | 452 | Alignment
40 | 452 | Stage Safe Position
40 | 452 | Vac Degradation < 3x10-5 mbar
40 | 452 | CZ Detector Cal
40 | 452 | Store Holder Position 1
40 | 453 | Check each individual segment in normal and inverse modes
40 | 453 | Analyze Segments
40 | 453 | Gain Ranges
40 | 453 | Pneumatics
40 | 453 | Alignment
40 | 453 | Stage Safe Position
40 | 453 | Vac Degradation < 3x10-5 mbar
40 | 453 | CZ Detector Cal
40 | 453 | Store Holder Position 1
40 | 1559 | Macros loaded/tested
40 | 1559 | $exchange set up
40 | 1559 | Rod Position Correctly Adjusted
40 | 1559 | Front Panel Control all Functional
40 | 1561 | EHT Cal
40 | 1561 | Spark Alarm
40 | 1561 | Resolution Preperation
40 | 1561 | Resolution
40 | 1564 | Align Needle
40 | 1564 | Pneumatics ok
40 | 1564 | Pressure Test (CC Valve fully open)
40 | 1564 | Take image of paper
44 | 454 | Interlock tested and working
41 | 10964 | MU Metal Shield fitted.
42 | 11047 | Actual X-Axis
42 | 11047 | Actual Y-Axis
43 | 345 | Refer to Work Instruction for column conditions
43 | 347 | 1kV Inlens
43 | 347 | 15kV Inlens
43 | 347 | 1kV ESB
43 | 10971 | Use Probe I Cal Graph
43 | 11049 | No EHT Trips or Errors During Test
43 | 11049 | Save as Dynamic EHT.MN2
44 | 348 | Chamber Door Interlock
44 | 348 | Service Switch
44 | 348 | Lens De-activation
44 | 348 | Detector De-activation
44 | 348 | EHT + IPPSU De-activation
44 | 349 | Carry out safety tests using the Secutest S111 pat Tester.
44 | 349 | Fill out completely the Product Safety Tests form in Section 13.
44 | 351 | Actual
44 | 353 | CalibSC.dat
44 | 353 | Customer Images
44 | 353 | License File
44 | 355 | Ensure the S/N and Part No. match the sales order
385 | 3286 | Fil I
385 | 3286 | Ext V
44 | 455 | Interlock tested and working
44 | 11064 | CalibDC.dat
44 | 11064 | Customer Images
44 | 11064 | License File
44 | 11090 | All Warning labels Fitted.
44 | 11235 | License File
79 | 1585 | Check wiring is neat and tidy and properly secured.
79 | 1585 | Check cable ties have been cut back sufficiently and no sharp edges remain.
20 | 10930 | Run the ‘Disable FIB–all GIS.reg’ file
40 | 450 | Configure VP BSD in Administrator
40 | 450 | Input Offset Cal
40 | 450 | Black Level Cal
40 | 450 | Diode Segment Test
40 | 450 | Diode Mirror Test
40 | 450 | All Gain Ranges working
40 | 450 | Check all channels in Invert mode
37 | 11314 | Tablefiller successfully completed.
21 | 256 | Save Full System Checker report
31 | 310 | Blanking Drift Distance Median (Limit <= 300nm) =
31 | 310 | Blanking Drift Time Median (Limit <= 3.5 Sec) =
31 | 310 | Corner Test Max Drift (Limit <= 300nm) =
195 | 11244 | Serial No. (or N/A)
195 | 11244 | No X-Ray Leakage detected.
37 | 11233 | 10kv VP
37 | 11234 | 10kv Widefield
102 | 11077 | ZEN Software Installed
31 | 310 | Blanking Drift Distance Median (Limit <= 300nm) =
31 | 310 | Blanking Drift Time Median (Limit <= 3.5 Sec) =
31 | 310 | Corner Test Max Drift (Limit <= 300nm) =
20 | 11255 | Check SCS Folder/Sales Order
20 | 11255 | SCS Number (If applicable)
44 | 11256 | Check SCS Folder/Sales Order
30 | 11213 | Calibration wizard completed.
40 | 450 | Configure VP BSD in Administrator
40 | 450 | Input Offset Cal
40 | 450 | Black Level Cal
40 | 450 | Diode Segment Test
40 | 450 | Diode Mirror Test
40 | 450 | All Gain Ranges working
40 | 450 | Check all channels in Invert mode
40 | 1555 | Leak check
40 | 1555 | Pneumatics OK
40 | 1555 | Diode Alignment OK
40 | 1555 | Good Repeatability
40 | 1555 | Bias Voltage Working
40 | 1555 | Vac Degradation < 3x10^-15 mbar
40 | 1555 | NanoVP Operation Working over full range
40 | 1555 | XVP Operation working over full range
41 | 343 | Before Cladding the system check that all column connections are securely fastened.
41 | 343 | Ensure Cladding geometry is correct and all parts aligned.
656 | 10664 | Table Top is Correct Version
656 | 10664 | Table Top is Free from Damage
656 | 10664 | Remove Punchout (Dependant on system type)
27 | 286 | Angle =
27 | 11072 | Low mag mode off Angle =
27 | 11072 | Low mag mode on Angle =
30 | 299 | Macro successfully completed.
30 | 299 | Save Image to Image directory.
20 | 11270 | Predictive Service S/W installed on Customer PC.
46 | 1567 | Remove the sample holder from the system.
46 | 1567 | Remove Docking Brackets and fit Transit clamps.
46 | 1567 | Airfeet deflated and free to move.
46 | 1567 | Backing lines re-configured so that PVSV valve is at-tached to flexi hose. Secure valve to frame.
46 | 1567 | If applicable, ensure that the A/L push rod is removed and blank fitted in its place. Ensure A/L is pumped down before powering down system.
46 | 1567 | Select 'Partial vent on standby' in Smartsem.
22 | 262 | Carry out overnight ‘Full service Bakeout’ and note actual duration.
22 | 262 | Confirm that all Thermocouples are reading.
22 | 262 | Confirm that all Ion pump heater elements and flange heaters are getting hot.
33 | 11062 | 1kV =
33 | 11062 | 15kV =
46 | 1568 | Power system down.
46 | 1568 | Blow water out of the cooling system.
46 | 1568 | Chamber secured to vibration plate
46 | 1568 | Secure Airfeet with foam blocks.
46 | 1568 | Fit foam blocks around column cladding.
46 | 1568 | Secure Turbo pump with foam blocks.
46 | 1568 | Fit Covers to all unused connectors.
46 | 1568 | Remove Cekon Plug From Mains Cable
46 | 1568 | Fit Live and Neutral Labels to Mains Cable
46 | 1568 | Mains and Earth cable stowed on Mains unit.
46 | 1568 | Fit shipping Eyes to transit blocks.
46 | 1568 | Confirm all column cladding is secure and screws Loctited in place.
46 | 1568 | Put thread lock on pendulum screws and transit block Screws.
46 | 1568 | All doors closed and secured.
46 | 1568 | All plinth panels removed and stored in transit trolley.
46 | 1568 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 1568 | Print shipping stickers.
46 | 1568 | Fit M8 screw to secure stage door to chamber.
40 | 1553 | VP Mode Macro selects VPSE when going to VP
40 | 1553 | Collector Bias working
40 | 1553 | Collector Bias breakdown occurs at >290v @ 30Pa.
40 | 1553 | Take Functional Image.
40 | 11084 | VP Mode Macro selects C2D when going to VP
40 | 11084 | Black Level calibration
40 | 11084 | Functional Checks ok
40 | 11084 | Take Functional Image
40 | 11273 | Correct Penning Guage Fitted
40 | 11273 | Chamber Pressure Regulation good
26 | 10965 | ESB Serial Number.
40 | 453 | Check each individual segment in normal and inverse modes
40 | 453 | Analyze Segments
40 | 453 | Gain Ranges
40 | 453 | Pneumatics
40 | 453 | Alignment
40 | 453 | Stage Safe Position
40 | 453 | Vac Degradation < 3x10-5 mbar
40 | 453 | CZ Detector Cal
40 | 453 | Store Holder Position 1
25 | 275 | Are Gun Head clamping screws secure?
25 | 275 | Are Condensor clamping screws secure?
25 | 275 | Save Emission image @1Kv (cond off)
25 | 275 | Save Emission image @10Kv (cond off)
25 | 11313 | Are Gun Head clamping screws secure?
25 | 11313 | Are Condensor clamping screws secure?
25 | 11313 | Save Emission image @1Kv (cond off)
25 | 11313 | Save Emission image @10Kv (cond off)
38 | 337 | Blanking Drift Distance Median (Limit <= 300nm) =
38 | 337 | Blanking Drift Time Median (Limit <= 3.5 Sec) =
38 | 337 | Corner Test Max Drift (Limit <= 300nm) =
38 | 11315 | Confirm that when the beam is blanked the measured I Probe = 0+/- 5pA
38 | 11315 | Difference between Eht on and off < 1pA.
27 | 11316 | No Image rotation with WD change
22 | 266 | Test Power Supply Serial Number
22 | 266 | Use the 35kV External supply for 30mins.
22 | 266 | Save Gun Moitor Log as GunConditioning_TIP SN
20 | 11270 | Predictive Service S/W installed on Customer PC.
21 | 256 | Save Full System Checker report
21 | 11321 | Correct Firmware installed
41 | 343 | Before Cladding the system check that all column connections are securely fastened.
41 | 343 | Ensure Cladding geometry is correct and all parts aligned.
40 | 11325 | Installation
40 | 11325 | Functional Test
28 | 409 | Gun X
28 | 409 | Gun Y
28 | 409 | Aper X
28 | 409 | Aper Y
30 | 297 | Check Mag cal errors are <1% in X & Y axes in all 3 Mag Ranges. Orthog. ± 1deg NOTE Gemini 560 is +-2% in all ranges
30 | 297 | Save 4 Images ( 3 Mag Ranges + Ortho.)
38 | 337 | Blanking Drift Distance Median (Limit <= 300nm) =
38 | 337 | Blanking Drift Time Median (Limit <= 3.5 Sec) =
38 | 337 | Corner Test Max Drift (Limit <= 300nm) =
31 | 306 | Actual Gun Axis X =
31 | 306 | Actual Gun Axis Y =
30 | 11069 | Dynamic focus fuction check - Whole image in focus
30 | 11069 | Y - Axis measurement (400μm +2%)
30 | 11069 | Orthogonality
45 | 11337 | Create Directory C:\Support\Manuals
45 | 11337 | Create Shortcut to above Directory on Desktops, called ‘Carl Zeiss User Documentation’.
45 | 11338 | C:\Support\Manuals
45 | 11338 | Customer USB documentation folder
45 | 11339 | Copy SmartSEM instruction manual from Customer USB to C:\Support\Manuals
27 | 434 | Image Shift (um) =
27 | 435 | Focus Deviation (um) =
27 | 436 | Focus Deviation (um) =
43 | 11341 | 1kV ESB
43 | 347 | 1kV Inlens
43 | 347 | 15kV Inlens
29 | 11381 | Run overnight Photograph Macro
29 | 11381 | Analyse images for drift, breakdown, ESB issues
668 | 10777 | Fitted To Correct ports
668 | 10777 | Connected to Correct ASAP Ports & Labeled
27 | 284 | 20nA-Emmission rotates CW with I increase
40 | 10868 | Connect Silver Box and Cables
40 | 10868 | Check Altas Test USB Stick Has The latest Version Of Software
40 | 10868 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
40 | 10868 | Request Licence File from Fibics
40 | 10868 | Install Atlas Software
40 | 10868 | Install Licence File
40 | 10868 | Calibrate SEM
40 | 10868 | Obtain Mosaic
40 | 10868 | Set ATLAS Rotation Offset
40 | 10868 | Save ATLAS Licence Document
27 | 289 | C2 Calibration Value
27 | 289 | C3 Calibration Value
27 | 434 | Image Shift (um) =
38 | 337 | Blanking Drift Distance Median (Limit <= 300nm) =
38 | 337 | Blanking Drift Time Median (Limit <= 3.5 Sec) =
38 | 337 | Corner Test Max Drift (Limit <= 300nm) =
30 | 297 | Check Mag cal errors are <1% in X & Y axes in all 3 Mag Ranges. Orthog. ± 1deg
30 | 297 | Save 4 Images ( 3 Mag Ranges + Ortho.)
43 | 347 | 1kV Inlens
43 | 347 | 15kV Inlens
43 | 11353 | 1kV Inlens
43 | 11353 | 15kV Inlens
43 | 11353 | 200V Inlens
43 | 11353 | 30KV SE2
102 | 11078 | Stage Moves To Correct Position When Arm Is Opened
102 | 11078 | Image Is Captured
102 | 11078 | Functional Check
40 | 10868 | Connect Silver Box and Cables
40 | 10868 | Check Altas Test USB Stick Has The latest Version Of Software
40 | 10868 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
40 | 10868 | Request Licence File from Fibics
40 | 10868 | Install Atlas Software
40 | 10868 | Install Licence File
40 | 10868 | Calibrate SEM
40 | 10868 | Obtain Mosaic
40 | 10868 | Set ATLAS Rotation Offset
40 | 10868 | Save ATLAS Licence Document
40 | 10868 | Connect Silver Box and Cables
40 | 10868 | Check Altas Test USB Stick Has The latest Version Of Software
40 | 10868 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
40 | 10868 | Request Licence File from Fibics
40 | 10868 | Install Atlas Software
40 | 10868 | Install Licence File
40 | 10868 | Calibrate SEM
40 | 10868 | Obtain Mosaic
40 | 10868 | Set ATLAS Rotation Offset
40 | 10868 | Save ATLAS Licence Document
40 | 10868 | Connect Silver Box and Cables
40 | 10868 | Check Altas Test USB Stick Has The latest Version Of Software
40 | 10868 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
40 | 10868 | Request Licence File from Fibics
40 | 10868 | Install Atlas Software
40 | 10868 | Install Licence File
40 | 10868 | Calibrate SEM
40 | 10868 | Obtain Mosaic
40 | 10868 | Set ATLAS Rotation Offset
40 | 10868 | Save ATLAS Licence Document
40 | 10868 | Connect Silver Box and Cables
40 | 10868 | Check Altas Test USB Stick Has The latest Version Of Software
40 | 10868 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
40 | 10868 | Request Licence File from Fibics
40 | 10868 | Install Atlas Software
40 | 10868 | Install Licence File
40 | 10868 | Calibrate SEM
40 | 10868 | Obtain Mosaic
40 | 10868 | Set ATLAS Rotation Offset
40 | 10868 | Save ATLAS Licence Document
42 | 344 | Actual X-Axis
42 | 344 | Actual Y-Axis
42 | 11013 | Actual X-Axis
42 | 11013 | Actual Y-Axis
40 | 10868 | Connect Silver Box and Cables
40 | 10868 | Check Altas Test USB Stick Has The latest Version Of Software
40 | 10868 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
40 | 10868 | Request Licence File from Fibics
40 | 10868 | Install Atlas Software
40 | 10868 | Install Licence File
40 | 10868 | Calibrate SEM
40 | 10868 | Obtain Mosaic
40 | 10868 | Set ATLAS Rotation Offset
40 | 10868 | Save ATLAS Licence Document
40 | 10868 | Connect Silver Box and Cables
40 | 10868 | Check Altas Test USB Stick Has The latest Version Of Software
40 | 10868 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
40 | 10868 | Request Licence File from Fibics
40 | 10868 | Install Atlas Software
40 | 10868 | Install Licence File
40 | 10868 | Calibrate SEM
40 | 10868 | Obtain Mosaic
40 | 10868 | Set ATLAS Rotation Offset
40 | 10868 | Save ATLAS Licence Document
102 | 1576 | Function checked in VP mode, signal intensity changes with VP pressure.
102 | 1576 | Signal intensity changes with collector voltage.
41 | 343 | Ensure Cladding geometry is correct and all parts aligned.
| 11391 | In-Lens and ESB
| 11391 | EO Dynamic PCB
| 11391 | EO Static PCB
| 11391 | Scan Gen PCB (ASAP or SVSEM)
| 11391 | Check that all column connections are securely fastened.
718 | 11392 | Heater & Thermocouple cables plugged into AC unit
718 | 11392 | Column mounting bolts all secured
718 | 11392 | IGP bracket screws all secured
718 | 11392 | All column cabling cable tied neatly
38 | 337 | Blanking Drift Distance Median (Limit <= 300nm) =
38 | 337 | Blanking Drift Time Median (Limit <= 3.5 Sec) =
38 | 337 | Corner Test Max Drift (Limit <= 300nm) =
42 | 11047 | Actual X-Axis
42 | 11047 | Actual Y-Axis
32 | 314 | Normal mode (5mm WD)
32 | 314 | High Current mode (5mm WD)
46 | 1567 | Remove the sample holder from the system.
46 | 1567 | Remove Docking Brackets and fit Transit clamps.
46 | 1567 | Airfeet deflated and free to move.
46 | 1567 | Backing lines re-configured so that PVSV valve is at-tached to flexi hose. Secure valve to frame.
46 | 1567 | If applicable, ensure that the A/L push rod is removed and blank fitted in its place. Ensure A/L is pumped down before powering down system.
46 | 1567 | Select 'Partial vent on standby' in Smartsem.
27 | 11072 | Low mag mode off Angle =
27 | 11072 | Low mag mode on Angle =
25 | 427 | Mechanically align the column using the Merlin Alignment Wizard
46 | 1568 | Power system down.
46 | 1568 | Blow water out of the cooling system.
46 | 1568 | Chamber secured to vibration plate
46 | 1568 | Secure Airfeet with foam blocks.
46 | 1568 | Fit foam blocks around column cladding.
46 | 1568 | Secure Turbo pump with foam blocks.
46 | 1568 | Fit Covers to all unused connectors.
46 | 1568 | Remove Cekon Plug From Mains Cable
46 | 1568 | Fit Live and Neutral Labels to Mains Cable
46 | 1568 | Mains and Earth cable stowed on Mains unit.
46 | 1568 | Fit shipping Eyes to transit blocks.
46 | 1568 | Confirm all column cladding is secure and screws Loctited in place.
46 | 1568 | Put thread lock on pendulum screws and transit block Screws.
46 | 1568 | All doors closed and secured.
46 | 1568 | All plinth panels removed and stored in transit trolley.
46 | 1568 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 1568 | Print shipping stickers.
46 | 1568 | Fit M8 screw to secure stage door to chamber.
46 | 11407 | Power system down.
46 | 11407 | Blow water out of the cooling system.
46 | 11407 | Chamber secured to vibration plate
46 | 11407 | Secure Airfeet with foam blocks.
46 | 11407 | Secure Turbo pump with foam blocks.
46 | 11407 | Fit Covers to all unused connectors.
46 | 11407 | Remove Cekon Plug From Mains Cable
46 | 11407 | Fit Live and Neutral Labels to Mains Cable
46 | 11407 | Mains and Earth cable stowed on Mains unit.
46 | 11407 | Fit shipping Eyes to transit blocks.
46 | 11407 | Put thread lock on pendulum screws and transit block Screws.
46 | 11407 | All doors closed and secured.
46 | 11407 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11407 | Print shipping stickers.
655 | 10662 | Check All Crimps Are Correctly Located And Secure.
655 | 10662 | Test PSU Ensure All LED’s Light Up Correctly
46 | 11408 | Remove the sample holder from the system.
46 | 11408 | Remove Test Tool SE2 and fit to warehouse stand. Fit SE2 blank plate to chamber and SHV connector to Spellman (00000-0302-367).
46 | 11408 | Confirm that Raith top mounted SE2 detector is mounted to a warehouse stand and that a protective cap is fitted over the light guide.
46 | 11408 | Set Stage position to Z=41mm & X & Y =65mm.
46 | 11408 | Remove SCM and pack for shipment (include screws)
46 | 11408 | Remove Docking Brackets and fit Transit clamps.
46 | 11408 | Airfeet deflated and free to move.
46 | 11408 | Backing lines re-configured so that PVSV valve is attached to flexi hose. Secure valve to frame.
46 | 11408 | Select 'Partial vent on standby' in Smartsem.
25 | 11035 | With Cond. Off Max Aper. X & Y < 45
25 | 11035 | Gun Head Clamping Screws Secure
25 | 11035 | Condenser Clamping Screws Secure
25 | 11036 | Large Aperture in Top Left Quadrant
25 | 11036 | Save Emission Image with Condenser On
25 | 11036 | Save Emission Image with Condenser Off
27 | 286 | Stage Angle =
656 | 10668 | VAC PCB S/No
46 | 11407 | Power system down.
46 | 11407 | Blow water out of the cooling system.
46 | 11407 | Chamber secured to vibration plate
46 | 11407 | Check that the Stage bearing blocks screws are secure.
46 | 11407 | Secure Airfeet with foam blocks.
46 | 11407 | Secure Turbo pump with foam blocks.
46 | 11407 | Fit Covers to all unused connectors.
46 | 11407 | Remove Cekon Plug From Mains Cable
46 | 11407 | Fit Live and Neutral Labels to Mains Cable
46 | 11407 | Mains and Earth cable stowed on Mains unit.
46 | 11407 | Fit shipping Eyes to transit blocks.
46 | 11407 | Put thread lock on pendulum screws and transit block Screws.
46 | 11407 | All doors closed and secured.
46 | 11407 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11407 | Print shipping stickers.
44 | 349 | Carry out safety tests using the Secutest S111 pat Tester.
44 | 349 | Fill out completely the Product Safety Tests form in Section 13.
44 | 349 | Carry out safety tests using the Seaward Prime 100 Tester.
44 | 349 | Fill out completely the Product Safety Tests form in Section 13.
26 | 281 | Type in measured value (Limit <= 30um)
45 | 362 | Em Server.db
45 | 362 | Em Server.txt Log
45 | 362 | Optiprobe.txt Log
45 | 362 | Smart Backup.log
45 | 362 | L500.ini
45 | 362 | License file.xml
45 | 362 | Registry file.reg
45 | 362 | CalibSc.dat
45 | 11024 | Em Server.db
45 | 11024 | Em Server.txt Log
45 | 11024 | Smart Backup.log
45 | 11024 | L500.ini
45 | 11024 | License file.xml
45 | 11024 | Registry file.reg
45 | 11024 | CalibDC.dat
45 | 11024 | Export Tablefiller
45 | 11058 | Em Server.db
45 | 11058 | Em Server.log Text File
45 | 11058 | Optiprobe Calibration.Log Text File
45 | 11058 | Mybackup.Smartbackup
45 | 11058 | Image Folder
45 | 11058 | L500.ini
45 | 11058 | Dynamic EHT.MN2
45 | 11058 | Probe Stability.MN2
45 | 11058 | License file.xml
45 | 11058 | Registry file.reg
45 | 11058 | System Checker Full Report
45 | 11058 | Export Alignment Data From Table Editor(.MERLINCDB)
24 | 273 | Level Frame
24 | 273 | Check pendulum height (with spacer)
24 | 273 | Set up ‘float’ height
24 | 273 | Check inlet / outlet valves
24 | 273 | Set up docking pins
30 | 11071 | Check Mag cal errors are <2% in X & Y axes in all 3 Mag Ranges. Orthog. ± 1deg
30 | 11071 | Save 4 Images ( 3 Mag Ranges + Ortho.)
656 | 10669 | ASAP S/No
46 | 1568 | Power system down.
46 | 1568 | Blow water out of the cooling system.
46 | 1568 | Chamber secured to vibration plate
46 | 1568 | Chamber Sercuring Screws Tightened to 25NM
46 | 1568 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 1568 | Secure Airfeet with foam blocks.
46 | 1568 | Fit foam blocks around column cladding.
46 | 1568 | Secure Turbo pump with foam blocks.
46 | 1568 | Fit Covers to all unused connectors.
46 | 1568 | Remove Cekon Plug From Mains Cable
46 | 1568 | Fit Live and Neutral Labels to Mains Cable
46 | 1568 | Mains and Earth cable stowed on Mains unit.
46 | 1568 | Fit shipping Eyes to transit blocks.
46 | 1568 | Confirm all column cladding is secure and screws Loctited in place.
46 | 1568 | Put thread lock on pendulum screws and transit block Screws.
46 | 1568 | All doors closed and secured.
46 | 1568 | All plinth panels removed and stored in transit trolley.
46 | 1568 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 1568 | Print shipping stickers.
46 | 1568 | Fit M8 screw to secure stage door to chamber.
46 | 11123 | Power system down.
46 | 11123 | Blow water out of the cooling system.
46 | 11123 | Chamber secured to vibration plate
46 | 11123 | Chamber Sercuring Screws tightened to 25NM
46 | 11123 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 11123 | Secure Airfeet with foam blocks.
46 | 11123 | Fit foam blocks around column cladding.
46 | 11123 | Secure Turbo pump with foam blocks.
46 | 11123 | Fit Covers to all unused connectors.
46 | 11123 | Mains and Earth cable stowed on Mains unit.
46 | 11123 | Fit shipping Eyes to transit blocks.
46 | 11123 | Confirm all column cladding is secure and screws Loctited in place.
46 | 11123 | Put thread lock on pendulum screws and transit block Screws.
46 | 11123 | All doors closed and secured.
46 | 11123 | All plinth panels removed and stored in transit trolley.
46 | 11123 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11123 | Print shipping stickers.
46 | 11123 | Fit M8 screw to secure stage door to chamber.
46 | 11407 | Power system down.
46 | 11407 | Blow water out of the cooling system.
46 | 11407 | Chamber secured to vibration plate
46 | 11407 | Chamber Sercuring Screws Tightened to 25NM
46 | 11407 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 11407 | Check that the Stage bearing blocks screws are secure.
46 | 11407 | Secure Airfeet with foam blocks.
46 | 11407 | Secure Turbo pump with foam blocks.
46 | 11407 | Fit Covers to all unused connectors.
46 | 11407 | Remove Cekon Plug From Mains Cable
46 | 11407 | Fit Live and Neutral Labels to Mains Cable
46 | 11407 | Mains and Earth cable stowed on Mains unit.
46 | 11407 | Fit shipping Eyes to transit blocks.
46 | 11407 | Put thread lock on pendulum screws and transit block Screws.
46 | 11407 | All doors closed and secured.
46 | 11407 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11407 | Print shipping stickers.
26 | 282 | Type in measured value (limit: <= 15um)
26 | 282 | Save image showing concentricity to image directory
26 | 282 | Type in measured value (limit: <= 15um)
26 | 282 | Save image showing concentricity to image directory
30 | 297 | Check Mag cal errors are <1% in X & Y axes in all 3 Mag Ranges. Orthoganality ± 1deg.
30 | 297 | Save 4 Images ( 3 Mag Ranges + Orthoganality)
670 | 10789 | All Connections Are Secured with Screws
670 | 10789 | Cables Are Neatly Routed and Secured
46 | 1567 | Remove the sample holder from the system.
46 | 1567 | Remove Docking Brackets and fit Transit clamps.
46 | 1567 | Airfeet deflated and free to move.
46 | 1567 | Backing lines re-configured so that PVSV valve is at-tached to flexi hose. Secure valve to frame.
46 | 1567 | If applicable, ensure that the A/L push rod is removed and blank fitted in its place. Ensure A/L is pumped down before powering down system.
46 | 1567 | If applicable, ensure that the Beamsleeve funnel has been removed from the detector
46 | 1567 | Select 'Partial vent on standby' in Smartsem.
46 | 11123 | Power system down.
46 | 11123 | Blow water out of the cooling system.
46 | 11123 | Chamber secured to vibration plate
46 | 11123 | Chamber Sercuring Screws tightened to 25NM
46 | 11123 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 11123 | Secure Airfeet with foam blocks.
46 | 11123 | Secure Turbo pump with foam blocks.
46 | 11123 | Fit Covers to all unused connectors.
46 | 11123 | Mains and Earth cable stowed on Mains unit.
46 | 11123 | Fit shipping Eyes to transit blocks.
46 | 11123 | Put thread lock on pendulum screws and transit block Screws.
46 | 11123 | All doors closed and secured.
46 | 11123 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11123 | Print shipping stickers.
46 | 11123 | Fit 8x M6 x 55mm cap head screws to stage door
22 | 260 | System Configuration correct?
22 | 260 | Build quality satisfactory?
37 | 11234 | 10kv Widefield Table entries complete for all Apertures?
37 | 11234 | Beamshift calibration complete for all Apertures.?
41 | 343 | Ensure Cladding geometry is correct and all parts aligned.
41 | 343 | Ensure the suppliers Cladding list for both kits is put into the customers pack.
21 | 255 | Air (Pressure = 6 Bar)
21 | 255 | Water
21 | 255 | Mains
21 | 255 | Safety Earth
21 | 255 | N2
21 | 11427 | Air (Pressure = 6 Bar)
21 | 11427 | Water
21 | 11427 | Mains
21 | 11427 | Safety Earth
21 | 11427 | N2
21 | 11427 | Remove Blank and fit Test Tool Stage
21 | 11427 | Fit Test Tool Stage Controller and cable up
21 | 11427 | Mount Test Tool SE2 Detector onto chamber
21 | 11427 | Mount PM Tube (with Immersion oil) onto SE2 Detector
21 | 11427 | Mount SCM to port on Stage Door
25 | 11035 | With Cond. Off Max Aper. X & Y < 30%
25 | 11035 | Gun Head Clamping Screws Secure
25 | 11035 | Condenser Clamping Screws Secure
27 | 289 | C2 Calibration Value
27 | 289 | C3 Calibration Value
27 | 435 | Focus Deviation (um) =
27 | 436 | Focus Deviation (um) =
81 | 1571 | Measured value =
81 | 1587 | Measured value =
656 | 10665 | Earth Plate
656 | 10665 | HV Cable Plate
656 | 10665 | Bellows
656 | 10671 | Stage PCB S/No
656 | 10675 | Specimen current test
671 | 10799 | Internal Surfaces
671 | 10799 | Tapped Holes
656 | 10784 | Turbo Controller Serial Number
670 | 10785 | EHT Cable PL5 is Plugged Into PL5
670 | 10785 | All Connectors Are Securely Fitted
670 | 10785 | Filter Earth Link Fitted
670 | 10785 | Cables Are Neatly Routed And Secured
670 | 10786 | All Connections Are Secured with Screws
670 | 10786 | Cables Are Neatly Routed and Secured
670 | 10787 | All Connections Are Secured with Screws
670 | 10787 | Cables Are Neatly Routed and Secured
670 | 10787 | Correct Version Of EO PCB Fitted
670 | 10788 | All Connections Are Secured with Screws
670 | 10788 | Cables Are Neatly Routed and Secured
670 | 10790 | Cables Are Neatly Routed and Secured
670 | 10790 | Plinth Is free From Debris
670 | 10790 | Table Top Is Free From Damage
670 | 10790 | Table Top Is Secured
670 | 10790 | EHT Set Is Secured
670 | 10790 | Clean and Tidy Assembly Bay
671 | 10794 | Chamber Is Free From Defects
671 | 10795 | Screws Fitted
671 | 10797 | Torque Chamber Bolts To 35Nm
671 | 10797
671 | 10798 | Internal Surfaces
671 | 10798 | Tapped Holes
671 | 10803 | TIV Serial Number
671 | 10803 | Function Check
671 | 10803 | Fit TIV
671 | 10808 | Turbo Pump Serial Number
672 | 10811 | All Vacuum Clamps are Tight
672 | 10811 | All Valves Are Fitted Corrctly
672 | 10811 | All Valves Are correctly Connected
672 | 10812 | Check Correct Routing Of Air Pipes
672 | 10812 | Check Correct Orientation Of Air Valves
672 | 10813 | Check Torque Of Chamber Bolts
672 | 10813 | Chamber Free From Any Detects
672 | 10813 | Screws Fitted To Internal Holes
672 | 10814 | Free From Any Defects
672 | 10814 | Table Top Secured
674 | 10816 | Correct Chamber Bung Fitted
674 | 10819 | Scan Cables
674 | 10819 | Lens Cables
674 | 10819 | CIV Cable
674 | 10819 | Penning Cable
668 | 10778 | Air Pipes Are Correctly Routed and Kink Free
674 | 10821 | Heater Cables
674 | 10821 | Ion Pump Supply Cables
674 | 10821 | EHT Cables
674 | 10821 | Liner Cable
674 | 10821 | Column Earth Cable
674 | 10822 | Bracket
674 | 10822 | Supply
674 | 10822 | Heater Blocks
674 | 10823 | 24V Supply Kit
674 | 10823 | Spellman EHT Set
674 | 10823 | EHT Set Serial Number
674 | 10823 | Earth Connection Torqued to 3Nm
675 | 10824 | All Cables Are Neatly Routed
675 | 10824 | All Plugs Are Secured Correctly
675 | 10824 | All Earth Cables Are Tightly Secured
676 | 10826 | Serial Number
676 | 10826 | Shim Stage Door
676 | 10827 | Stage Cable
676 | 10827 | Chamberscope
676 | 10827 | Specimen Touch Cable
668 | 10780 | Gun Switch Cable is Pluuged into VICE PCB
677 | 10844 | All Cables Are Neatly Routed
677 | 10844 | All Plugs Are Secured Correctly
677 | 10844 | All Earth Cables Are Tightly Secured
677 | 10844 | Specimen Current Test
677 | 10845 | Fitted To Correct Ports
677 | 10845 | Detectors Connected To Correct ASAP Port And Labelled
678 | 10847 | HAL Test
678 | 10847 | PAT Test
678 | 10849 | Rotary Pump Pumping
678 | 10849 | Penning Gauge LED is On
678 | 10849 | Light Is Lit On EHT Set
678 | 10849 | The Distance Between The Vac Cable Loom AND The Plinth Base Is Greater Than 30mm
678 | 10850 | Vacuum PCB Test Passed
678 | 10850 | EO PCB Test Passed
678 | 10850 | Stage PCB Test Passed
678 | 10851 | ASAP FW/HW Present
678 | 10851 | Stage FW/HW Present
678 | 10851 | Vacuum FW/HW Present
678 | 10851 | EO PCB FW/HW Present
678 | 10851 | Vacuum Reading OK
678 | 10851 | Initialise Stage
678 | 10852 | PSU Tests Passed
678 | 10852 | EO Tests Passed
678 | 10853 | Vac 5.15V (P1)
678 | 10853 | Vac 5.00V (P10)
678 | 10853 | Stage 5.15V
678 | 10853 | Stage 21.5V
678 | 10853 | Stage -21.5V
678 | 10853 | EO 5.15V
678 | 10854 | Stage Links Set Correctly
678 | 10855 | No Leaks Detected
678 | 10856 | No Leaks Detected
678 | 10857 | No Leaks Detected
679 | 10858 | Picking List
679 | 10858 | Turbo Pump
679 | 10858 | Turbo Pump Controller
679 | 10858 | Chamber
679 | 10858 | Plinth
679 | 10858 | EO PCB
679 | 10858 | Stage PCB
679 | 10858 | Stage Logbook
679 | 10859 | All Cables Are Neatly Routed
679 | 10859 | All Plugs Are Secured Correctly
679 | 10859 | All Earth Cables Are Tightly Secured
679 | 10859 | Check Rear Panel Cables Are Correctly Located
679 | 10859 | Check Opto Devices Are Not Damaged On All PCBs & Mains Box
679 | 10859 | Check Stage Board Link Configuration
679 | 10860 | Fitted To Correct ports
679 | 10860 | Connected to Correct ASAP Ports & Labeled
679 | 10861 | Air Pipes Are Correctly Routed and Kink Free
679 | 10861 | Pressure Regulators Are Fitted In THe Correct Orientation
679 | 10861 | WaterPipes Are Correctly Routed And Free From Kinks
679 | 10862 | Chamber is Free From Defects
679 | 10862 | Chamber Meshes Are Fitted
679 | 10862 | Chamber Pumping Pipe Is Fitted Centrally To The Table Top
679 | 10862 | Table Top Is Free From Defects
679 | 10863 | Plinth Base Is Clean and Free From Debris
679 | 10864 | Clean & Tidy Assembly Bay
656 | 10663 | Damper Plate is Correct Version
656 | 10663 | Damper Plate is Located Correctlu
656 | 10664 | Table Top is Correct Version
656 | 10664 | Table Top is Free from Damage
656 | 10664 | Remove Punchout (Dependant on system type)
656 | 10670 | EO PCB S/No
656 | 10670 | Confirm Material No. is $[SigmaEOMaterial]
671 | 10793 | Turbo Stack Assembled
671 | 10793 | Turbo M8 Bolts Torqued to 7Nm
659 | 11106 | Serial Number
655 | 10660 | Serial Number
656 | 11105 | Serial Number
655 | 10662 | Check All Crimps Are Correctly Located And Secure.
655 | 10662 | Test PSU Ensure All LED’s Light Up Correctly
655 | 10662 | To be Checked By Another Engineer - Name:
656 | 11117 | Check SCS Folder/Sales Order
679 | 11118 | Check SCS Folder/Sales Order
656 | 10670 | EO PCB S/No
656 | 10670 | Correct Version Of EO PCB Fitted
670 | 10785 | EHT Cable PL5 is Plugged Into PL10
670 | 10785 | All Connectors Are Securely Fitted
670 | 10785 | Filter Earth Link Fitted
670 | 10785 | Cables Are Neatly Routed And Secured
670 | 10787 | All Connections Are Secured with Screws
670 | 10787 | Cables Are Neatly Routed and Secured
670 | 10787 | Correct Version Of EO PCB Fitted
678 | 10847 | HAL Test
678 | 10847 | PAT Test
671 | 10797 | Torque Chamber Bolts To 35Nm
656 | 11117 | Check SCS Folder/Sales Order
679 | 11118 | Check SCS Folder/Sales Order
656 | 11117 | Check SCS Folder/Sales Order
656 | 11117 | Check SCS Folder/Sales Order
656 | 11117 | SCS Number (If applicable)
676 | 11194 | Check that the Assembly is the Correct Part Number
676 | 11194 | Correct Collector Beak Fitted
676 | 11194 | Fit Detector Assembly
676 | 11194 | Apply Emersion Oil to PM Tube Face
676 | 11194 | Fit PM Head Amp Assembly
655 | 10660 | Serial Number
656 | 11105 | Serial Number
656 | 11105 | Serial Number
676 | 11194 | Check that the Assembly is the Correct Part Number
676 | 11194 | Correct Collector Beak Fitted
676 | 11194 | Fit Detector Assembly
676 | 11194 | Apply Emersion Oil to PM Tube Face
676 | 11194 | Fit PM Head Amp Assembly
678 | 11248 | Print Plinth Seperators
678 | 11248 | Sort Documentation with Seperators
676 | 11268 | Detector Type
676 | 11268 | Detector Serial Number
676 | 11268 | Detector Type
676 | 11268 | Detector Serial Number
717 | 11386 | Time =
717 | 11386 | Date =
717 | 11388 | Time =
717 | 11388 | Date =
676 | 11400 | Fit Chamberscope
658 | 10680 | Serial Number
658 | 10681 | Cable Up PSU
658 | 10682 | Check All Crimps Are Correctly Located And Secure.
658 | 10682 | Test PSU Ensure All LED’s Light Up Correctly
659 | 10683 | Damper Plate is Correct Version
659 | 10683 | Damper Plate is Located Correctlu
659 | 10684 | Table Top is Correct Version
659 | 10684 | Table Top is Free from Damage
659 | 10684 | Remove Punchout (Dependant on system type)
659 | 10685 | Earth Plate
659 | 10685 | Cable Plates
659 | 10685
659 | 10685 | Belllows
659 | 10686 | Turbo Controller Serial Number
659 | 10689 | Vacuum PCB Serial Number
659 | 10690 | ASAP Serial Number
377 | 3261 | Bake Out Hours
659 | 10691 | EO PCB Serial Number
659 | 10692 | Stage PCB Serial Number
659 | 10695 | Install PSU Loom
659 | 10695
659 | 10695 | Specimen Current Test
660 | 10701 | EHT Cable PL5 is Plugged Into PL5
660 | 10701 | All Connectors Are Securely Fitted
660 | 10701 | Filter Earth Link Fitted
660 | 10701 | Cables Are Neatly Routed And Secured
660 | 10702 | All Connections Are Secured with Screws
660 | 10702 | Cables Are Neatly Routed and Secured
660 | 10703 | All Connections Are Secured with Screws
660 | 10703 | Cables Are Neatly Routed and Secured
660 | 10704 | All Connections Are Secured with Screws
660 | 10704 | Cables Are Neatly Routed and Secured
660 | 10705 | All Connections Are Secured with Screws
660 | 10705 | Cables Are Neatly Routed and Secured
660 | 10706 | Cables Are Neatly Routed and Secured
660 | 10706 | Plinth Is free From Debris
660 | 10706 | Table Top Is Free From Damage
660 | 10706 | Table Top Is Secured
660 | 10706 | EHT Set Is Secured
660 | 10706 | Clean and Tidy Assembly Bay
661 | 10710 | Torque Chamber Bolts to 35Nm
390 | 3317 | Pressure (<3e-6mbar)
661 | 10711 | Internal Surfaces
661 | 10711 | Tapped Holes
661 | 10712 | Internal Surfaces
661 | 10712 | Tapped Holes
661 | 10716 | TIV Serial Number
661 | 10716 | Function Check
661 | 10716 | Fit TIV
661 | 10720 | Turbo Pump Serial Number
662 | 10722 | All Vacuum Clamps are Tight
662 | 10722 | All Valves Are Fitted Corrctly
662 | 10722 | All Valves Are correctly Connected
662 | 10723 | Check Correct Routing Of Air Pipes
662 | 10723 | Check Correct Orientation Of Air Valves
662 | 10724 | Check Torque Of Chamber Bolts
662 | 10724 | Chamber Free From Any Detects
662 | 10725 | Free From Any Defects
662 | 10725 | Table Top Secured
663 | 10728 | Scan Cables
663 | 10728 | Lens Cables
663 | 10728 | EHT Cables
663 | 10728 | Penning Cable
663 | 10730 | Valves Cables
663 | 10730 | Heater Cables
663 | 10730 | Ion Pump Cable
663 | 10730 | Ion Pump PSU Serial Number
663 | 10730 | Fit Ion Pump PSU Bracket
663 | 10730 | Fit Ion Pump PSU
663 | 10731 | EHT Serial Number
663 | 10731 | Torque Earth Connection to 3Nm
664 | 10732 | All Cables Are Neatly Routed
664 | 10732 | All Plugs Are Secured Correctly
664 | 10732 | All Earth Cables Are Tightly Secured
664 | 10733 | Clean and Tidy Assembly Bay
665 | 10734 | Stage Serial Number
665 | 10734 | Shim Stage Door
665 | 10734 | Set Stage Clamp Height
665 | 10735 | Fit Stage Cable
665 | 10735 | Fit Specimen Touch Cable
666 | 10754 | All Cables Are Neatly Routed
666 | 10754 | All Plugs Are Secured Correctly
666 | 10754 | All Earth Cables Are Tightly Secured
666 | 10755 | Fitted To Correct Ports
666 | 10755 | Detectors Connected To Correct ASAP Port And Labelled
666 | 10756 | Clean and tidy Assembly Bay
667 | 10765 | HAL Test
667 | 10765 | PAT Test
667 | 10767 | Rotary Pump Pumping
667 | 10767 | Penning Gauge LED is On
667 | 10768 | Vacuum PCB Test Passed
667 | 10768 | EO PCB Test Passed
667 | 10768 | Stage PCB Test Passed
659 | 10685 | Earth Plate
659 | 10685 | Cable Plates
659 | 10685
659 | 10685 | Bellows
667 | 10769 | ASAP FW/HW Present
667 | 10769 | Stage FW/HW Present
667 | 10769 | Vacuum FW/HW Present
667 | 10769 | EO PCB FW/HW Present
667 | 10769 | Vacuum Reading OK
667 | 10769 | Initialise Stage
667 | 10770 | PSU Tests Passed
667 | 10770 | EO Tests Passed
667 | 10771 | Vac 5.15V (P1)
667 | 10771 | Vac 5.00V (P10)
667 | 10771 | Stage 5.15V
667 | 10771 | Stage 21.5V
667 | 10771 | Stage -21.5V
667 | 10771 | EO 5.15V
667 | 10772 | Stage Links Set Correctly
667 | 10773 | No Leaks Detected
667 | 10774 | No Leaks Detected
668 | 10775 | Picking List
668 | 10775 | Turbo Pump
668 | 10775 | Turbo Pump Controller
668 | 10775 | Chamber
668 | 10775 | Plinth
668 | 10775 | EO PCB
668 | 10775 | Stage PCB
668 | 10775 | Stage Logbook
668 | 10776 | All Cables Are Neatly Routed
668 | 10776 | All Plugs Are Secured Correctly
668 | 10776 | All Earth Cables Are Tightly Secured
668 | 10776 | Check Rear Panel Cables Are Correctly Located
668 | 10776 | Check Opto Devices Are Not Damaged On All PCBs & Mains Box
668 | 10781 | Chamber is Free From Defects
668 | 10781 | Chamber Meshes Are Fitted
668 | 10781 | Chamber Pumping Pipe Is Fitted Centrally To The Table Top
668 | 10781 | Table Top Is Free From Defects
668 | 10782 | Plinth Base Is Clean and Free From Debris
668 | 10783 | Clean & Tidy Assembly Bay
659 | 10683 | Damper Plate is Correct Version
659 | 10683 | Damper Plate is Located Correctly
661 | 10707 | Turbo Stack Assembled
661 | 10707 | Turbo M8 Bolts Torqued to 7Nm
663 | 10730 | Valves Cables
663 | 10730 | Heater Cables
663 | 10730 | Ion Pump Cable
663 | 10730 | Ion Pump PSU Serial Number
663 | 10730 | Fit Ion Pump PSU Bracket
663 | 10730 | Fit Ion Pump PSU
665 | 10734 | Stage Serial Number
665 | 10734 | Shim Stage Door
665 | 10734 | Set Stage Clamp Height
659 | 10685 | Earth Plate
659 | 10685 | Cable Plates
659 | 10685 | Bellows
659 | 10695 | Install PSU Loom
659 | 10695 | Specimen Current Test
659 | 11115 | Check SCS Folder/Sales Order
668 | 11116 | Check SCS Folder/Sales Order
659 | 11115 | Check SCS Folder/Sales Order
660 | 10706 | Cables Are Neatly Routed and Secured
660 | 10706 | Plinth Is free From Debris
660 | 10706 | Table Top Is Free From Damage
660 | 10706 | Table Top Is Secured
660 | 10706 | Clean and Tidy Assembly Bay
664 | 10732 | All Cables Are Neatly Routed
664 | 10732 | All Plugs Are Secured Correctly
664 | 10732 | All Earth Cables Are Tightly Secured
664 | 10732 | EHT Set Is Secured
659 | 11115 | Check SCS Folder/Sales Order
659 | 11115 | SCS Number (If applicable)
664 | 10733 | Clean and Tidy Assembly Bay
661 | 10716 | TIV Serial Number
661 | 10716 | Function Check
661 | 10716 | Fit TIV
665 | 11193 | Check that the Assembly is the Correct Part Number
665 | 11193 | Correct Collector Beak Fitted
665 | 11193 | Fit Detector Assembly
665 | 11193 | Apply Emersion Oil to PM Tube Face
665 | 11193 | Fit PM Head Amp Assembly
658 | 10680 | Serial Number
658 | 10681 | Cable Up PSU
659 | 11106 | Serial Number
665 | 11193 | Check that the Assembly is the Correct Part Number
665 | 11193 | Correct Collector Beak Fitted
665 | 11193 | Fit Detector Assembly
665 | 11193 | Apply Emersion Oil to PM Tube Face
665 | 11193 | Fit PM Head Amp Assembly
667 | 11247 | Print Plinth Seperators
667 | 11247 | Sort Documentation with Seperators
665 | 11267 | Detector Type
665 | 11267 | Detector Serial Number
665 | 11267 | Detector Type
665 | 11267 | Detector Serial Number
665 | 10741 | EP/Water Kit Fitted
665 | 10741 | TTL Kit Fitted
665 | 10742 | Check all valve cables have been labelled
665 | 10742 | All cable earth screens are in electrical contact with the grounding bracket
665 | 10741 | EP/Water Kit Fitted
665 | 10741 | TTL Kit Fitted
659 | 10683 | Damper Plate is Correct Version
659 | 10683 | Damper Plate is Located Correctly
659 | 10684 | Table Top is Correct Version
659 | 10684 | Table Top is Free from Damage
659 | 10684 | Remove Punchout (Dependant on system type)
666 | 10756 | Clean and tidy Assembly Bay
659 | 10685 | Earth Plate
659 | 10685 | Cable Plates
659 | 10685 | Bellows
659 | 10686 | Turbo Controller Serial Number
659 | 10689 | Vacuum PCB Serial Number
659 | 10690 | ASAP Serial Number
659 | 10691 | EO PCB Serial Number
376 | 3250 | Sales order is latest revision
376 | 3250 | All accessories are accounted for and are fitted to correct ports
376 | 3250 | Identify any un-described items on sales order
659 | 10692 | Stage PCB Serial Number
659 | 11398 | Fit Vac/ASAP PSU
659 | 11398 | Vac/ASAP PSU Serial Number
659 | 11398 | Connect Turbo Controller PCB Power (P2)
659 | 11398 | Fit EO/Stage PCB
659 | 11398 | EO/Stage PSU Serial Number
659 | 11399 | Fit Quiet Mode PCB
659 | 11399 | Fit Cable X45
659 | 11399 | Fit Cable X11
660 | 10701 | EHT Cable PL5 is Plugged Into PL5
660 | 10701 | All Connectors Are Securely Fitted
660 | 10701 | Filter Earth Link Fitted
660 | 10701 | Cables Are Neatly Routed And Secured
665 | 10735 | Fit Stage Cable
665 | 10735 | Fit Specimen Touch Cable
660 | 10702 | All Connections Are Secured with Screws
660 | 10702 | Cables Are Neatly Routed and Secured
660 | 10703 | All Connections Are Secured with Screws
660 | 10703 | Cables Are Neatly Routed and Secured
660 | 10704 | All Connections Are Secured with Screws
660 | 10704 | Cables Are Neatly Routed and Secured
660 | 10705 | All Connections Are Secured with Screws
660 | 10705 | Cables Are Neatly Routed and Secured
660 | 10706 | Cables Are Neatly Routed and Secured
660 | 10706 | Plinth Is free From Debris
660 | 10706 | Table Top Is Free From Damage
660 | 10706 | Table Top Is Secured
660 | 10706 | Clean and Tidy Assembly Bay
661 | 10710 | Torque Chamber Bolts to 35Nm
376 | 3251 | F/W Version
662 | 10722 | All Vacuum Clamps are Tight
662 | 10722 | All Valves Are Fitted Corrctly
662 | 10722 | All Valves Are correctly Connected
662 | 10723 | Check Correct Routing Of Air Pipes
662 | 10723 | Check Correct Orientation Of Air Valves
662 | 10724 | Check Torque Of Chamber Bolts
662 | 10724 | Chamber Free From Any Detects
662 | 10725 | Free From Any Defects
662 | 10725 | Table Top Secured
663 | 10728 | Scan Cables
663 | 10728 | Lens Cables
663 | 10728 | EHT Cables
663 | 10728 | Penning Cable
663 | 10730 | Valves Cables
663 | 10730 | Heater Cables
663 | 10730 | Ion Pump Cable
663 | 10730 | Ion Pump PSU Serial Number
663 | 10730 | Fit Ion Pump PSU Bracket
663 | 10730 | Fit Ion Pump PSU
663 | 10731 | EHT Serial Number
663 | 10731 | Torque Earth Connection to 3Nm
666 | 10754 | All Cables Are Neatly Routed
666 | 10754 | All Plugs Are Secured Correctly
666 | 10754 | All Earth Cables Are Tightly Secured
666 | 10755 | Fitted To Correct Ports
666 | 10755 | Detectors Connected To Correct ASAP Port And Labelled
664 | 10732 | All Cables Are Neatly Routed
664 | 10732 | All Plugs Are Secured Correctly
664 | 10732 | All Earth Cables Are Tightly Secured
664 | 10732 | EHT Set Is Secured
664 | 10733 | Clean and Tidy Assembly Bay
665 | 10734 | Stage Serial Number
665 | 10734 | Shim Stage Door
665 | 10734 | Set Stage Clamp Height
665 | 10735 | Fit Stage Cable
665 | 10735 | Fit Specimen Touch Cable
665 | 10741 | EP/Water Kit Fitted
665 | 10741 | TTL Kit Fitted
665 | 11390 | Fit Valves and Pipework
665 | 11390 | Connect Valve Cable to PL3 on Vac PCB
665 | 11405 | Select Number of Ways
285 | 3247 | Set Air Suspension Height
285 | 3247 | Check Correct Operation
667 | 10765 | HAL Test
667 | 10765 | PAT Test
667 | 10771 | Vac 5.15V (P1)
667 | 10771 | Vac 5.00V (P10)
667 | 10771 | Stage 5.15V
667 | 10771 | Stage 21.5V
667 | 10771 | Stage -21.5V
667 | 10771 | EO 5.15V
659 | 10683 | Damper Plate is Correct Version
659 | 10683 | Damper Plate is Located Correctly
659 | 10684 | Table Top is Correct Version
659 | 10684 | Table Top is Free from Damage
659 | 10684 | Remove Punchout (Dependant on system type)
285 | 3248 | Plate
285 | 3248 | Screws
285 | 3248 | Spacers
659 | 10685 | Earth Plate
659 | 10685 | Cable Plates
659 | 10685 | Bellows
659 | 10686 | Turbo Controller Serial Number
385 | 3283 | EXT Voltage
659 | 10689 | Vacuum PCB Serial Number
659 | 10690 | ASAP Serial Number
659 | 10691 | EO PCB Serial Number
659 | 10692 | Stage PCB Serial Number
661 | 10710 | Torque Chamber Bolts to 35Nm
385 | 3284 | Max specimen current
663 | 10728 | Scan Cables
663 | 10728 | Lens Cables
663 | 10728 | EHT Cables
663 | 10728 | Penning Cable
663 | 10730 | Valves Cables
663 | 10730 | Heater Cables
663 | 10730 | Ion Pump Cable
663 | 10730 | Ion Pump PSU Serial Number
663 | 10730 | Fit Ion Pump PSU Bracket
663 | 10730 | Fit Ion Pump PSU
663 | 10731 | EHT Serial Number
663 | 10731 | Torque Earth Connection to 3Nm
665 | 10734 | Stage Serial Number
665 | 10734 | Shim Stage Door
665 | 10734 | Set Stage Clamp Height
665 | 10741 | EP/Water Kit Fitted
665 | 10741 | TTL Kit Fitted
390 | 3318 | Pressure (<5e-9mbar)
667 | 10771 | Vac 5.15V (P1)
667 | 10771 | Vac 5.00V (P10)
667 | 10771 | ASAP 5.25V (Set on PSU)
667 | 10771 | Stage 5.15V
667 | 10771 | Stage 21.5V
667 | 10771 | Stage -21.5V
667 | 10771 | EO 5.15V
285 | 3242 | Select PC type
285 | 3242 | PC S/No.
285 | 3242 | Hard Drive Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Activate Windows
285 | 3242 | Record Windows Product Key
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Smart SEM Version
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3243 | Language
285 | 3245 | Check Cap is Fitted
285 | 3245 | Check Earth Cable is Fitted
285 | 3246 | Air Pressure O/P = 6bar
285 | 3246 | Water
285 | 3246 | Mains
381 | 3270 | Data Sheet Vext
381 | 3270 | Vext
381 | 3270 | Speciment Current
381 | 3271 | Data Sheet Vext
381 | 3271 | Vext
381 | 3271 | Speciment Current
385 | 3287 | 1kV X
385 | 3287 | 1kV Y
385 | 3287 | 10kV X
385 | 3287 | 10kV Y
385 | 3287 | 20kV X
385 | 3287 | 20kV Y
385 | 3288 | 1kV X
385 | 3288 | 1kV Y
385 | 3288 | 10kV X
385 | 3288 | 10kV Y
385 | 3288 | 20kV X
385 | 3288 | 20kV Y
386 | 3415 | Serial No. (or or N/A)
386 | 3415 | No X-Ray Leackage detected.
388 | 3291 | X =
388 | 3291 | Y =
388 | 3292 | 1kV (Spec: $[Resolution1kV]nm)
388 | 3292 | 15kV (Spec: $[Resolution15kV]nm)
388 | 3293 | Actual (Spec: $[ResolutionBSD]nm)
388 | 3294 | Actual (Spec: $[ResolutionVP]nm)
389 | 3296 | Function Test
389 | 3297 | BSE Collector Voltage Test
389 | 3297 | BSE Macro
389 | 3297 | BSE 0V Bias Image
389 | 3297 | BSE 620V Bias Image
389 | 3298 | Detector Alignment Check
389 | 3298 | Collector Bias Check
389 | 3298 | Function Test
389 | 3299 | Edit Chamberscope Macro
389 | 3299 | Set Chamberscope Focus
389 | 3300 | Function Test
389 | 3301 | Function Test
389 | 3302 | Run ‘Quick Sample Clean’
389 | 3303 | Function Test
389 | 3303 | Disable Confirmation Check
389 | 3303 | Turbo Load Check
389 | 3304 | BSD Diode Pre-test
389 | 3304 | Check all segments
389 | 3304 | Check all Gain Ranges
389 | 3304 | Take 3 X Proof Images
389 | 3304 | Calibrate ADC 4 Card
389 | 3305 | BSD Diode Pre-test
389 | 3305 | Check all Gain Ranges
389 | 3305 | Take 3 X Proof Images
389 | 3306 | Check Segments
389 | 3306 | Check all Gain Ranges
389 | 3306 | Proof Images
389 | 3307 | Set Rod Retract Switch
389 | 3307 | Edit Airlock Macro
389 | 3307 | Set Up Stage Exchange
389 | 3307 | Airlock Function Test
389 | 3308 | Gatan Function Test
389 | 3308 | Resolution
389 | 3308 | Ripple
389 | 3308 | Contamination
390 | 3320 | Set Rod Retract Switch
390 | 3320 | Set Rod Retract Switch
390 | 3320 | Set Rod Retract Switch
390 | 3320 | Set Rod Retract Switch
390 | 3321 | Load customer licences and check against order
390 | 3321 | Test customer licenses
390 | 3322 | Check System And Dongle S/No. Match In Service Panel
390 | 3322 | Un-tick Omit Dongle In Factory Panel
390 | 3326 | All verification & Safety tests done and within threshold limits defined
390 | 3326 | Sales Order is latest version
390 | 3326 | All items on Sales order accounted for
390 | 3327 | Em Server.accdb
390 | 3327 | Em Server.txt Log
390 | 3327 | Optiprobe.txt Log
390 | 3327 | All Images
391 | 3328 | Two detectors can be used simultaneously
391 | 3329 | Two Mags can be achieved simultaneously
391 | 3330 | Feature can be centred at high and low mag
391 | 3331 | Ensure mode works without distortions or rotation.
391 | 3332 | 5+ annotation
391 | 3333 | Able to edit LUT
391 | 3334 | Optiprobe can be enabled and specimen current varied
391 | 3335 | Application is installed and can be opened
391 | 3336 | Plasma Clean can be enabled
393 | 3344 | Images in TIFF format can be saved on disk and Hard disk and reloaded without loss of information
394 | 3349 | Reading (< 3.0e-6 mbar) =
394 | 3350 | Reading (< 5e-9mbar) =
394 | 3351 | Chamber vacuum can be changed continuously between 10 and 133 PA (Beam On)
396 | 3358 | Reading (220 +/- 10pA)
396 | 3359 | Reading (300 +/- 10pA)
396 | 3360
396 | 3360 | Filament Serial number
396 | 3360 | Filament heating Current
396 | 3360 | Extrator Power at 0.2 mA/sr
396 | 3360 | Total Current at 0.2 mA/sr
398 | 3380 | Stigmator value X (< 30%)
398 | 3380 | Stigmator value Y (< 30%)
399 | 3397 | Em Server.mdb
399 | 3397 | Em Server.txt Log
399 | 3397 | Optiprobe.txt Log
399 | 3397 | All Images
399 | 3398 | Uninstall MS Security Essentials
399 | 3398 | Remove test macro files
399 | 3398 | Empty Recycle bin
399 | 3398 | Remove images from image gallery
399 | 3400 | Em Server.mdb
399 | 3400 | Licence File
382 | 6928 | Macro Result ($[StageRepeatability])
382 | 6929 | X
382 | 6929 | Y
382 | 6931 | Drift
382 | 6932 | Spacer Height Setup
382 | 6932 | Holder Height
382 | 6932 | Holder Rotation
389 | 3304 | BSD Diode Pre-test
389 | 3304 | Check all segments
389 | 3304 | Check all Gain Ranges
389 | 3304 | Take 3 X Proof Images
389 | 3304 | Calibrate ADC 4 Card
397 | 10925 | Mineralogic Software Installed and Tested
12 | 164 | Replace the M8 nuts on the gun head ground with 2x Zinc equivalents
285 | 3242 | Select PC type
285 | 3242 | PC S/No.
285 | 3242 | Hard Drive Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Activate Windows
285 | 3242 | Record Windows Product Key
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3245 | Check Cap is Fitted
285 | 3245 | Check Earth Cable is Fitted
285 | 3246 | Air Pressure O/P = 6bar
285 | 3246 | Water
285 | 3246 | Mains
285 | 3247 | Set Air Suspension Height
285 | 3247 | Check Correct Operation
376 | 3250 | Sales order is latest revision
376 | 3250 | All accessories are accounted for and are fitted to correct ports
376 | 3250 | Identify any un-described items on sales order
376 | 3251 | F/W Version
376 | 3252 | F/W Version
376 | 3252 | Hardware Version
376 | 3253 | F/W Version
376 | 3253 | Hardware Version
376 | 3255 | Controller Version
376 | 3255 | Axes Version
376 | 3255 | Hardware Version
397 | 10926 | ZEN Software Installed
377 | 3261 | Bake Out Hours
381 | 3270 | Data Sheet Vext
381 | 3270 | Vext
381 | 3270 | Speciment Current
381 | 3271 | Data Sheet Vext
381 | 3271 | Vext
381 | 3271 | Speciment Current
382 | 6928 | Macro run time
382 | 6928 | Macro Result ($[StageRepeatability])
399 | 3396 | Special Fitted
399 | 3396 | If Fitted, Checked by SCS Team Member
393 | 6934 | PC S/N
393 | 6934 | Processor
393 | 6934 | RAM size
393 | 6934 | Hard Drive Size
393 | 6934 | Operating System
393 | 6934 | SmartSEM Version
393 | 6934 | Stage Firmware Version
393 | 6934 | Vacuum Firmware Version
393 | 6934 | EO Firmware Version
396 | 3361 | Measured Value
382 | 6929 | X
382 | 6929 | Y
376 | 3254 | F/W Version
376 | 3254 | Hardware Version
382 | 6931 | Drift
382 | 6932 | Spacer Height Setup
382 | 6932 | Holder Height
382 | 6932 | Holder Rotation
385 | 3284 | Max specimen current
385 | 3286 | Fil I
385 | 3286 | Ext V
385 | 3287 | 1kV X
385 | 3287 | 1kV Y
385 | 3287 | 10kV X
385 | 3287 | 10kV Y
385 | 3287 | 20kV X
385 | 3287 | 20kV Y
385 | 3288 | 1kV X
385 | 3288 | 1kV Y
385 | 3288 | 10kV X
385 | 3288 | 10kV Y
385 | 3288 | 20kV X
385 | 3288 | 20kV Y
386 | 3415 | Serial No. (or or N/A)
386 | 3415 | No X-Ray Leackage detected.
394 | 3349 | Reading (< 3.00e-06 mbar) =
388 | 3291 | X =
388 | 3291 | Y =
397 | 10919 | Stage Moves To Correct Position When Arm Is Opened
397 | 10919 | Image Is Captured
397 | 10919 | Functional Check
388 | 3292 | 1kV (Spec: $[Resolution1kV]nm)
388 | 3292 | 15kV (Spec: $[Resolution15kV]nm)
389 | 3297 | BSE Collector Voltage Test
389 | 3297 | BSE Macro
389 | 3297 | BSE 0V Bias Image
389 | 3297 | BSE 620V Bias Image
389 | 3298 | Detector Alignment Check
389 | 3298 | Collector Bias Check
389 | 3298 | Function Test
389 | 3299 | Edit Chamberscope Macro
389 | 3299 | Set Chamberscope Focus
389 | 3300 | Function Test
389 | 3301 | Function Test
389 | 3302 | Run ‘Quick Sample Clean’
389 | 3303 | Function Test
389 | 3303 | Disable Confirmation Check
389 | 3303 | Turbo Load Check
389 | 3304 | BSD Diode Pre-test
389 | 3304 | Check all segments
389 | 3304 | Check all Gain Ranges
389 | 3304 | Take 3 X Proof Images
389 | 3304 | Calibrate ADC 4 Card
389 | 3307 | Set Rod Retract Switch
389 | 3307 | Edit Airlock Macro
389 | 3307 | Set Up Stage Exchange
389 | 3307 | Airlock Function Test
394 | 3350 | Reading (< 5.00e-09mbar) =
397 | 10921 | Check Aquire Spectrum
397 | 10921 | Check Line Scanning
397 | 10921 | Check Mapping
397 | 10921 | Check EDX Icon Is Present In All User Toolbars
397 | 10921 | EDAX Calibration File Saved To SmartSEM USB Stick
389 | 3314 | Fit Detector to Chamber
389 | 3314 | Install EDS PCBs into PC
389 | 3314 | Install Cables
389 | 3314 | Install EDAX Software
389 | 3314 | Disable Fast Startup
389 | 3314 | Install Smart API Client from SmartSEM USB stick, setup, configure and Run API test
389 | 3314 | Configure imaging – Run SGShell
389 | 3314 | Calibrate Detector – Run eDPPShell
389 | 3314 | Save calibration data on exit
389 | 3314 | Run SmartEDX software – check Spectrum, Mapping and Line Scan
389 | 3314 | Save aquired Mapping and Spectrum Images to Image Folder
389 | 3314 | Save Calibration Data to SmartSem USB stick and in the Image folder.
389 | 3314
390 | 3327 | Em Server.accdb
390 | 3327 | Em Server.txt Log
390 | 3327 | Optiprobe.txt Log
390 | 3327 | Em Server Screening.txt
390 | 3327 | All Images
400 | 3403 | Fit LH Bracket
400 | 3403 | Fit RH Bracket
400 | 3403 | Fit Rear Bracket
400 | 3403 | Fit Front Bracket
401 | 10623 | Check LH Bracket
401 | 10623 | Check RH Bracket
401 | 10623 | Check Rear Bracket
401 | 10623 | Check Front Bracket
386 | 3414 | Accelerating Voltage=30KV
386 | 3414 | Specimen Current
386 | 3414 | Monitor Model
386 | 3414 | Serial Number
386 | 3414 | Cal. Expiry Date
386 | 3414 | No X-Ray Leaks Detected
386 | 3414 | Complete & Sign X-Ray Certificate
386 | 3415 | Serial No. (or N/A)
386 | 3415 | No X-Ray Leackage detected.
285 | 3242 | Select PC type
285 | 3242 | PC S/No.
285 | 3242 | Hard Drive Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Activate Windows
285 | 3242 | Record Windows Product Key
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Record MAC address
285 | 3242 | Network access requested
386 | 3414 | Accelerating Voltage=30KV
386 | 3414 | Specimen Current
386 | 3414 | Monitor Model
386 | 3414 | Serial Number
386 | 3414 | Cal. Expiry Date
386 | 3414 | No X-Ray Leaks Detected
386 | 3414 | Complete & Sign X-Ray Certificate
388 | 3293 | Actual (Spec: $[ResolutionBSD]nm)
396 | 3360
396 | 3360 | Filament Serial number
396 | 3360 | Filament heating Current
396 | 3360 | Extrator Power at 0.2 mA/sr
396 | 3360 | Total Current at 0.2 mA/sr
396 | 10865 | Filament heating Current
396 | 10865 | Extrator Power at 0.2 mA/sr
396 | 10865 | Total Current at 0.2 mA/sr
389 | 3304 | BSD Diode Pre-test
389 | 3304 | Check all segments
389 | 3304 | Check all Gain Ranges
389 | 3304 | Take 3 X Proof Images
389 | 3309 | Connect Silver Box and Cables
389 | 3309 | Check Altas Test USB Stick Has The latest Version Of Software
389 | 3309 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
389 | 3309 | Request Licence File from Fibics
389 | 3309 | Install Atlas Software
389 | 3309 | Install Licence File
389 | 3309 | Calibrate SEM
389 | 3309 | Obtain Mosaic
389 | 3309 | Set ATLAS Rotation Offset
389 | 3309 | Save ATLAS Licence Document
389 | 3316 | Install SmartSEM Licences
389 | 3316 | Install MIL Redistribution
389 | 3316 | Fit MIL Dongle
389 | 3316 | Install SmartPI Software
389 | 3316 | Calibrate I/O-SCAN card
389 | 3316 | Setup & Run Application
389 | 3316 | Check Results & Image Montage
389 | 10867 | Install Hard Disk
389 | 10867 | Format Hard Disk (NTFS)
389 | 10867 | Hard Disk Is Seen In File Explorer
389 | 10867 | Copy A File To The Hard Disk, Then Delete It
389 | 10867 | Install Cloning Software
389 | 10870 | Camera Model
389 | 10870 | Install Camera Software
389 | 10872 | Install EMO Box
389 | 10872 | Check All Mains Outputs
389 | 10872 | Check Cut Off Switch
389 | 10873 | Number of Detectors
389 | 10873 | Detector Serial Number(s)
389 | 10873 | Detector Resolution
389 | 10873 | Detector Area
389 | 10873 | Detector Installed
389 | 10873 | In/Out Stops Set
389 | 10873 | Isolation Check of Detector & Column >2MΩ
389 | 10873 | Install Remoting & Setup Network Connection
389 | 10873 | Install EDS Software
389 | 10873 | Install EDS Manager
389 | 10873 | Install EDS Licences
389 | 10873 | Setup Pulse Processor
389 | 10873 | Check Timing Cards
389 | 10874 | Install SmartSEM Licences
389 | 10874 | Instal MIL Redistribution
389 | 10874 | Fit MIL Dongle
389 | 10874 | Install Mineralogic Software
389 | 10874 | Calibrate I/O-SCAN card
389 | 10874 | Setup & Run Application
389 | 10874 | Check Results & Image Montage
389 | 10874 | 24 Hour Run
390 | 3317 | Pressure (<3.00e-9mbar)
390 | 3318 | Pressure (<5.00e-9mbar)
390 | 3320 | 1KV Resolution Image
390 | 3320 | 15KV Resolution Image
390 | 3320 | Licence File
390 | 3320 | EM Server ACCDB
390 | 3320 | EM Server Screening .TXT
393 | 6934 | PC S/N
393 | 6934 | Processor
393 | 6934 | RAM size
393 | 6934 | Hard Drive Size
393 | 6934 | Operating System
393 | 6934 | SmartSEM Version
393 | 6934 | Stage Firmware Version
393 | 6934 | Vacuum Firmware Version
393 | 6934 | EO Firmware Version
393 | 6934 | PC S/N
393 | 6934 | Processor
393 | 6934 | RAM size
393 | 6934 | Hard Drive 1 Size
393 | 6934 | Hard Drive 2 Size
393 | 6934 | Operating System
393 | 6934 | SmartSEM Version
393 | 6934 | Stage Firmware Version
393 | 6934 | Vacuum Firmware Version
393 | 6934 | EO Firmware Version
397 | 10918 | Atlas Hardware Installed
397 | 10918 | Atlas Software Installed
397 | 10918 | Atlas Licence Saved To SmartSEM USB Stick
389 | 10879 | SmartTouch Software Version
389 | 10879 | Install SmartTouch Software
389 | 10879 | Install Easy VP Aperture
389 | 10879 | Capture Sample Image Using the Nav Cam
389 | 10879 | Select Sample Holder and Sample Type
389 | 10879 | Set Yellow Line To 8.5mm WD
389 | 10879 | Sample Setup Complete
389 | 10879 | In the Image Tab, Check Sliders Function
389 | 10879 | Save Image - Photomode 1, Scan speed 8
389 | 10879 | Save Image - Photomode 2, Single Image
389 | 10879 | Save Image - Photomode 3, Scan Speed 2, Multiple Frames
389 | 10879 | Check Auto Saturation (W Filament Only)
389 | 10879 | Check Auto Stigmation
389 | 10879 | Check Manual Stigmation
389 | 10879 | Check Annotate
389 | 10879 | Check All Detectors Are Presaent
389 | 10879 | Select the Automate tab & Create Project as Zeiss
389 | 10879 | Setup 3 Point Calibration (Nav Cam )
389 | 10879 | Automation Run Completed
389 | 10879 | Select Browse Tab and Check Functions
389 | 10883 | ZEN Version
389 | 10883 | Correct Version of Windows 10 Installed
389 | 10883 | Install ZEN
389 | 10883 | Install Licence
389 | 10883 | Install Remote API Client
389 | 10883 | Functional Test
285 | 3242 | Select PC type
285 | 3242 | PC S/No.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Activate Windows
285 | 3242 | Record Windows Product Key
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Record MAC address
285 | 3242 | Network access requested
378 | 3265 | Serial No.
386 | 3414 | Accelerating Voltage=30KV
386 | 3414 | Specimen Current
386 | 3414 | Monitor Model
386 | 3414 | Serial Number
386 | 3414 | Cal. Expiry Date
386 | 3414 | No X-Ray Leaks Detected
389 | 3305 | BSD Diode Pre-test
389 | 3305 | Check all Gain Ranges
389 | 3305 | Take 3 X Proof Images
389 | 3306 | Check Segments
389 | 3306 | Check all Gain Ranges
389 | 3306 | Proof Images
389 | 11012 | Fit Test tool BB Electronis
389 | 11012 | Function Test
389 | 11012 | Remove Test tool BB Electronics
390 | 3317 | Pressure (<3.00e-6mbar)
390 | 3318 | Pressure (<5.00e-9 mbar)
390 | 3319 | Time (<5 min)
390 | 3320 | 1KV Resolution Image
390 | 3320 | 15KV Resolution Image
390 | 3320 | Licence File
390 | 3320 | EM Server ACCDB
390 | 3320 | EM Server Screening .TXT
390 | 3321 | Load customer licences and check against order
390 | 3321 | Test customer licenses
390 | 3327 | Em Server.accdb
390 | 3327 | Em Server.txt Log
390 | 3327 | Optiprobe.txt Log
390 | 3327 | Em Server Screening.txt
390 | 3327 | All Images
396 | 10885 | Check That all detectors are listed under and signed off in XRay checks, If any extra detectors have been fitted perform XRAY checks on those detectors
396 | 10885 | Sign XRay Certificate
397 | 3374 | Video Input Kit Fitted
397 | 10886 | Functional Check
397 | 10908 | Power Off Button Tested
397 | 10909 | Software Loaded
397 | 10910 | Hard Disk Is Seen In File Manager
397 | 10910 | Cloning Software Installed
397 | 10911 | C2D Is Selected When Going Into VP Mode
397 | 10911 | Image Brightness Changes When Changing Bias
397 | 10911 | Breakdown Of Bias Occurs at 100%
397 | 10912 | Check correct operation of Quiet Mode. Deselect Quiet Mode and ensure rotary pump starts
397 | 10913 | Run Up & Obtain An Image
399 | 3396 | Special Fitted
399 | 3396 | If Fitted, Checked by SCS Team Member
399 | 3398 | Uninstall MS Security Essentials
399 | 3398 | Remove test macro files
399 | 3398 | Empty Recycle bin
399 | 3398 | Remove images from image gallery
399 | 3400 | Em Server.mdb
399 | 3400 | Licence File
399 | 3402 | Set Customer Language
400 | 3403 | Fit LH Bracket
400 | 3403 | Fit RH Bracket
400 | 3403 | Fit Rear Bracket
400 | 3403 | Fit Front Bracket
389 | 3300 | Function Test
389 | 3301 | Function Test
389 | 3303 | Function Test
389 | 3303 | Disable Confirmation Check
389 | 3303 | Turbo Load Check
382 | 11008 | Request 5200
382 | 11008 | Request 5240
389 | 3314 | Fit Detector to Chamber
389 | 3314 | Install EDS PCBs into PC
389 | 3314 | Install Cables
389 | 3314 | Install EDAX Software
389 | 3314 | Disable Fast Startup
389 | 3314 | Install Smart API Client from SmartSEM USB stick, setup, configure and Run API test
389 | 3314 | Configure imaging – Run SGShell
389 | 3314 | Calibrate Detector – Run eDPPShell
389 | 3314 | Save calibration data on exit
389 | 3314 | Run SmartEDX software – check Spectrum, Mapping and Line Scan
389 | 3314 | Save aquired Mapping and Spectrum Images to Image Folder
389 | 3314 | Save Calibration Data to SmartSem USB stick and in the Image folder.
389 | 3314
389 | 10927 | Install YAG BSD Lightpipe
389 | 10927 | Alignment
389 | 10927 | Select YAG BSD in the Administrator
389 | 10927 | 30KV Function Image
389 | 10927 | Obtain a Image at 3KV
400 | 3403 | Fit LH Bracket
400 | 3403 | Fit RH Bracket
400 | 3403 | Fit Rear Bracket
400 | 3403 | Fit Front Bracket
400 | 3403 | Checked By Name:
384 | 3280 | Record Figures in Gun-Aperture Alignment Table in customer Documentation
384 | 3281 | Record Figures in Gun-Aperture Alignment Table in customer Documentation
386 | 3415 | Serial No. (or N/A)
386 | 3415 | No X-Ray Leakage detected.
387 | 11010 | Ensure EHT Earth Connection is Torqued to 3Nm
389 | 11011 | Function Test
285 | 3242 | Select PC type
285 | 3242 | PC S/No.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Activate Windows
285 | 3242 | Record Windows Product Key
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Record MAC address
285 | 3242 | Network access requested
| 11357 | C1 Coil
| 11357 | C2 Coil
389 | 3298 | Detector Alignment Check
389 | 3298 | Collector Bias Check
389 | 3298 | Function Test
389 | 3299 | Edit Chamberscope Macro
389 | 3299 | Set Chamberscope Focus
384 | 11044 | Set Aperture sizes in alignment panel
384 | 11044 | Set Aperture default Gun and Aperture alignment figures
384 | 11044 | Re-align aperture 1
397 | 3369 | Focused chamber-scope image can be obtained at predefined contrast and brightness
397 | 3370 | Focused chamber-scope image can be obtained at predefined contrast and brightness
399 | 3398 | Uninstall MS Security Essentials
399 | 3398 | Remove test macro files
399 | 3398 | Delete Contents of Log Folder
399 | 3398 | Download Folder Images Deleted
399 | 3398 | Remove images from image gallery
399 | 3398 | Empty Recycle Bin
389 | 11074 | Check each individual segment in normal and inverse modes
389 | 11074 | Analyze Segments
389 | 11074 | Gain Ranges
389 | 11074 | Pneumatics
389 | 11074 | Alignment
389 | 11074 | Vac Degradation < 3x10-5 mbar
389 | 11074 | CZ Detector Cal
389 | 11074 | Z Adjustment
399 | 3398 | Uninstall MS Security Essentials
399 | 3398 | Remove test macro files
399 | 3398 | Delete Contents of Log Folder
399 | 3398 | Download Folder Images Deleted
399 | 3398 | Remove images from image gallery
399 | 3398 | Empty Recycle Bin
399 | 3398 | Clear Internet History
397 | 11075 | Pneumatics functional
397 | 11075 | SmartSEM Control tested
397 | 11075 | Images saved
389 | 11074 | Check each individual segment in normal and inverse modes
389 | 11074 | Analyze Segments
389 | 11074 | Gain Ranges
389 | 11074 | Pneumatics
389 | 11074 | Alignment
389 | 11074 | Vac Degradation < 3x10-5 mbar
389 | 11074 | CZ Detector Cal
389 | 11074 | Z Adjustment
397 | 11075 | Pneumatics functional
397 | 11075 | SmartSEM Control tested
397 | 11075 | Images saved
389 | 11082 | Functional Check
389 | 11082 | Function Image
394 | 3351 | Chamber vacuum can be changed continuously between 10 and 133 PA (Beam On)
394 | 3351 | Check Function of VP Detector (VPSE or C2D)
399 | 3396 | Special Fitted
399 | 3396 | If Fitted, Checked by SCS Team Member
285 | 11101 | Check SCS Folder/Sales Order
390 | 11102 | Check SCS Folder/Sales Order
285 | 11101 | Check SCS Folder/Sales Order
390 | 11102 | Check SCS Folder/Sales Order
399 | 3396 | Special Fitted
399 | 3396 | If Fitted, Checked by SCS Team Member
285 | 3242 | Select PC type
285 | 3242 | PC S/No.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Record MAC address
285 | 3242 | Network access requested
285 | 11101 | Check SCS Folder/Sales Order
285 | 11101 | SCS Number (If applicable)
397 | 11075 | Pneumatics functional
397 | 11075 | SmartSEM Control tested
397 | 11075 | Images saved
285 | 3242 | Select PC type
285 | 3242 | PC S/No.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Record MAC address
285 | 3242 | Network access requested
389 | 11082 | Functional Check
389 | 11082 | Function Image
285 | 3247 | Set Air Suspension Height
285 | 3247 | Check Correct Operation
285 | 11214 | Install EMO Box
285 | 11214 | Check All Mains Outputs
285 | 11214 | Check Cut Off Switch
389 | 3308 | Gatan Function Test
389 | 3308 | Resolution
389 | 3308 | Ripple
389 | 3308 | Contamination
390 | 11102 | Check SCS Folder/Sales Order
399 | 3396 | Special Fitted
399 | 3396 | If Fitted, Checked by SCS Team Member
389 | 3304 | BSD Diode Pre-test
389 | 3304 | Check each individual segment in normal and inverse modes
389 | 3304 | Analyze Segments
389 | 3304 | Gain Ranges
389 | 3305 | BSD Diode Pre-test
389 | 3305 | Check each individual segment in normal and inverse modes
389 | 3305 | Analyze Segments
389 | 3305 | Gain Ranges
389 | 11074 | Check each individual segment in normal and inverse modes
389 | 11074 | Analyze Segments
389 | 11074 | Gain Ranges
389 | 11074 | Pneumatics
389 | 11074 | Alignment
389 | 11074 | Vac Degradation < 3x10-5 mbar
389 | 11074 | CZ Detector Cal
389 | 11074 | Z Adjustment
285 | 11214 | Install EMO Box
285 | 11214 | Check All Mains Outputs
285 | 11214 | Check Cut Off Switch
389 | 3316 | Print Out SmartPI Check List (Link in Documents)
389 | 3316 | Install SmartSEM Licences
389 | 3316 | Install MIL Redistribution
389 | 3316 | Fit MIL Dongle
389 | 3316 | Install SmartPI Software
389 | 3316 | Calibrate I/O-SCAN card
389 | 3316 | Setup & Run Application
389 | 3316 | Check Results & Image Montage
389 | 10867 | Install Hard Disk
389 | 10867 | Format Hard Disk (NTFS)
389 | 10867 | Hard Disk Is Seen In File Explorer
389 | 10867 | Copy A File To The Hard Disk, Then Delete It
389 | 10867 | Install Cloning Software
389 | 10874 | Print Out Mineralogic Check List (Link in Documents)
389 | 10874 | Install SmartSEM Licences
389 | 10874 | Instal MIL Redistribution
389 | 10874 | Fit MIL Dongle
389 | 10874 | Install Mineralogic Software
389 | 10874 | Calibrate I/O-SCAN card
389 | 10874 | Setup & Run Application
389 | 10874 | Check Results & Image Montage
389 | 10874 | 24 Hour Run
389 | 3316 | Correct PC - 347823-9024-660 Installed
389 | 3316 | Print Out SmartPI Check List (Link in Documents)
389 | 3316 | Install SmartSEM Licences
389 | 3316 | Install MIL Redistribution
389 | 3316 | Fit MIL Dongle
389 | 3316 | Install SmartPI Software
389 | 3316 | Calibrate I/O-SCAN card
389 | 3316 | Setup & Run Application
389 | 3316 | Check Results & Image Montage
389 | 10874 | Print Out Mineralogic Check List (Link in Documents)
389 | 10874 | Correct PC - Installed
389 | 10874 | Install SmartSEM Licences
389 | 10874 | Instal MIL Redistribution
389 | 10874 | Fit MIL Dongle
389 | 10874 | Install Mineralogic Software
389 | 10874 | Calibrate I/O-SCAN card
389 | 10874 | Setup & Run Application
389 | 10874 | Check Results & Image Montage
389 | 10874 | 24 Hour Run
385 | 11259 | Fil I
385 | 11259 | Ext V
385 | 11259 | Fil I
385 | 11259 | Ext V
397 | 10919 | Stage Moves To Correct Position When Arm Is Opened
397 | 10919 | Image Is Captured
397 | 10919 | Functional Check
383 | 3278 | Black Level Cal
383 | 3278 | C2D Bias Target Check
383 | 3278 | Gain Range Check
383 | 3278 | Auto Level Check
383 | 3278 | Auto Range Check
383 | 3278 | Function Image
285 | 11271 | Predictive Service S/W installed on Customer PC.
522 | 10953 | Fit an Aperture Insert Into Each Pole Piece
522 | 10953 | Check Aperture Inserts are Securely Fitted
389 | 3314 | Type of Detector
389 | 3314 | Serial number
389 | 3314 | Fit Flange & Mount Detector
389 | 3314 | Install EDS PCBs into PC
389 | 3314 | Install Cables
389 | 3314 | Install EDAX Software
389 | 3314 | Disable Fast Startup
389 | 3314 | Install Smart API Client from SmartSEM USB stick, setup, configure and Run API test
389 | 3314 | Configure imaging – Run SGShell
389 | 3314 | Calibrate Detector – Run eDPPShell
389 | 3314 | Save Calibration Report
389 | 3314 | Detector Resolution - ≥ 129 eV
389 | 3314 | Save calibration data on exit
389 | 3314 | Run SmartEDX software – check Spectrum, Mapping and Line Scan
389 | 3314 | Save aquired Mapping and Spectrum Images to Image Folder
389 | 3314 | Save Calibration Data to SmartSem USB stick and in the Image folder.
389 | 3314
389 | 3314 | Type of Detector
389 | 3314 | Detector Serial number
389 | 3314 | DDP Module Serial Number
389 | 3314 | Fit Flange & Mount Detector
389 | 3314 | Install EDS PCBs into PC
389 | 3314 | Install Cables
389 | 3314 | Install EDAX Software
389 | 3314 | Disable Fast Startup
389 | 3314 | Install Smart API Client from SmartSEM USB stick, setup, configure and Run API test
389 | 3314 | Configure imaging – Run SGShell
389 | 3314 | Calibrate Detector – Run eDPPShell
389 | 3314 | Save Calibration Report
389 | 3314 | Detector Resolution - ≥ 129 eV
389 | 3314 | Save calibration data on exit
389 | 3314 | Run SmartEDX software – check Spectrum, Mapping and Line Scan
389 | 3314 | Save aquired Mapping and Spectrum Images to Image Folder
389 | 3314 | Save Calibration Data to SmartSem USB stick and in the Image folder.
389 | 3314
389 | 3314 | Type of Detector
389 | 3314 | Detector Serial number
389 | 3314 | DDP Module Serial Number
389 | 3314 | Fit Flange & Mount Detector
389 | 3314 | Install EDS PCBs into PC
389 | 3314 | Installation of DPP Module & Cables
389 | 3314 | Configue PC Network & Test
389 | 3314 | Disable Windows10 Fast Startup
389 | 3314 | Install Licences
389 | 3314 | Fit Dongle
389 | 3314 | APEX Software Version
389 | 3314 | Install APEX Software
389 | 3314 | Install Remoting
389 | 3314 | Setup RConfigure
389 | 3314 | Run API Test
389 | 3314 | Configure imaging – Run SGShell
389 | 3314 | Calibrate Detector – Run eDPPShell
389 | 3314 | Save Calibration Report
389 | 3314 | Detector Resolution - ≥ 129 eV
389 | 3314 | Save calibration data on exit
389 | 3314 | Run APEX
389 | 3314 | Check Spectrum
389 | 3314 | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected
389 | 3314 | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected
389 | 3314 | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected
389 | 3314 | Check Mapping
389 | 3314 | Check Line Scanning
389 | 3314 | Save Aquired Spectrum, Mapping and Line Scanning Images to Image Folder
389 | 3314 | Save Calibration Data to the Licence Folderon the SmartSem USB stick.
389 | 3314 | Save Calibration Data to the Image Folder
389 | 3314 | Type of Detector
389 | 3314 | Detector Serial number
389 | 3314 | DDP Module Serial Number
389 | 3314 | Fit Flange & Mount Detector
389 | 3314 | Install EDS PCBs into PC
389 | 3314 | Installation of DPP Module & Cables
389 | 3314 | Configure PC Network & Test
389 | 3314 | Disable Windows10 Fast Startup
389 | 3314 | Install Licences
389 | 3314 | Fit Dongle
389 | 3314 | APEX Software Version
389 | 3314 | Install APEX Software
389 | 3314 | Install Remoting
389 | 3314 | Setup RConfigure
389 | 3314 | Run API Test
389 | 3314 | Configure imaging – Run SGShell
389 | 3314 | Calibrate Detector – Run eDPPShell
389 | 3314 | Save Calibration Report
389 | 3314 | Detector Resolution - ≥ 129 eV
389 | 3314 | Save calibration data on exit
389 | 3314 | Run APEX
389 | 3314 | Check Spectrum
389 | 3314 | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected
389 | 3314 | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected
389 | 3314 | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected
389 | 3314 | Check Mapping
389 | 3314 | Check Line Scanning
389 | 3314 | Save Aquired Spectrum, Mapping and Line Scanning Images to Image Folder
389 | 3314 | Save Calibration Data to the Licence Folderon the SmartSem USB stick.
389 | 3314 | Save Calibration Data to the Image Folder
698 | 10957 | Clean and Prepare the Jig and Fit the O ring
285 | 3242 | Select PC type
285 | 3242 | PC S/No.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Record MAC address
285 | 3242 | Network access requested
389 | 10873 | Number of Detectors
389 | 10873 | Detector Serial Number(s)
389 | 10873 | Detector Resolution
389 | 10873 | Detector Area
389 | 10873 | Detector Installed
389 | 10873 | In/Out Stops Set
389 | 10873 | Isolation Check of Detector & Column >2MΩ
389 | 10873 | Install Remoting & Setup Network Connection
389 | 10873 | Install EDS Software
389 | 10873 | Install EDS Manager
389 | 10873 | Install Bruker Licences
389 | 10873 | Setup Pulse Processor
389 | 10873 | Check Timing Cards
389 | 10873 | Install Detector Data
389 | 10873 | Check Microscope driver
389 | 10873 | Setup SEM Conditions
389 | 10873 | Determine Analytical Working Distance
389 | 10873 | Calibrate Detector
389 | 10873 | Calibrate Image in Espirit
389 | 11011 | Function Test
389 | 11012 | Fit Test tool BB Electronis
389 | 11012 | Function Test
389 | 11012 | Remove Test tool BB Electronics
389 | 11074 | Check each individual segment in normal and inverse modes
389 | 11074 | Analyze Segments
389 | 11074 | Gain Ranges
389 | 11074 | Pneumatics
389 | 11074 | Alignment
389 | 11074 | Vac Degradation < 3x10-5 mbar
389 | 11074 | CZ Detector Cal
389 | 11074 | Z Adjustment
389 | 3299 | Edit Chamberscope Macro
389 | 3299 | Set Chamberscope Focus
389 | 3299 | Chamberscope 2 - Select Port or Not Fitted
389 | 3299 | Chamberscope 2 - Check Camera 2 for Focus
389 | 3299 | Chamberscope 2 - Functional Check-Black/White + Colour Mode
389 | 3299 | Chamberscope 2 - Run wxPropview to Remove Defective Pixels
389 | 3299 | Chamberscope 2 - Add Icon to Toolbar
698 | 10968 | Fit C3 Cone into Jig (Vac seal facing you)
388 | 3292 | 1kV (Spec: $[Resolution1kV]nm)
388 | 3292 | 15kV (Spec: $[Resolution15kV]nm)
389 | 3307 | Set Rod Retract Switch
389 | 3307 | Edit Airlock Macro
389 | 3307 | Set Up Stage Exchange
389 | 3307 | Airlock Function Test
522 | 10954 | Inspect Bores of Plate and Clamp
522 | 10954 | Fit 300um Aperture
522 | 10954 | Clean and Inspect O Ring and Fit to the Underside of Plate
389 | 3309 | Connect Silver Box and Cables
389 | 3309 | Check Altas Test USB Stick Has The latest Version Of Software
389 | 3309 | Run Atlas5UpgradeAuditTool.exe and Save Computer Info
389 | 3309 | Request Licence File from Fibics
389 | 3309 | Install Atlas Software
389 | 3309 | Install Licence File
389 | 3309 | Calibrate SEM
389 | 3309 | Obtain Mosaic
389 | 3309 | Set ATLAS Rotation Offset
389 | 3309 | Save ATLAS Licence Document
389 | 3316 | Correct PC - 347823-9024-660 Installed
389 | 3316 | Print Out SmartPI Check List (Link in Documents)
389 | 3316 | Install SmartSEM Licences
389 | 3316 | Install MIL Redistribution
389 | 3316 | Fit MIL Dongle
389 | 3316 | Install SmartPI Software
389 | 3316 | Calibrate I/O-SCAN card
389 | 3316 | Setup & Run Application
389 | 3316 | Check Results & Image Montage
391 | 3333 | Able to edit LUT
397 | 10919 | Stage Moves To Correct Position When Arm Is Opened
397 | 10919 | Image Is Captured
397 | 10919 | Functional Check
397 | 11075 | Pneumatics functional
397 | 11075 | SmartSEM Control tested
397 | 11075 | Images saved
399 | 3396 | Special Fitted
399 | 3396 | If Fitted, Checked by SCS Team Member
389 | 11347 | Unpack Gatan + check against packing list (Photograph each layer before removing)
389 | 11347 | Set up stage disable Macro in SmartSEM + assign icon on toolbar
389 | 11347 | Remove transit clamps from Gatan stage and fit tie rods
389 | 11347 | Adjust stage rails so both Klein and Gatan stages can pump without assistance
389 | 11347 | If Nav Cam fitted please contact Peter Lewis to have in – line connection made up
| 11379 | Continuity (Close Circuit)
| 11379 | Coils to GND (Open Circuit)
389 | 11348 | Plasma clean Gatan stage (20W 2Hrs)
389 | 11348 | Fit On Point detector and diode
389 | 11348 | Remove Mc Affee software from Gatan PC
389 | 11348 | Disable Firewalls on Gatan and Zeiss PC
389 | 11348 | Install and configure remote client on Gatan PC
389 | 11348 | Configure ‘R configure‘ server on Zeiss PC
389 | 11348 | Ping IP adresses
389 | 11348 | Configure Gatan settings
389 | 11349 | Align BSD and obtain image
389 | 11349 | Check SCM with Gatan Stage
389 | 11349 | Ripple X Spec 3nm (600Kx Mag 15KV WD2)
389 | 11349 | Ripple Y Spec 3nm (600Kx Mag 15KV WD2)
389 | 11349 | Resolution 1KV Spec 2.1nm
389 | 11349 | Resolution 15KV Spec 1.2nm
389 | 11349 | Contamination
389 | 11349 | Pump Down Time
389 | 11349 | Ultimate Vacuum
389 | 11349 | Focal CC fitted to 3View stage
389 | 11349 | Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar
389 | 11349 | Charging paper test carried out and images saved.
389 | 11350 | Assemble Gatan vessel pumping accessories.
389 | 11350 | Re-package system
389 | 11350 | ManEng/ Senior Eng/ Production Manger Sign off
389 | 3297 | BSE Collector Voltage Test
389 | 3297 | BSE Macro
389 | 3297 | BSE 0V Bias Image
389 | 3297 | BSE 620V Bias Image
382 | 6928 | Macro run time
382 | 6928 | Macro Result ($[StageRepeatability])
382 | 6929 | X
382 | 6929 | Y
382 | 6931 | Drift
382 | 6932 | Spacer Height Setup
382 | 6932 | Holder Height
382 | 6932 | Holder Rotation
399 | 3400 | Em Server.accdb
399 | 3400 | Licence File
522 | 10949 | Clean and Inspect O Ring
522 | 10949 | Fit O Ring to the Underside of the Body
522 | 10949 | Fit Pumping Guide
522 | 10949 | Check Pumping Guide Alignment
522 | 10949 | Fit M3x10 Countersunk Screws
522 | 10949 | Fit Chamber Pumping Shield
522 | 10949 | Clean and Inspect O Ring
522 | 10949 | Fit O Ring to Upperside of the Body
285 | 3242 | Select PC type
285 | 3242 | PC S/No.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Upgrade ASAP Firmware and API
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Record MAC address
285 | 3242 | Network access requested
389 | 11011 | Function Test
388 | 3291 | X =
388 | 3291 | Y =
388 | 3292 | 1kV (Spec: $[Resolution1kV]nm)
388 | 3292 | 15kV (Spec: $[Resolution15kV]nm)
389 | 10883 | ZEN Version
389 | 10883 | Correct Version of Windows 10 Installed
389 | 10883 | Install ZEN
389 | 10883 | Install Licence
389 | 10883 | Install Remote API Client
389 | 10883 | Functional Test
389 | 3316 | Correct PC - 347823-9024-660 Installed
389 | 3316 | Print Out SmartPI Check List (Link in Documents)
389 | 3316 | Install SmartSEM Licences
389 | 3316 | Install MIL Redistribution
389 | 3316 | Fit MIL Dongle
389 | 3316 | SmartPI Software Version
389 | 3316 | Install SmartPI Software
389 | 3316 | Calibrate I/O-SCAN card
389 | 3316 | Setup & Run Application
389 | 3316 | Check Results & Image Montage
522 | 10950 | Inspect Aperture Blade
522 | 10950 | Install the Apertures
522 | 10951 | Disassemble the Firing Unit
522 | 10951 | Clean Listed Parts in Ultrasonic Tank
522 | 10951 | Reassemble Firing Unit and Fit Filament
522 | 10951 | Check Filament Height and Centre
522 | 10952 | Inspect Bores of Plate and Clamp
522 | 10952 | Fit 600um Aperture
522 | 10952 | Clean and Inspect O Ring and Fit to the Underside of Plate
389 | 10873 | Number of Detectors
389 | 10873 | Detector Serial Number(s)
389 | 10873 | Detector Resolution
389 | 10873 | Detector Area
389 | 10873 | Detector Installed
389 | 10873 | In/Out Stops Set
389 | 10873 | Isolation Check of Detector & Column >2MΩ
389 | 10873 | Install Remoting & Setup Network Connection
389 | 10873 | Bruker Esprit Software Version
389 | 10873 | Install Bruker Esprit Software
389 | 10873 | Bruker EDS Manager Version
389 | 10873 | Install EDS Manager
389 | 10873 | Install Bruker Licences
389 | 10873 | Setup Pulse Processor
389 | 10873 | Check Timing Cards
389 | 10873 | Install Detector Data
389 | 10873 | Check Microscope driver
389 | 10873 | Setup SEM Conditions
389 | 10873 | Determine Analytical Working Distance
389 | 10873 | Calibrate Detector
389 | 10873 | Calibrate Image in Espirit
388 | 11425 | 15kv 0.9nm
388 | 11425 | 15KV result
388 | 11425 | 1kv 1.3nm
388 | 11425 | 1KV result
388 | 11425 | 500V 2nm
388 | 11425 | 500V result
388 | 3292 | 1kV (Spec: $[Resolution1kV]nm)
388 | 3292 | 15kV (Spec: $[Resolution15kV]nm)
| 12818 | Datei =
40 | 121 | Collector Working
40 | 121 | Scintillator Working
388 | 3294 | Actual (Spec: $[ResolutionVP]nm)
388 | 3294 | Store VPSE/C2D Resolution Image and Pate image
522 | 10942 | Clean and Inspect O Ring
522 | 10942 | Fit O Ring into Blank
522 | 10943 | Clean and Inspect O Ring
522 | 10943 | Fit O Ring into Blank
522 | 10944 | Clean and Inspect O Rings
522 | 10944 | Fit O Rings into Plate
522 | 10945 | Clean and Inspect O Ring
522 | 10945 | Fit O Ring into Vac Seal
522 | 10945 | Insert Vac Seal into C3 Cone
522 | 10946 | Clean and Inspect O Ring
522 | 10946 | Fit O Ring it to the Lens Cap
522 | 10946 | Mount the End Cap Onto the C3 Cone
522 | 10946 | Tighten the Screws
522 | 10946 | End Cap and C3 Cone Gap Checked
522 | 10947 | Liner Tube Inspected
522 | 10947 | Inspect O Ring Sealing Surfaces
522 | 10947 | Liner Tube Cleaned
522 | 10947 | Clean and Inspect O Rings
522 | 10947 | Fit O Rings onto Coils
522 | 10948 | Liner Tube Inspected
522 | 10948 | Inspect O Ring Sealing Surfaces
522 | 10948 | Liner Tube Cleaned
522 | 10948 | Clean and Inspect O Rings
522 | 10948 | Fit O Rings onto Coils
522 | 10979 | Select System Type
522 | 10979 | Column Type
388 | 11426 | 15kv 0.7nm
388 | 11426 | 15KV result
388 | 11426 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 1kv 1.2nm
388 | 11426 | 1KV result
388 | 11426 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 500V 1.9nm
388 | 11426 | 500V result
388 | 11426 | Store 500V Resolution Images and Pate image for 2 x consecutive images
698 | 10969 | Fit the Sealing Plate onto the Inner C3 Cone
698 | 10969 | Fit Handles and Lower into the Outer Cone
698 | 10969 | Secure the Outer C3 Cone to the Jig
698 | 10969 | Secure the Inner C£ Cone to the Outer Cone
698 | 10969 | Fit Scan Coils
698 | 10969 | Fit FoV Coils
698 | 10969 | Fit O Ring to Top of Scan Coil
698 | 10969 | Perform Vacuum Leak Check
698 | 10970 | Fit Column Pumping Body onto the C3 Assembly
698 | 10970 | Fit Test Square Blanking Plate
698 | 10970 | Fit Mu-metal Shield
698 | 10970 | Fit Aperture Changer
698 | 10970 | Fit Circular Blank (PCM)
698 | 10970 | Fit Rectangular Blank (CIV)
698 | 10974 | Fit C1/C2 Body
698 | 10974 | Perform Vacuum Leak Check
698 | 10975 | Scan Coil - LED is On
698 | 10975 | Press Button and LED Goes Off
698 | 10976 | Fit Firing Unit into Gun Head
698 | 10976 | Fit O Ring into Top of C!/C2 Body
698 | 10976 | Fit Gun Head to C1/C2 Body
698 | 10977 | Fit Gun Safety Switch
698 | 10977 | Fit Gun Align Cable To Bracket
698 | 10977 | Fit Assembled Bracket To C1/C2 Body
698 | 10978 | Fit Pole Pieces
698 | 10978 | Fit Bottom Plate
698 | 10978 | Fit Gun Align Coils
698 | 10978 | Fit Anode
698 | 10978 | Check for Cleanliness Using UV Torch
698 | 10980 | Complete Column Vauum Checked
698 | 10980 | Vacuum Isolation Valve Closed
698 | 10957 | Clean and Prepare the Jig and Fit the O ring
698 | 10968 | Fit C3 Cone into Jig (Vac seal facing you)
698 | 10969 | Fit the Sealing Plate onto the Inner C3 Cone
698 | 10969 | Fit Handles and Lower into the Outer Cone
698 | 10969 | Secure the Outer C3 Cone to the Jig
698 | 10969 | Secure the Inner C£ Cone to the Outer Cone
698 | 10969 | Fit Scan Coils
698 | 10969 | Fit FoV Coils
698 | 10969 | Fit O Ring to Top of Scan Coil
698 | 10969 | Perform Vacuum Leak Check
698 | 10970 | Fit Column Pumping Body onto the C3 Assembly
698 | 10970 | Fit Test Square Blanking Plate
698 | 10970 | Fit Mu-metal Shield
698 | 10970 | Fit Aperture Changer
698 | 10970 | Fit Circular Blank (PCM)
698 | 10970 | Fit Rectangular Blank (CIV)
698 | 10974 | Fit C1/C2 Body
698 | 10974 | Perform Vacuum Leak Check
698 | 10975 | Scan Coil - LED is On
698 | 10975 | Press Button and LED Goes Off
698 | 10976 | Fit Firing Unit into Gun Head
698 | 10976 | Fit O Ring into Top of C!/C2 Body
698 | 10976 | Fit Gun Head to C1/C2 Body
698 | 10977 | Fit Gun Safety Switch
698 | 10977 | Fit Gun Align Cable To Bracket
698 | 10977 | Fit Assembled Bracket To C1/C2 Body
698 | 10978 | Fit Pole Pieces
698 | 10978 | Fit Bottom Plate
698 | 10978 | Fit Gun Align Coils
698 | 10978 | Fit Anode
698 | 10978 | Check for Cleanliness Using UV Torch
698 | 10980 | Complete Column Vauum Checked
698 | 10980 | Vacuum Isolation Valve Closed
522 | 10942 | Clean and Inspect O Ring
522 | 10942 | Fit O Ring into Blank
522 | 10943 | Clean and Inspect O Ring
522 | 10943 | Fit O Ring into Blank
522 | 10944 | Clean and Inspect O Rings
522 | 10944 | Fit O Rings into Plate
522 | 10945 | Clean and Inspect O Ring
522 | 10945 | Fit O Ring into Vac Seal
522 | 10945 | Insert Vac Seal into C3 Cone
522 | 10946 | Clean and Inspect O Ring
522 | 10946 | Fit O Ring it to the Lens Cap
522 | 10946 | Mount the End Cap Onto the C3 Cone
522 | 10946 | Tighten the Screws
522 | 10946 | End Cap and C3 Cone Gap Checked
522 | 10947 | Liner Tube Inspected
522 | 10947 | Inspect O Ring Sealing Surfaces
522 | 10947 | Liner Tube Cleaned
522 | 10947 | Clean and Inspect O Rings
522 | 10947 | Fit O Rings onto Coils
| 11380 | S/N =
522 | 10948 | Liner Tube Inspected
522 | 10948 | Inspect O Ring Sealing Surfaces
522 | 10948 | Liner Tube Cleaned
522 | 10948 | Clean and Inspect O Rings
522 | 10948 | Fit O Rings onto Coils
522 | 10949 | Clean and Inspect O Ring
522 | 10949 | Fit O Ring to the Underside of the Body
522 | 10949 | Fit Pumping Guide
522 | 10949 | Check Pumping Guide Alignment
522 | 10949 | Fit M3x10 Countersunk Screws
522 | 10949 | Fit Chamber Pumping Shield
522 | 10949 | Clean and Inspect O Ring
522 | 10949 | Fit O Ring to Upperside of the Body
| 11364 | Continuity (Close Circuit)
| 11364 | Coils to GND (Open Circuit)
522 | 10950 | Inspect Aperture Blade
522 | 10950 | Install the Apertures
522 | 10951 | Disassemble the Firing Unit
522 | 10951 | Clean Listed Parts in Ultrasonic Tank
522 | 10951 | Reassemble Firing Unit and Fit Filament
522 | 10951 | Check Filament Height and Centre
522 | 10952 | Inspect Bores of Plate and Clamp
522 | 10952 | Fit 600um Aperture
522 | 10952 | Clean and Inspect O Ring and Fit to the Underside of Plate
522 | 10953 | Fit an Aperture Insert Into Each Pole Piece
522 | 10953 | Check Aperture Inserts are Securely Fitted
522 | 10954 | Inspect Bores of Plate and Clamp
522 | 10954 | Fit 300um Aperture
522 | 10954 | Clean and Inspect O Ring and Fit to the Underside of Plate
522 | 11198 | Set Up Ski jump on Jig
522 | 11198 | Fit Ski Jump Onto Plate
522 | 11198 | Using Jig Align Ski Slope Bore With Plate Bore.
522 | 11198 | Fit Bottom Plate Using 4x M4X8 CSK Screws
522 | 11198 | Fit Metal Shield
522 | 11198 | Insert Ski Jump Plate and Secure
522 | 11198 | Fit O Ring
522 | 11198 | Clean and Check CIV Sealing O Ring
522 | 11198 | Clean, Check and Fit CIV Body O Ring
522 | 11198 | Check CIV Height Using Jig
522 | 11198 | Fit CIV Onto Pumping Body
522 | 11198 | Check CIV Can Open and Close Without Touching Insert Plate
522 | 11199 | Fit Heater and Cover Plate
522 | 11199 | Fit the Large Cover Plate and Temperature Label
522 | 11199 | Remove Ion Pump Transit Blank
522 | 11199 | Fit Copper Seal
522 | 11199 | Fit and Secure Baffle
522 | 11199 | Align and Secure T Pipe
522 | 11199 | Fit Ion Pump Assy. Onto Column
| 11356 | Drying time of 12hrs.
698 | 10957 | Clean and Prepare the Jig and Fit the O ring
698 | 10968 | Fit C3 Cone into Jig (Vac seal facing you)
698 | 11246 | All Paperwork Sorted and Stapled Together
698 | 11246 | All Paperwork Sorted and Stapled Together
| 11365 | Continuity (Close Circuit)
| 11365 | Coils to GND (Open Circuit)
| 11366 | Continuity (Close Circuit)
| 11366 | Coils to GND (Open Circuit)
522 | 11198 | Fit Ski Jump Onto Plate
522 | 11198 | Using Jig Align Ski Slope Bore With Plate Bore.
522 | 11198 | Fit Bottom Plate Using 4x M4X8 CSK Screws
522 | 11198 | Fit Metal Shield
522 | 11198 | Insert Ski Jump Plate and Secure
522 | 11198 | Fit O Ring
522 | 11198 | Clean and Check CIV Sealing O Ring
522 | 11198 | Clean, Check and Fit CIV Body O Ring
522 | 11198 | Check CIV Height Using Letterbox Jig
522 | 11198 | Fit CIV Onto Pumping Body
522 | 11198 | Check CIV Can Open and Close Without Touching Insert Plate
709 | 11322 | Weight
709 | 11322 | Dimensions Lenght
709 | 11322 | Dimensions Width
709 | 11322 | Dimensions Height
709 | 11322 | Security Seals Serial No
709 | 11322 | Shock Watch Serial Number
709 | 11322 | Tilt Watch Serial No.
709 | 11322 | Weight
709 | 11322 | Dimensions Length
709 | 11322 | Dimensions Width
709 | 11322 | Dimensions Height
709 | 11322 | Security Seals Serial No
709 | 11322 | Shock Watch Serial Number
709 | 11322 | Tilt Watch Serial No.
709 | 11322 | Weight
709 | 11322 | Dimensions Length
709 | 11322 | Dimensions Width
709 | 11322 | Dimensions Height
709 | 11322 | Security Seals Serial No
709 | 11322 | Shock Watch Serial Number
709 | 11322 | Tilt Watch Serial No.
709 | 11322 | Packed By
| 11375 | Position A:
| 11375 | Position B:
| 11375 | Position C:
| 11375 | Position D:
| 11375 | Max. Difference
| 11377 | Has Transzorb been fitted
| 11378 | Continuity (Close Circuit)
| 11378 | Coils to GND (Open Circuit)
| 12834 | Wert =
| 11358 | C1 Coil
| 11358 | C2 Coil
| 11361 | Inspect
| 11361 | Clean
| 11367 | Continuity (Close Circuit)
| 11367 | Coils to GND (Open Circuit)
| 11371 | A - B (< 100kΩ)
| 11371 | C - D (< 10Ω)
| 11371 | B - D (Open Circuit)
| 11373 | S/N =
| 12807 | Wert =
| 12809 | Datei =
| 12810 | Wert =
| 12811 | Datei =
| 12812 | Wert =
| 12813 | Datei =
| 12814 | Wert =
| 12815 | Wert =
| 12835 | Wert =
| 12816 | Wert =
| 12817 | Wert =
| 12819 | Datei =
| 12820 | Wert =
| 12821 | Datei =
| 12822 | Wert =
| 12823 | Datei =
| 12824 | Wert =
| 12825 | Wert =
| 12826 | Wert =
| 12829 | Wert =
| 12830 | Wert =
| 12831 | Wert =
| 12832 | Wert =
| 12833 | Wert =
| 12837 | Wert =
| 12838 | Wert =
| 12839 | Wert =
| 12845 | Datei =
| 12848 | Wert =
| 12849 | Wert =
| 12850 | Datei =
| 12851 | Wert =
| 12852 | Wert =
| 12987 | Wert =
| 12853 | Wert =
| 12854 | Wert =
| 12855 | Wert =
| 12859 | Wert =
| 12861 | Wert =
| 12863 | Wert =
| 12864 | Wert =
| 13077 | Wert =
| 13389 | Wert =
| 13433 | Wert =
| 13444 | Wert =
| 13445 | Wert =
| 13498 | Wert =
| 13499 | Wert =
| 13520 | Wert =
| 13521 | Datei =
| 12983 | Wert =
| 12986 | Wert =
| 13081 | Wert =
| 13109 | Wert =
| 13592 | Wert =
| 13595 | Wert =
| 13679 | Wert =
| 13239 | Wert =
| 13415 | Wert =
| 12596 | Wert =
| 12582 | Wert =
| 12597 | Wert =
| 12614 | Wert =
| 12616 | Wert =
| 12617 | Wert =
| 13299 | Wert =
| 13300 | Wert =
| 13702 | Wert =
| 13335 | Wert =
| 13337 | Wert =
| 13247 | Wert =
| 13559 | Wert =
| 13560 | Wert =
| 13561 | Datei =
| 13562 | Datei =
| 13564 | Wert =
| 13565 | Wert =
| 13267 | Wert =
| 13268 | Wert =
| 13270 | Wert =
| 13271 | Wert =
| 13276 | Wert =
| 13283 | Wert =
| 13393 | Wert =
| 13394 | Wert =
| 13395 | Wert =
| 13396 | Wert =
| 13397 | Wert =
| 13398 | Wert =
| 13399 | Wert =
| 13400 | Wert =
| 13414 | Wert =
| 13452 | Datei =
| 13545 | Wert =
| 13546 | Wert =
| 12628 | Wert =
| 12629 | Wert =
| 12630 | Wert =
| 12631 | Wert =
| 12632 | Wert =
| 12633 | Wert =
| 12634 | Wert =
| 12635 | Wert =
| 12652 | Wert =
| 13465 | Wert =
| 12674 | Wert =
| 12675 | Wert =
| 12676 | Wert =
| 12678 | Wert =
| 12679 | Wert =
| 12685 | Wert =
| 13421 | Wert =
| 13422 | Wert =
| 13423 | Wert =
| 13424 | Wert =
| 13426 | Wert =
| 13441 | Wert =
| 13497 | Wert =
| 13506 | Wert =
| 13508 | Wert =
| 13509 | Wert =
| 13510 | Wert =
| 13531 | Wert =
| 13541 | Wert =
| 12719 | Wert =
| 13676 | Wert =
| 13677 | Wert =
| 13678 | Wert =
| 13684 | Wert =
| 13701 | Wert =
| 12691 | Wert =
| 12692 | Wert =
| 12693 | Wert =
| 12694 | Wert =
| 12695 | Wert =
| 12720 | Wert =
| 12710 | Wert =
| 12712 | Wert =
| 12713 | Wert =
| 12714 | Wert =
| 12715 | Wert =
| 12716 | Wert =
| 12717 | Wert =
| 12718 | Wert =
| 12721 | Wert =
| 12722 | Wert =
| 12723 | Wert =
| 12724 | Wert =
| 12725 | Datei =
| 12726 | Wert =
| 12727 | Wert =
| 12730 | Wert =
| 12743 | Wert =
| 12744 | Wert =
| 12805 | Wert =
| 13450 | Wert =
| 13501 | Wert =
| 13502 | Wert =
| 13525 | Wert =
| 13526 | Wert =
| 13624 | Wert =
| 13664 | Wert =
| 13665 | Wert =
| 12973 | Wert =
| 13666 | Wert =
| 12752 | Wert =
| 12772 | Wert =
| 12773 | Wert =
| 12774 | Wert =
| 12775 | Wert =
| 12776 | Wert =
| 12777 | Wert =
| 12778 | Wert =
| 12779 | Wert =
| 12780 | Wert =
| 12781 | Wert =
| 12782 | Wert =
| 12783 | Wert =
| 12784 | Wert =
| 12785 | Wert =
| 12786 | Wert =
| 12787 | Wert =
| 12788 | Wert =
| 12789 | Wert =
| 12790 | Wert =
| 12791 | Wert =
| 12792 | Wert =
| 12793 | Wert =
| 12794 | Wert =
| 12795 | Wert =
| 12799 | Wert =
| 12800 | Wert =
| 12802 | Wert =
| 12804 | Wert =
| 13429 | Wert =
| 13430 | Wert =
| 13431 | Wert =
| 13432 | Wert =
| 12895 | Wert =
| 12876 | Wert =
| 12877 | Wert =
| 12878 | Wert =
| 12879 | Wert =
| 12880 | Wert =
| 12881 | Wert =
| 12882 | Wert =
| 12883 | Wert =
| 12886 | Wert =
| 13512 | Wert =
| 13591 | Wert =
| 13705 | Wert =
| 12896 | Wert =
| 12900 | Wert =
| 13076 | Wert =
| 13385 | Wert =
| 13496 | Wert =
| 13513 | Wert =
| 13530 | Wert =
| 13601 | Wert =
| 12912 | Wert =
| 12920 | Wert =
| 13466 | Wert =
| 12924 | Wert =
| 12925 | Wert =
| 12926 | Wert =
| 13593 | Wert =
| 12929 | Wert =
| 12931 | Wert =
| 12933 | Wert =
| 12935 | Wert =
| 12936 | Wert =
| 12938 | Wert =
| 12940 | Wert =
| 12942 | Wert =
| 12944 | Wert =
| 12946 | Wert =
| 12948 | Wert =
| 13527 | Wert =
| 13528 | Wert =
| 13611 | Wert =
| 13612 | Wert =
| 13669 | Wert =
| 12955 | Wert =
| 12956 | Wert =
| 12960 | Wert =
| 12967 | Wert =
| 12979 | Wert =
| 12988 | Wert =
| 12990 | Wert =
| 12991 | Wert =
| 12992 | Wert =
| 12996 | Wert =
| 12997 | Wert =
| 12998 | Wert =
| 12999 | Wert =
| 13000 | Wert =
| 13008 | Wert =
| 13012 | Wert =
| 13442 | Wert =
| 13029 | Wert =
| 13030 | Wert =
| 13031 | Wert =
| 13035 | Wert =
| 13036 | Wert =
| 13037 | Wert =
| 13041 | Wert =
| 13042 | Wert =
| 13046 | Wert =
| 13047 | Wert =
| 13048 | Wert =
| 13052 | Wert =
| 13053 | Wert =
| 13054 | Wert =
| 13056 | Wert =
| 13057 | Wert =
| 13058 | Wert =
| 13059 | Wert =
| 13062 | Wert =
| 13063 | Wert =
| 13064 | Wert =
| 13065 | Wert =
| 13067 | Wert =
| 13068 | Wert =
| 13069 | Wert =
| 13070 | Wert =
| 13071 | Wert =
| 13388 | Wert =
| 13455 | Wert =
| 13456 | Wert =
| 13457 | Wert =
| 13458 | Wert =
| 13459 | Wert =
| 13514 | Wert =
| 13599 | Wert =
| 13652 | Wert =
| 13653 | Wert =
| 13654 | Wert =
| 13655 | Wert =
| 13113 | Wert =
| 13115 | Wert =
| 13116 | Wert =
| 13118 | Wert =
| 13123 | Wert =
| 13606 | Wert =
| 13132 | Wert =
| 13133 | Wert =
| 13134 | Wert =
| 13135 | Wert =
| 13136 | Wert =
| 13137 | Wert =
| 13143 | Wert =
| 13144 | Wert =
| 13145 | Wert =
| 13146 | Wert =
| 13151 | Wert =
| 13152 | Wert =
| 13153 | Wert =
| 13515 | Wert =
| 13154 | Wert =
| 13156 | Wert =
| 13157 | Wert =
| 13158 | Wert =
| 13160 | Wert =
| 13162 | Datei =
| 13163 | Datei =
| 13386 | Wert =
| 13516 | Wert =
| 13517 | Wert =
| 13662 | Wert =
| 13663 | Wert =
| 13170 | Wert =
| 13172 | Wert =
| 13179 | Wert =
| 13182 | Wert =
| 13183 | Wert =
| 13185 | Wert =
| 13187 | Wert =
| 13188 | Wert =
| 13201 | Wert =
| 13202 | Wert =
| 13203 | Wert =
| 13204 | Wert =
| 13205 | Wert =
| 13670 | Wert =
| 13208 | Wert =
| 13210 | Wert =
| 13537 | Wert =
| 13570 | Wert =
| 13625 | Wert =
| 13671 | Wert =
| 13672 | Wert =
| 13348 | Wert =
| 13349 | Wert =
| 13353 | Wert =
| 13359 | Wert =
| 13360 | Wert =
| 13366 | Wert =
| 13371 | Wert =
| 13418 | Wert =
| 13375 | Wert =
| 13377 | Datei =
| 13475 | Wert =
| 13476 | Wert =
| 13490 | Wert =
| 13682 | Wert =
| 13481 | Wert =
| 13483 | Datei =
| 13628 | Wert =
| 13629 | Wert =
| 13630 | Wert =
| 13631 | Wert =
| 13632 | Wert =
| 13633 | Wert =
| 13634 | Wert =
| 13635 | Wert =
| 13656 | Wert =
| 13658 | Wert =
2 | 12 | Serial No.
2 | 12 | Version
2 | 12 | Qualification Marks Checked
4 | 39 | Turbo Pump S/No.
4 | 39 | Qualification Marks Checked
24 | 272 | Chamber secured to vibration plate (N/A to XB 550L)
| 11373 | S/N =
| 11380 | S/N =
30 | 300 | X =
30 | 300 | Y =
2 | 7 | Record IPPSU S/N
2 | 7 | Set DIP switches
2 | 7 | Qualification Marks Checked
2 | 8 | Record IPPSU S/N
2 | 8 | Set DIP switches
2 | 8 | Qualification Marks Checked
2 | 10 | Record AC unit S/N
2 | 10 | Qualification Marks Checked
2 | 12 | Serial No.
2 | 12 | Version
2 | 12 | Qualification Marks Checked
4 | 39 | Turbo Pump S/No.
4 | 39 | Qualification Marks Checked
24 | 272 | Chamber lowered onto foam blocks (N/A to XB 550L)
| 11356 | Drying time of 12hrs.
| 11357 | C1 Coil
| 11357 | C2 Coil
| 11358 | C1 Coil
| 11358 | C2 Coil
| 11361 | Inspect
| 11361 | Clean
| 11364 | Continuity (Close Circuit)
| 11364 | Coils to GND (Open Circuit)
| 11365 | Continuity (Close Circuit)
| 11365 | Coils to GND (Open Circuit)
| 11366 | Continuity (Close Circuit)
| 11366 | Coils to GND (Open Circuit)
| 11367 | Continuity (Close Circuit)
| 11367 | Coils to GND (Open Circuit)
| 11371 | A - B (< 100kΩ)
| 11371 | C - D (< 10Ω)
| 11371 | B - D (Open Circuit)
| 11373 | S/N =
2 | 7 | Record IPPSU S/N
2 | 7 | Set DIP switches
2 | 7 | Qualification Marks Checked
2 | 8 | Record IPPSU S/N
2 | 8 | Set DIP switches
2 | 8 | Record IPPSU S/N
2 | 8 | Set DIP switches
2 | 8 | Qualification Marks Checked
2 | 10 | Qualification Marks Checked
2 | 10 | Record AC unit S/N
2 | 10 | Qualification Marks Checked
7 | 1650 | S/N =
389 | 10873 | Number of Detectors
389 | 10873 | Detector Serial Number(s)
389 | 10873 | Detector Resolution
389 | 10873 | Detector Area
389 | 10873 | Detector Installed
389 | 10873 | In/Out Stops Set
389 | 10873 | Isolation Check of Detector & Column >2MΩ
389 | 10873 | Install Remoting & Setup Network Connection
389 | 10873 | Bruker Esprit Software Version
389 | 10873 | Install Bruker Esprit Software
389 | 10873 | Bruker EDS Manager Version
389 | 10873 | Install EDS Manager
389 | 10873 | Install Bruker Licences
389 | 10873 | Setup Pulse Processor
389 | 10873 | Check Standard Timing Cards Fitted
389 | 10873 | Install Detector Data
389 | 10873 | Check Microscope driver
389 | 10873 | Setup SEM Conditions
389 | 10873 | Determine Analytical Working Distance
389 | 10873 | Calibrate Detector
389 | 10873 | Calibrate Image in Espirit
646 | 10653 | Number of Detectors
646 | 10653 | Detector Serial Number
646 | 10653 | Detector Resolution
646 | 10653 | Detector Area
646 | 10653 | Detector Installed
646 | 10653 | In/Out Stops Set
646 | 10653 | Isolation Check of Detector & Column >2MΩ
646 | 10653 | Install Remoting & Setup Network Connection
646 | 10653 | Bruker Esprit Software Version
646 | 10653 | Install Bruker Esprit Software
646 | 10653 | Bruker EDS Manager Version
646 | 10653 | Install EDS Manager
646 | 10653 | Install Bruker Licences
646 | 10653 | Setup Pulse Processor
646 | 10653 | Check Standard Timing Cards Fitted
646 | 10653 | List Optional Timing Cards Fitted
646 | 10653 | Install Detector Data
646 | 10653 | Check Microscope driver
646 | 10653 | Setup SEM Conditions
646 | 10653 | Determine Analytical Working Distance
646 | 10653 | Calibrate Detector
646 | 10653 | Calibrate Image in Espirit
646 | 84 | Language Set to English (United States
646 | 84 | Ethernet Connection Setup
646 | 84 | Install Aztec Software
646 | 84 | AZtec Software Version
646 | 84 | Install Licences
646 | 84 | Run AZtec Assist+
646 | 84 | Connected to Test Tool Module
389 | 10873 | Number of Detectors
389 | 10873 | Detector Serial Number(s)
389 | 10873 | Detector Resolution
389 | 10873 | Detector Area
389 | 10873 | Detector Installed
389 | 10873 | In/Out Stops Set
389 | 10873 | Isolation Check of Detector & Column >2MΩ
389 | 10873 | Install Remoting & Setup Network Connection
389 | 10873 | Bruker Esprit Software Version
389 | 10873 | Install Bruker Esprit Software
389 | 10873 | Bruker EDS Manager Version
389 | 10873 | Install EDS Manager
389 | 10873 | Install Bruker Licences
389 | 10873 | Setup Pulse Processor
389 | 10873 | Check Standard Timing Cards Fitted
389 | 10873 | List Optional Timing Cards Fitted
389 | 10873 | Install Detector Data
389 | 10873 | Check Microscope driver
389 | 10873 | Setup SEM Conditions
389 | 10873 | Determine Analytical Working Distance
389 | 10873 | Calibrate Detector
389 | 10873 | Calibrate Image in Espirit
20 | 11038 | PC S/N
20 | 11038 | Hard disk Image (Version)
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Record Windows Product Key
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
20 | 11270 | Predictive Service S/W installed on Customer PC.
25 | 80 | Ext V =
25 | 80 | Measured Speciment Current =
20 | 11038 | PC S/N
20 | 11038 | Hard disk Image (Version)
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
389 | 67 | Install ZEN core EM.Exe
389 | 67 | ZEN core Software Version
389 | 67 | Setup Exchange Position
389 | 67 | Capture Image
20 | 252 | PC S/N
20 | 252 | Hard disk Image (Version)
20 | 252 | PC/Model (SAP No.)
20 | 252 | Record Operating System
20 | 252 | Disable Renesas Power Management
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Register COM API in ASAP engine
20 | 252 | Run ASAP Screening Program
20 | 252 | Approved Datazone Loaded
46 | 11123 | Power system down.
46 | 11123 | Blow water out of the cooling system.
46 | 11123 | Chamber secured to vibration plate (N/A to XB 550L)
46 | 11123 | Chamber Sercuring Screws tightened to 25NM
46 | 11123 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 11123 | Secure Airfeet with foam blocks.
46 | 11123 | Secure Turbo pump with foam blocks.
46 | 11123 | Fit Covers to all unused connectors.
46 | 11123 | Mains and Earth cable stowed on Mains unit.
46 | 11123 | Fit shipping Eyes to transit blocks.
46 | 11123 | Put thread lock on pendulum screws and transit block Screws.
46 | 11123 | All doors closed and secured.
46 | 11123 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11123 | Print shipping stickers.
46 | 11123 | Fit 8x M6 x 55mm cap head screws to stage door
46 | 11123 | Power system down.
46 | 11123 | Blow water out of the cooling system.
46 | 11123 | Chamber secured to vibration plate (N/A to XB 550L)
46 | 11123 | Chamber securing screws tightened to 25nm (N/A to XB 550L)
46 | 11123 | Chamber Sercuring Screws Tightness Checked By Name:
46 | 11123 | Secure Airfeet with foam blocks.
46 | 11123 | Secure Turbo pump with foam blocks.
46 | 11123 | Fit Covers to all unused connectors.
46 | 11123 | Mains and Earth cable stowed on Mains unit.
46 | 11123 | Fit shipping Eyes to transit blocks.
46 | 11123 | Put thread lock on pendulum screws and transit block Screws.
46 | 11123 | All doors closed and secured.
46 | 11123 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11123 | Print shipping stickers.
46 | 11123 | Fit 8x M6 x 55mm cap head screws to stage door
46 | 11123 | Power system down.
46 | 11123 | Blow water out of the cooling system.
46 | 11123 | Chamber secured to vibration plate (N/A to XB 550L)
46 | 11123 | Chamber securing screws tightened to 25nm (N/A to XB 550L)
46 | 11123 | Chamber securing screws tightness checked. Type your initials (For XB 550L type N/A)
46 | 11123 | Secure Airfeet with foam blocks.
46 | 11123 | Secure Turbo pump with foam blocks.
46 | 11123 | Fit Covers to all unused connectors.
46 | 11123 | Mains and Earth cable stowed on Mains unit.
46 | 11123 | Fit shipping Eyes to transit blocks.
46 | 11123 | Put thread lock on pendulum screws and transit block Screws.
46 | 11123 | All doors closed and secured.
46 | 11123 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11123 | Print shipping stickers.
46 | 11123 | Fit 8x M6 x 55mm cap head screws to stage door
102 | 11077 | ZEN Software Installed
22 | 268 | Carry Out Gun Run UP
22 | 268 | Save Gun Monitor Log as Gun Run Up_TIP SN
40 | 1552 | Collector Working
40 | 1552 | Scintillator Working
40 | 85 | Install ZEN core EM.Exe
40 | 85 | ZEN core Software Version
40 | 85 | Setup Exchange Position
40 | 85 | Capture Image
84 | 86 | Chamber vaccuum tracks demand
41 | 343 | Ensure Cladding geometry is correct and all parts aligned.
41 | 343 | Ensure the suppliers Cladding list for both kits is put into the customers pack.
22 | 79 | Carry Out Gun Run UP
22 | 79 | Save Gun Monitor Log as Gun Run Up_TIP SN
22 | 79 | Confirm that you are aware that the Filament I value used on Raith systems is 20mA less than the Datasheet value.
22 | 79 | Carry Out Gun Run UP
22 | 79 | Save Gun Monitor Log as Gun Run Up_TIP SN
22 | 79 | Confirm that you are aware that the Filament I value used on Raith systems is 20mA less than the Datasheet value.
22 | 79 | Carry Out Gun Run UP
22 | 79 | Save Gun Monitor Log as Gun Run Up_TIP SN
22 | 79 | Confirm that you are aware that the Filament I value used on Raith systems is 20mA less than the Datasheet value.
22 | 79 | Carry Out Gun Run UP
22 | 79 | Save Gun Monitor Log as Gun Run Up_TIP SN
25 | 80 | Specimen I measured with SCM
25 | 80 | Specimen I measured with Pico-ammeter
25 | 80 | Specimen I measured with SCM
25 | 80 | Specimen I measured with Pico-ammeter
25 | 80 | Is the SCM measurement within +/-5pa of the Pico-ammeter measurement?
25 | 80 | Specimen I measured with SCM
25 | 80 | Specimen I measured with Pico-ammeter
25 | 80 | Is the SCM measurement within +/-5pa of the Pico-ammeter measurement?
84 | 86 | Chamber vaccuum tracks demand
397 | 10926 | ZEN Software Installed
102 | 11077 | ZEN Software Installed
41 | 343 | Ensure Cladding geometry is correct and all parts aligned.
41 | 343 | Ensure the suppliers Cladding list for both kits is put into the customers pack.
41 | 343 | Ensure Cladding geometry is correct and all parts aligned.
41 | 343 | Ensure the suppliers Cladding list for both kits is put into the customers pack.
30 | 303 | Actual
41 | 343 | Ensure Cladding geometry is correct and all parts aligned.
41 | 343 | Ensure the suppliers Cladding list for both kits is put into the customers pack.
22 | 79 | Confirm that you are aware that the Filament I value used on Raith systems is 20mA less than the Datasheet value.
22 | 79 | Carry Out Gun Run UP
22 | 79 | Save Gun fxsdvbgfdsgvdsgdf Monitor Log as Gun Run Up_TIP SN
389 | 89 | Test Valves Using Raman Test Interlock Box
388 | 11425 | 15kv 0.9nm
388 | 11425 | 15KV result
388 | 11425 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 1kv 1.3nm
388 | 11425 | 1KV result
388 | 11425 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 500V 2nm
388 | 11425 | 500V result
388 | 11425 | Store 500V Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 15kv 0.7nm
388 | 11426 | 15KV result
388 | 11426 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 1kv 1.2nm
388 | 11426 | 1KV result
388 | 11426 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 500V 1.9nm
388 | 11426 | 500V result
388 | 11426 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
22 | 262 | Carry out overnight ‘Full service Bakeout’ and note actual duration.
22 | 262 | Confirm that all Thermocouples are reading.
22 | 262 | Confirm that all Ion pump heater elements and flange heaters are getting hot.
22 | 262 | Gun Vacuum post Bakeout. Ensure column has cooled prior to taking reading.
22 | 262 | After completion of Bakeout Torque the Gun head bolts to 25NM.
37 | 329 | Stig: 30μm, All EHT’s is ≤ 20%
37 | 329 | Stig: Off Axis Aper, All EHT’s is ≤ 30%
37 | 329 | EHT <1kv: 30μm Aperture Gun/Aperture Aligns ≤ 60%
37 | 329 | EHT ≥1kv: 30μm Aperture Gun/Aperture Aligns ≤ 50%
25 | 275 | Are Gun Head clamping screws secure?
25 | 275 | Are Condensor clamping screws secure?
25 | 275 | Save Emission image @1Kv (cond off)
25 | 275 | Check emission image against 'emissions distortion criteria
25 | 275 | Save Emission image @10Kv (cond off)
388 | 3293 | Actual (Spec: $[ResolutionBSD]nm)
388 | 3293 | Store ASB/BSD Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 15kv 0.7nm
388 | 11426 | 15KV result
388 | 11426 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 1kv 1.2nm
388 | 11426 | 1KV result
388 | 11426 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 500V 1.9nm
388 | 11426 | 500V result
388 | 11426 | Store 500V Resolution Images and Pate image for 2 x consecutive images
388 | 3292 | 1kV (Spec: $[Resolution1kV]nm)
388 | 3292 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 3292 | 15kV (Spec: $[Resolution15kV]nm)
388 | 3292 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
5 | 98 | Torque done up to 25Nm
674 | 113 | Torque done up to 25Nm
674 | 113 | Torque done up to 25Nm
389 | 3299 | Edit Chamberscope Macro
389 | 3299 | Set Chamberscope Focus
389 | 3299 | Chamberscope 2 - Select Port or Not Fitted
389 | 3299 | Chamberscope 2 - Check Camera 2 for Focus
389 | 3299 | Chamberscope 2 - Functional Check-Black/White + Colour Mode
389 | 3299 | Chamberscope 2 - Run wxPropview to Remove Defective Pixels
389 | 3299 | Chamberscope 2 - Add Icon to Toolbar
389 | 3299 | Edit Chamberscope Macro
389 | 3299 | Set Chamberscope Focus
389 | 3299 | Chamberscope 2 - Select Port or Not Fitted
389 | 3299 | Chamberscope 2 - Check Camera 2 for Focus
389 | 3299 | Chamberscope 2 - Functional Check-Black/White + Colour Mode
389 | 3299 | Chamberscope 2 - Run wxPropview to Remove Defective Pixels
389 | 3299 | Chamberscope 2 - Add Icon to Toolbar
7 | 11258 | Continuity ok
389 | 11074 | Check each individual segment in normal and inverse modes
389 | 11074 | Analyze Segments
389 | 11074 | Gain Ranges
389 | 11074 | Pneumatics
389 | 11074 | Alignment
389 | 11074 | Vac Degradation < 3x10-5 mbar
389 | 11074 | CZ Detector Cal
389 | 11074 | Z Adjustment
389 | 3314 | Type of Detector
389 | 3314 | Detector Serial number
389 | 3314 | DDP Module Serial Number
389 | 3314 | Fit Flange & Mount Detector
389 | 3314 | Install EDS PCBs into PC
389 | 3314 | Installation of DPP Module & Cables
389 | 3314 | Configure PC Network & Test
389 | 3314 | Disable Windows10 Fast Startup
389 | 3314 | Install ApexKey.edx Licences
389 | 3314 | Install Genesis.edx Licences (if Applicable)
389 | 3314 | Fit Dongle
389 | 3314 | APEX Software Version
389 | 3314 | Install APEX Software
389 | 3314 | Install Remoting
389 | 3314 | Setup RConfigure
389 | 3314 | Run API Test
389 | 3314 | Configure imaging – Run SGShell
389 | 3314 | Calibrate Detector – Run eDPPShell
389 | 3314 | Save Calibration Report
389 | 3314 | Detector Resolution - ≥ 129 eV
389 | 3314 | Save calibration data on exit
389 | 3314 | Run APEX
389 | 3314 | Check Spectrum
389 | 3314 | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected
389 | 3314 | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected
389 | 3314 | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected
389 | 3314 | Check Mapping
389 | 3314 | Check Line Scanning
389 | 3314 | Save Aquired Spectrum, Mapping and Line Scanning Images to Image Folder
389 | 3314 | Save Calibration Data to the Licence Folderon the SmartSem USB stick.
389 | 3314 | Save Calibration Data to the Image Folder
675 | 10824 | All Cables Are Neatly Routed
675 | 10824 | All Plugs Are Secured Correctly
675 | 10824 | All Earth Cables Are Tightly Secured
675 | 10824 | Verify the EHT gun assembly earth is torqued p to 25Nm.
646 | 10530 | Calibration of Bias
646 | 10530 | Functional Check
646 | 10530 | Function Image
646 | 10530 | Store VPSE Functional Image
646 | 10529 | BSD Diode Pre-test
646 | 10529 | Select Type
646 | 10529 | Black Level Calibration
646 | 10529 | Check Diode Segment Function
646 | 10529 | Check All Gain ranges
646 | 10529 | Check Auto range
646 | 10529 | Check Auto Level
646 | 10529 | Setup and Check TOPO
646 | 10529 | Take 3 images - Atomic, 5pA, 2KV
646 | 10529 | Store BSD Functional Images
646 | 10529 | BSD Diode Pre-test
646 | 10529 | Select Type
646 | 10529 | Black Level Calibration
646 | 10529 | Check Diode Segment Function
646 | 10529 | Check All Gain ranges
646 | 10529 | Check Auto range
646 | 10529 | Check Auto Level
646 | 10529 | Setup and Check TOPO
646 | 10529 | Take 3 images - Atomic, 5pA, 2KV
646 | 10529 | Store BSD Functional Images-Atomic, 5pA, 2KV
646 | 10531 | Black Level Cal
646 | 10531 | C2D Bias Target Check
646 | 10531 | Gain Range Check
646 | 10531 | Auto Level Check
646 | 10531 | Auto Range Check
646 | 10531 | Function Image
646 | 10531 | Store C2D Functional Image
646 | 10534 | Alignment of Ring (Using Jig)
646 | 10534 | Black Level Calibration
646 | 10534 | Functional Check
646 | 10534 | Store C2DX Functional Images - Air, water
646 | 10536 | Check 5KV is achievable using the Bias Slider
646 | 10536 | safety Check - Remove PL1 (Vac Ready from vac switch
646 | 10536 | Safety Check - Remove PL2 (Turbo Speed signal)
646 | 10536 | Safety Check - Remove PL4 (CAN Bus signal)
646 | 10536 | Functional Check and Images (Using the BSD Detector)
646 | 10536 | Store Beam De-Cel Functional Images - OV and 5KV
43 | 11341 | 1kV ESB
43 | 11341 | Store 1kV ESB Customer Image
646 | 10535 | Black Level Calibration
646 | 10535 | Functional Check
646 | 10535 | Function Image
646 | 10535 | Store SCD Functional Image
646 | 10537 | Functional Check
646 | 10537 | Function Image
646 | 10537 | Store CL Functional Image
646 | 10544 | Installation
646 | 10544 | Functional Test
646 | 10544 | Store SmartStitch Functional Image
646 | 10652 | Type of Detector
646 | 10652 | Detector Serial number
646 | 10652 | DDP Module Serial Number
646 | 10652 | Fit Flange & Mount Detector
646 | 10652 | Install EDS PCB's Into PC
646 | 10652 | Installation of DPP Module & Cables
646 | 10652 | Configure PC Network & Test
646 | 10652 | Disable Windows10 Fast Startup
646 | 10652 | Install Licences
646 | 10652 | Fit Dongle
646 | 10652 | APEX Software Version
646 | 10652 | Install APEX Software
646 | 10652 | Install Remoting
646 | 10652 | Setup RConfigure
646 | 10652 | Run API Test
646 | 10652 | Configure Imaging-Run SGShell
646 | 10652 | Calibrate Detector-Run eDPP Shell
646 | 10652 | Save Calibration Report
646 | 10652 | Detector Resolution (measured using Cu) ≈ 129 eV
646 | 10652 | Save Calibration on Exit
646 | 10652 | Run APEX
646 | 10652 | Check Spectrum
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected
646 | 10652 | Check Mapping
646 | 10652 | Check Line Scanning
646 | 10652 | Save Aquired Spectrum, Mappimg & Line Scanning to the Image Folder
646 | 10652 | Store Aquired Spectrum, Mappimg & Line Scanning Images
646 | 10652 | Save Calibration Data to the Licence Folder on the SmartSEM USB Stick
646 | 10652 | Save Calibration Data to the Image Folder
646 | 10652 | Type of Detector
646 | 10652 | Detector Serial number
646 | 10652 | DDP Module Serial Number
646 | 10652 | Fit Flange & Mount Detector
646 | 10652 | Install EDS PCB's Into PC
646 | 10652 | Installation of DPP Module & Cables
646 | 10652 | Configure PC Network & Test
646 | 10652 | Disable Windows10 Fast Startup
646 | 10652 | Install Licences
646 | 10652 | Fit Dongle
646 | 10652 | APEX Software Version
646 | 10652 | Install APEX Software
646 | 10652 | Install Remoting
646 | 10652 | Setup RConfigure
646 | 10652 | Run API Test
646 | 10652 | Configure Imaging-Run SGShell
646 | 10652 | Calibrate Detector-Run eDPP Shell
646 | 10652 | Save Calibration Report
646 | 10652 | Detector Resolution (measured using Cu) ≈ 129 eV
646 | 10652 | Store Calibration Report
646 | 10652 | Save Calibration on Exit
646 | 10652 | Run APEX
646 | 10652 | Check Spectrum
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected
646 | 10652 | Check Mapping
646 | 10652 | Check Line Scanning
646 | 10652 | Save Aquired Spectrum, Mappimg & Line Scanning to the Image Folder
646 | 10652 | Store Aquired Spectrum (STD, Quick, HD), Mappimg & Line Scanning Images
646 | 10652 | Save Calibration Data to the Licence Folder on the SmartSEM USB Stick
646 | 10652 | Store Calibration Data File
646 | 10652 | Save Calibration Data to the Image Folder
| 11391 | In-Lens and ESB
| 11391 | EO Dynamic PCB
| 11391 | EO Static PCB
| 11391 | Scan Gen PCB (ASAP or SVSEM)
| 11391 | Check that all column connections are securely fastened.
| 11391 | Ensure torque of gun earth is 20NM
646 | 10547 | ZEN Version
646 | 10547 | Correct Version of Windows 10 Installed
646 | 10547 | Install ZEN
646 | 10547 | Install Licence
646 | 10547 | Install Remote API Client
646 | 10547 | Functional Test
646 | 10547 | Save Image
646 | 10547 | ZEN Version
646 | 10547 | Correct Version of Windows 10 Installed
646 | 10547 | Install ZEN
646 | 10547 | Install Licence
646 | 10547 | Install Remote API Client
646 | 10547 | Functional Test
646 | 10547 | Save Snapped Image
40 | 121 | Collector Working
646 | 10547 | ZEN Version
646 | 10547 | Correct Version of Windows 10 Installed
646 | 10547 | Install ZEN
646 | 10547 | Install Licence
646 | 10547 | Install Remote API Client
646 | 10547 | Functional Test
646 | 10547 | Save Snapped Image
646 | 10547 | Store Snapped Image
646 | 10545 | Correct PC - 347823-9024-660 Installed
646 | 10545 | Print Out SmartPI Check List (Link in Documents)
646 | 10545 | Install SmartSEM Licences
646 | 10545 | Install MIL Redistribution
646 | 10545 | Fit MIL Dongle
646 | 10545 | SmartPI Software Version
646 | 10545 | Install SmartPI Software
646 | 10545 | I/O-SCAN Card Fitted (Bruker Only)
646 | 10545 | I/O-SCAN Card Calibrated
646 | 10545 | Setup & Run Application
646 | 10545 | Check Results & Image Montage
646 | 10545 | Store Compressed Database and Montage Image
646 | 10545 | Correct PC - 347823-9024-660 Installed
646 | 10545 | Print Out SmartPI Check List (Link in Documents)
646 | 10545 | Install SmartSEM Licences
646 | 10545 | Install MIL Redistribution
646 | 10545 | Fit MIL Dongle
646 | 10545 | SmartPI Software Version
646 | 10545 | Install SmartPI Software
646 | 10545 | I/O-SCAN Card Fitted (Bruker Only)
646 | 10545 | I/O-SCAN Card Calibrated
646 | 10545 | Setup & Run Application
646 | 10545 | Check Results & Image Montage
646 | 10545 | Store Compressed Database and Montage Image
646 | 10545 | Correct PC - 347823-9024-660 Installed
646 | 10545 | Print Out SmartPI Check List (Link in Documents)
646 | 10545 | Install SmartSEM Licences
646 | 10545 | Install MIL Redistribution
646 | 10545 | Fit MIL Dongle
646 | 10545 | SmartPI Software Version
646 | 10545 | Install SmartPI Software
646 | 10545 | I/O-SCAN Card Fitted (Bruker Only)
646 | 10545 | I/O-SCAN Card Calibrated
646 | 10545 | Setup & Run Application
646 | 10545 | Check Results & Image Montage
646 | 10545 | Store Compressed Database,Montage Image, Log files etc
26 | 281 | Type in measured value (Limit <= 30um)
26 | 281 | Store ESB Consentricity Image
43 | 11341 | 1kV ESB
43 | 11341 | Store 1kV ESB Customer Image
646 | 10654 | Print Out Mineralogic Check List (Link in Documents)
646 | 10654 | Correct PC - Installed
646 | 10654 | Install SmartSEM Licences
646 | 10654 | Instal MIL Redistribution
646 | 10654 | Fit MIL Dongle
646 | 10654 | Install Mineralogic Software
646 | 10654 | Calibrate I/O-SCAN card
646 | 10654 | Setup & Run Application
646 | 10654 | Check Results & Image Montage
646 | 10654 | 24 Hour Run
646 | 10654 | Store Compressed Database,Montage Image, Log files etc
25 | 278 | Actual specimen current
25 | 278 | Nominal Extractor Voltage
25 | 278 | Actual specimen current
25 | 278 | Nominal Extractor Voltage (Extractor voltage must be no higher than 500V over datasheet value)
25 | 278 | Actual specimen current
25 | 278 | Ext V (Extractor voltage must be no higher than 500V over datasheet value)
21 | 12567 | Enter the FIB Extension number using EEPromUpdater.exe
646 | 10652 | Type of Detector
646 | 10652 | Detector Serial number
646 | 10652 | DDP Module Serial Number
646 | 10652 | Fit Flange & Mount Detector
646 | 10652 | Install EDS PCB's Into PC
646 | 10652 | Installation of DPP Module & Cables
646 | 10652 | Configure PC Network & Test
646 | 10652 | Disable Windows10 Fast Startup
646 | 10652 | Install ApexKey.edx Licences
646 | 10652 | Install Genesis.edx Licences (if Applicable)
646 | 10652 | Fit Dongle
646 | 10652 | APEX Software Version
646 | 10652 | Install APEX Software
646 | 10652 | Install Remoting
646 | 10652 | Setup RConfigure
646 | 10652 | Run API Test
646 | 10652 | Configure Imaging-Run SGShell
646 | 10652 | Calibrate Detector-Run eDPP Shell
646 | 10652 | Save Calibration Report
646 | 10652 | Detector Resolution (measured using Cu) ≈ 129 eV
646 | 10652 | Store Calibration Report
646 | 10652 | Save Calibration on Exit
646 | 10652 | Run APEX
646 | 10652 | Check Spectrum
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected
646 | 10652 | Check Mapping
646 | 10652 | Check Line Scanning
646 | 10652 | Save Aquired Spectrum, Mappimg & Line Scanning to the Image Folder
646 | 10652 | Store Aquired Spectrum (STD, Quick, HD), Mappimg & Line Scanning Images
646 | 10652 | Save Calibration Data to the Licence Folder on the SmartSEM USB Stick
646 | 10652 | Store Calibration Data File
646 | 10652 | Save Calibration Data to the Image Folder
26 | 282 | Type in measured value (limit: <= 15um)
26 | 282 | Store Inlens to ESB image
21 | 256 | Store System checker full report
26 | 282 | Type in measured value (limit: <= 15um)
26 | 282 | Store Inlens to ESB image
26 | 281 | Type in measured value (Limit <= 30um)
26 | 281 | Store ESB Consentricity Image
26 | 396 | Type in measured value (Limit: <= 50um)
26 | 396 | Store ESB Concentricity image
25 | 275 | Are Gun Head clamping screws secure?
25 | 275 | Are Condensor clamping screws secure?
25 | 275 | Store 1kV emission image (Condensor off)
25 | 275 | Check emission image against 'emissions distortion criteria'
25 | 275 | Store 10kV emission image (Condensor off)
25 | 11313 | Are Gun Head clamping screws secure?
25 | 11313 | Are Condensor clamping screws secure?
25 | 11313 | Store 1kV emission image (Condensor off)
25 | 11313 | Store 10kV emission image (Condensor off)
30 | 297 | Check Mag cal errors are <1% in X & Y axes in all 3 Mag Ranges. Orthoganality ± 1deg.
30 | 297 | Store accuracy images (Mag 1, Mag 2, Mag 3 and Orthog)
30 | 11071 | Check Mag cal errors are <2% in X & Y axes in all 3 Mag Ranges. Orthog. ± 1deg
30 | 11071 | Store accuracy images (Mag 1, Mag 2, Mag 3 and Orthog)
6 | 1253 | Confirm all vacuum modules are Green in Vac Test
30 | 11071 | Check Mag cal errors are <2% in X & Y axes in all 3 Mag Ranges. Orthog. ± 1deg
30 | 11071 | Store accuracy images (Mag 1, Mag 2, Mag 3 and Orthog)
30 | 299 | Macro successfully completed.
30 | 299 | Store Kompass results window
30 | 303 | Actual
30 | 303 | Store stage drift macro results image
32 | 311 | Actual =
32 | 311 | Store objective drift results
33 | 316 | Actual for 1kV =
33 | 316 | Store 1KV Resolution Images and Pate image
33 | 316 | Actual for 15kV =
33 | 316 | Store 15KV Resolution Images and Pate image
33 | 11062 | 1kV =
33 | 11062 | Store 1KV Resolution Images and Pate image
33 | 11062 | 15kV =
33 | 11062 | Store 15KV Resolution Images and Pate image
33 | 316 | Actual for 1kV =
33 | 316 | Store 1KV Resolution Images and Pate image
33 | 316 | Actual for 15kV =
33 | 316 | Store 15KV Resolution Images and Pate image
38 | 335 | Measured Specimen Current @ 20kV =
38 | 335 | Store Optiprobe log
38 | 337 | Blanking Drift Distance Median (Limit <= 300nm) =
38 | 337 | Blanking Drift Time Median (Limit <= 3.5 Sec) =
38 | 337 | Corner Test Max Drift (Limit <= 300nm) =
38 | 337 | Store Autobeamshift results PDF
39 | 339 | Actual =
39 | 339 | After conditioning confirm there are no EMServer error messages
39 | 339 | Save Gun Monitor as Probestability.MN2
39 | 339 | Store Gun monitor trace as Probestability.mn2
43 | 345 | Refer to Work Instruction for column conditions
43 | 345 | Store contamination shot
43 | 346 | 1kV =
43 | 346 | Store 1KV Resolution Images and Pate image
43 | 346 | 15kV =
43 | 346 | Store 1KV Resolution Images and Pate image
39 | 339 | Actual =
39 | 339 | After conditioning confirm there are no EMServer error messages
39 | 339 | Store Gun monitor trace as Probestability.mn2
42 | 344 | Actual X-Axis
42 | 344 | Store X ripple shots and results
42 | 344 | Actual Y-Axis
42 | 344 | Store Y ripple shots and results
45 | 368 | Analytic Fil I
45 | 368 | Analytic Extractor V:
42 | 11013 | Actual X-Axis
42 | 11013 | Store X ripple shots and results
42 | 11013 | Actual Y-Axis
42 | 11013 | Store Y ripple shots and results
42 | 11047 | Actual X-Axis
42 | 11047 | Store X ripple shots and results
42 | 11047 | Actual Y-Axis
42 | 11047 | Store Y ripple shots and results
43 | 347 | 1kV Inlens
43 | 347 | Store 1kV Customer Image
43 | 347 | 15kV Inlens
43 | 347 | Store 15kV Customer Image
43 | 347 | 1kV Inlens
43 | 347 | Store 1kV Inlens Customer Image
43 | 347 | 15kV Inlens
43 | 347 | Store 15kV Inlens Customer Image
43 | 11353 | 1kV Inlens
43 | 11353 | Store 1kV Inlens Customer Image
43 | 11353 | 15kV Inlens
43 | 11353 | Store 15kV Inlens Customer Image
43 | 11353 | 200V Inlens
43 | 11353 | Store 200V Inlens Customer Image
43 | 11353 | 30KV SE2
43 | 11353 | Store 30kV SE2 Customer Image
43 | 347 | 1kV Inlens
43 | 347 | Store 1kV Inlens Customer Image
43 | 347 | 15kV Inlens
43 | 347 | Store 15kV Inlens Customer Image
5 | 1256 | Fit SCM
43 | 11049 | No EHT Trips or Errors During Test
43 | 11049 | Store Dynamic EHT gun trace
21 | 256 | Store System checker full report
42 | 344 | Actual X-Axis
42 | 344 | Store X ripple shots and results
42 | 344 | Actual Y-Axis
42 | 344 | Store Y ripple shots and results
42 | 11013 | Actual X-Axis
42 | 11013 | Store X ripple shots and results
42 | 11013 | Actual Y-Axis
42 | 11013 | Store Y ripple shots and results
42 | 344 | Actual X-Axis
42 | 344 | Store X ripple shots and results
42 | 344 | Actual Y-Axis
42 | 344 | Store Y ripple shots and results
42 | 11013 | Actual X-Axis
42 | 11013 | Store X ripple shots and results
42 | 11013 | Actual Y-Axis
42 | 11013 | Store Y ripple shots and results
42 | 11047 | Actual X-Axis
42 | 11047 | Store X ripple shots and results
42 | 11047 | Actual Y-Axis
42 | 11047 | Store Y ripple shots and results
646 | 10960 | Install 3DSM Software and Record Software Version
646 | 10960 | Install 3DSM Software and Record Software Version
646 | 10960 | Set Up Image Using One segment of the BSD
646 | 10960 | Run 3DSM and Setup
646 | 10960 | Save 3DSM Project
646 | 10960 | Install 3DSM Software and Record Software Version
646 | 10960 | Set Up Image Using One segment of the BSD
646 | 10960 | Run 3DSM and Setup
646 | 10960 | Save 3DSM Project
646 | 10960 | Store 3DSM Project
646 | 10549 | Install YAG BSD Lightpipe
646 | 10549 | Alignment
646 | 10549 | Select YAG BSD in the Administrator
646 | 10549 | 30KV Function Image
646 | 10549 | Store YAG BSD Functional Image
646 | 10549 | Obtain a Image at 3KV
647 | 10551 | Macro Passed
647 | 10551 | Gun Monitor Saved to Product Folder
647 | 10551 | Store CIV Gun Monitor Files (Parram. Mon file & Mon Export File)
647 | 10552 | Conditions set - 20KV, 1nA, Long Filament Life
647 | 10552 | Gun Monitor Setup and Running
647 | 10552 | Stability Run Passed
647 | 10552 | Gun Monitor Saved to Product Folder
647 | 10552 | Store STAB Gun Monitor Files (Parram. Mon file & Mon Export File)
648 | 10553 | 30KV Res- ≤2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 1KV Res- ≤20nm
648 | 10553 | 200V Res- Not measured
648 | 10553 | Store Standard Resolution Images
648 | 10553 | 30KV Res- ≤2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 1KV Res- ≤20nm
648 | 10553 | 200V Res- Not measured
648 | 10553 | Store Standard Resolution Images
81 | 1587 | Measured value =
648 | 10611 | 30KV Res- ≤2nm
648 | 10611 | 3KV Res- ≤6nm
648 | 10611 | 1KV Res- ≤9nm
648 | 10611 | 200V Res- Not Measured
648 | 10611 | Store Standard LaB6 Resolution Images
648 | 10554 | 2KV Res - 100um VP Aperture- ≤16nm
648 | 10554 | 2KV Res - 400um Easy VP Aperture- ≤16nm
648 | 10554 | Store VP & Easy VP Aperture Resolution Images
648 | 10612 | 2KV Res - 100um VP Apertutre- ≤12nm
648 | 10612 | Store VP Aperture (100u) Resolution Images
648 | 10881 | 2KV Res - 750um Easy VP Aperture- ≤12nm
648 | 10881 | Store Easy VP Aperture (750u) Resolution Images
648 | 10614 | 2KV Res - 750um Easy Align Aperture- ≤12nm
648 | 10614 | Store Easy Align Aperture (750u) Resolution Images
648 | 10875 | 2KV Res - 750um Easy Align Aperture- ≤16nm
648 | 10875 | Store Easy Align Aperture (750u) Resolution Images
648 | 10657 | 2KV - 100um VP Aperture- ≤16nm
648 | 10657 | 2KV - 140um VP Aperture- ≤16nm
648 | 10657 | Store 2KV Resolution Images
24 | 273 | Level Frame
24 | 273 | Check pendulum height (with spacer)
24 | 273 | Set up ‘float’ height
24 | 273 | Check inlet / outlet valves
24 | 273 | Set up docking pins
648 | 10658 | 2KV - 100um VP Aperture- ≤12nm
648 | 10658 | 2KV - 140um VP Aperture- ≤12nm
648 | 10658 | Store 2KV Aperture Resolution Images
648 | 10657 | 2KV - 100um VP Aperture- ≤16nm
648 | 10657 | 2KV - 140um VP Aperture- ≤16nm
648 | 10657 | Store 2KV Resolution Images
648 | 10657 | 2KV - 100um VP Aperture- ≤16nm
648 | 10657 | 2KV - 140um VP Aperture- ≤16nm
648 | 10657 | Store 2KV Aperture Resolution Images
649 | 10555 | BSD Resolution ≤4nm
649 | 10555 | Store BSDResolution Image
649 | 10556 | 30KV Resolution - Easy VP- ≤3.4nm
649 | 10556 | Store VPSE Resolution Image
649 | 11206 | 30KV Res- ≤2nm
649 | 11206 | Store 30KV ( Large ZTR)Resolution Image
649 | 10659 | 30KV Resolution- ≤5.5nm
649 | 10659 | Store VPSE Resolution Image
649 | 10557 | Resolution @ 7KV- ≤9nm
649 | 10557 | Store Beam Deceleration Resolution Image
25 | 278 | Ext V =
25 | 278 | Ext V (Extractor voltage must be no higher than 500V over datasheet value)
650 | 10567 | Install Customer Licences
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer to the Product Folder
650 | 10567 | Store EM Server
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Add Licence to the SmartSEM USB Stick
650 | 10567 | Store Easy Align Aperture (750u) Resolution Images
650 | 10567 | Install Customer Licences
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer to the Product Folder
650 | 10567 | Store EM Server
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Add Licence to the SmartSEM USB Stick
650 | 10567 | Store Easy Align Aperture (750u) Resolution Images
650 | 10567 | Install Customer Licences
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer to the Product Folder
650 | 10567 | Store EM Server
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Sreening
650 | 10567 | Add Licence to the SmartSEM USB Stick
650 | 10567 | Store Easy Align Aperture (750u) Resolution Images
650 | 10567 | Install Customer Licences
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer to the Product Folder
650 | 10567 | Store EM Server
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
650 | 10567 | Add Licence to the SmartSEM USB Stick
650 | 10568 | Calibration File Added to SmartSEM USB Stick
650 | 10568 | Copy Calibration File to the Product Folder
650 | 10568 | Store SmartEDX Calibration File
650 | 10568 | Calibration File Added to SmartSEM USB Stick
650 | 10568 | Copy Calibration File to the Product Folder
650 | 10568 | Store SmartEDX Calibration File
27 | 404 | 100nA-Emmision rotates ACW with I increase
6 | 61 | Safety Earth
6 | 61 | Compressed Air (Set to 0.5MPa)
6 | 61 | PVP Valve
6 | 61 | Power
6 | 61 | Rotary Pump
6 | 61 | Nitrogen (Set to 0.05 MPa)
6 | 61 | Cooling Water
6 | 61 | Set Water Flow switch to 0.9 LPM
6 | 61 | EHT Set Serial Number SN:
6 | 61 | EHT Lens output resistance ( < 2.2kΩ)
22 | 262 | Carry out overnight ‘Full service Bakeout’ and note actual duration.
22 | 262 | Confirm that all Thermocouples are reading.
22 | 262 | Confirm that all Ion pump heater elements and flange heaters are getting hot.
22 | 262 | Gun Vacuum post Bakeout. Ensure column has cooled prior to taking reading.
22 | 262 | After completion of Bakeout Torque the Gun head bolts to 25NM.
| 11380 | S/N =
389 | 114 | Black Level Cal
389 | 114 | C2D Bias Target Check
389 | 114 | Gain Range Check
389 | 114 | Auto Level Check
389 | 114 | Auto Range Check
389 | 114 | Function Image
389 | 114 | Store C2D Functional Image
| 11375 | Position A:
| 11375 | Position B:
| 11375 | Position C:
| 11375 | Position D:
| 11375 | Max. Difference
| 11377 | Has Transzorb been fitted
| 11378 | Continuity (Close Circuit)
| 11378 | Coils to GND (Open Circuit)
| 11379 | Continuity (Close Circuit)
| 11379 | Coils to GND (Open Circuit)
389 | 115
12 | 150 | Check Configuration is correct for $[MaxCurrent]nA Column.
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
403 | 1167 | Download Licence from SharePoint
403 | 1167 | Delete the Licence in SharePoint
389 | 115 | Setup Ethernet Connection
389 | 115 | Install Aztec Software
389 | 115 | Aztec Software Version
389 | 115 | Licence Installation
389 | 115 | Run Aztec Assist+
389 | 115 | Install Detector
389 | 115 | Imaging Setup
389 | 115 | Stage Calibration Setup
389 | 115 | Calibrate Detector
389 | 115 | Save Aztec Sign Off PDF and Calibration Folder
12 | 149 | Confirm Configuration is correct for $[MaxCurrent]nA Column
12 | 148 | Check Configuration is correct for $[MaxCurrent]nA Column.
12 | 151 | Check Configuration is correct for $[MaxCurrent]nA Column.
12 | 149 | Confirm Configuration is correct for $[MaxCurrent]nA Column
12 | 148 | Check Configuration is correct for $[MaxCurrent]nA Column.
12 | 150 | Check Configuration is correct for $[MaxCurrent]nA Column.
12 | 151 | Check Configuration is correct for $[MaxCurrent]nA Column.
13 | 168 | Driver A – Liner (>1000MΩ)
13 | 168 | Driver B – Liner (>1000MΩ)
13 | 168 | Driver A – Driver B (>1000MΩ)
13 | 171 | Driver A – Liner (>1000MΩ)
13 | 171 | Driver B – Liner (>1000MΩ)
13 | 171 | Driver A – Driver B (>1000MΩ)
40 | 1552 | Collector Working
40 | 1552 | Scintillator Working
25 | 278 | Ext V =
25 | 278 | Measured Speciment Current =
6 | 72 | Routed though Plinth secured on top of Plinth
4 | 11389 | Fit chamber pumping pipe
4 | 11389 | Fit chamber to damper
4 | 11389 | Fit chamber bolts
4 | 11389 | Fit earthquake bolts
652 | 10621 | Check LH Bracket
652 | 10621 | Check RH Bracket
652 | 10621 | Check Rear Bracket
652 | 10621 | Check Front Bracket
5 | 1257 | Fit SCM
707 | 11156 | Ensure size of hole is 20um (+/- 2um)
707 | 11156 | Use LR conditions (1024x768, N=1)
707 | 11156 | Save to TIFF as Multi_TS_20um
707 | 11156 | Inspection Passed and Images Saved
707 | 11158 | Ensure size of hole is 30um (+/- 2um)
707 | 11158 | Use LR conditions (1024x768, N=1)
707 | 11158 | Save to TIFF as Multi_TS_30um
707 | 11158 | Inspection Passed and Images Saved
707 | 11159 | Ensure size of hole is 300um (+/- 10um)
707 | 11159 | Use LR conditions (1024x768, N=1)
707 | 11159 | Save to TIFF as Multi_TS_300um
707 | 11159 | Inspection Passed and Images Saved
707 | 11160 | Ensure size of hole is 120um (+/- 3um)
707 | 11160 | Use LR conditions (1024x768, N=1)
707 | 11160 | Save to TIFF as Multi_TS_120um
707 | 11160 | Inspection Passed and Images Saved
707 | 11161 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11161 | Use HR conditions (6144x4608, N=5)
707 | 11161 | Save to TIFF as Multi_US_HRQA
707 | 11161 | Inspection Passed and Images Saved
707 | 11162 | (There is no specific checks not regarding the search for particles)
707 | 11162 | Use HR conditions (6144x4608, N=5)
707 | 11162 | Save to TIFF as Anode_TS_40um
707 | 11162 | Inspection Passed and Images Saved
707 | 11163 | (There is no specific checks not regarding the search for particles)
707 | 11163 | Use LR conditions (1024x768, N=1)
707 | 11163 | Save to TIFF as Anode_TS_Bore1
707 | 11163 | Inspection Passed and Images Saved
707 | 11164 | (There is no specific checks not regarding the search for particles)
707 | 11164 | Use LR conditions (1024x768, N=1)
707 | 11164 | Save to TIFF as Anode_TS_Bore2
707 | 11164 | Inspection Passed and Images Saved
707 | 11165 | Ensure size of hole is 40um (+/- 2um)
707 | 11165 | Use LR conditions (1024x768, N=1)
707 | 11165 | Save to TIFF as Anode_TS_40um
707 | 11165 | Inspection Passed and Images Saved
707 | 11166 | Ensure size of hole is 55um (+/- 2um
707 | 11166 | Use LR conditions (1024x768, N=1)
707 | 11166 | Save to TIFF as Anode_TS_55um
707 | 11166 | Inspection Passed and Images Saved
707 | 11167 | Ensure size of hole is 90um (+/- 2um)
707 | 11167 | Use LR conditions (1024x768, N=1)
707 | 11167 | Save to TIFF as Anode_TS_90um
707 | 11167 | Inspection Passed and Images Saved
707 | 11187 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11187 | Use HR conditions (6144x4608, N=5)
707 | 11187 | Save to TIFF as Multi_TS_HRQA
707 | 11187 | Inspection Passed and Images Saved
707 | 11166 | Ensure size of hole is 55um (+/- 2um
707 | 11166 | Use LR conditions (1024x768, N=1)
707 | 11166 | Save to TIFF as Anode_TS_55um
707 | 11166 | Inspection Passed and Images Saved
707 | 11167 | Ensure size of hole is 90um (+/- 5um)
707 | 11167 | Use LR conditions (1024x768, N=1)
707 | 11167 | Save to TIFF as Anode_TS_90um
707 | 11167 | Inspection Passed and Images Saved
707 | 11168 | Ensure size of hole is 150um (+/- 5um)
707 | 11168 | Use LR conditions (1024x768, N=1)
707 | 11168 | Save to TIFF as Anode_TS_150um
707 | 11168 | Inspection Passed and Images Saved
707 | 11169 | (There is no specific checks not regarding the search for particles)
707 | 11169 | Use HR conditions (6144x4608, N=5)
707 | 11169 | Save to TIFF as Anode_US_HRQA
707 | 11169 | Inspection Passed and Images Saved
707 | 11170 | (Check surface finish on the inner bore edge is free from marks)
707 | 11170 | Use HR conditions (6144x4608, N=5)
707 | 11170 | Save to TIFF as VP_TS_HRQA
707 | 11170 | Inspection Passed and Images Saved
707 | 11171 | (Check surface finish on the inner bore edge is free from marks)
707 | 11171 | Use LR conditions (1024x768, N=1))
707 | 11171 | Save to TIFF as VP_TS_Bore
707 | 11171 | Inspection Passed and Images Saved
707 | 11172 | (Check there is no engraved ring around the aperture, this should be on the underside.
707 | 11172 | Use HR conditions (6144x4608, N=5)
707 | 11172 | Save to TIFF as SF_TS_HRQA
707 | 11172 | Inspection Passed and Images Saved
707 | 11173 | Ensure size of hole is 35um (+/- 2 um)
707 | 11173 | Use HR conditions (6144x4608, N=5)
707 | 11173 | Save to TIFF as SF_TS_35um
707 | 11173 | Inspection Passed and Images Saved
707 | 11174 | Ensure size of hole is 750um (+/- 15um)
707 | 11174 | Use HR conditions (6144x4608, N=5)
707 | 11174 | Save to TIFF as Plat_TS_750um_HRQA
707 | 11174 | Inspection Passed and Images Saved
707 | 11185 | Hold aperture holder, look for holes on the sides.No particles are to be observed
707 | 11186 | Hold aperture holder, look for holes on the sides.No particles are to be observed
708 | 11236 | Hold aperture holder, look for holes on the sides.No particles are to be observed
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
25 | 279 | Fil I
25 | 279 | Ext V
708 | 11238 | Make sure there are no cut-outs on the edge of the anode. No particles are to be observed.
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
708 | 11241 | Make sure there are no cut-outs on the edge of the anode. No particles are to be observed.
25 | 275 | Are Gun Head clamping screws secure?
25 | 275 | Are Condensor clamping screws secure?
25 | 275 | Save Emission image @1Kv (cond off)
25 | 275 | Save Emission image @10Kv (cond off)
25 | 276 | Actual
707 | 11187 | Select PC
707 | 11187 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11187 | Use HR conditions (6144x4608, N=5)
707 | 11187 | Save to TIFF as Multi_TS_HRQA
707 | 11187 | Inspection Passed and Images Saved
707 | 11156 | Select PC
707 | 11156 | Ensure size of hole is 20um (+/- 2um)
707 | 11156 | Use LR conditions (1024x768, N=1)
707 | 11156 | Save to TIFF as Multi_TS_20um
707 | 11156 | Inspection Passed and Images Saved
707 | 11185 | Select PC
707 | 11185 | Hold aperture holder, look for holes on the sides.No particles are to be observed
707 | 11159 | Select PC
707 | 11159 | Ensure size of hole is 300um (+/- 10um)
707 | 11159 | Use LR conditions (1024x768, N=1)
707 | 11159 | Save to TIFF as Multi_TS_300um
707 | 11159 | Inspection Passed and Images Saved
707 | 11160 | Select PC
707 | 11160 | Ensure size of hole is 120um (+/- 3um)
707 | 11160 | Use LR conditions (1024x768, N=1)
707 | 11160 | Save to TIFF as Multi_TS_120um
707 | 11160 | Inspection Passed and Images Saved
707 | 11161 | Select PC
707 | 11161 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11161 | Use HR conditions (6144x4608, N=5)
707 | 11161 | Save to TIFF as Multi_US_HRQA
707 | 11161 | Inspection Passed and Images Saved
707 | 11187 | Select PC
707 | 11187 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11187 | Use HR conditions (6144x4608, N=5)
707 | 11187 | Save to TIFF as Multi_TS_HRQA
707 | 11187 | Inspection Passed and Images Saved
707 | 11252 | Select PC
707 | 11252 | Ensure size of hole is 30um (+/- 2um)
707 | 11252 | Use LR conditions (1024x768, N=1)
707 | 11252 | Use LR conditions (1024x768, N=1)
707 | 11252 | Inspection Passed and Images Saved
707 | 11163 | Select PC
707 | 11163 | (There is no specific checks not regarding the search for particles)
707 | 11163 | Use LR conditions (1024x768, N=1)
707 | 11163 | Save to TIFF as Anode_TS_Bore1
707 | 11163 | Inspection Passed and Images Saved
707 | 11164 | Select PC
707 | 11164 | (There is no specific checks not regarding the search for particles)
707 | 11164 | Use LR conditions (1024x768, N=1)
707 | 11164 | Save to TIFF as Anode_TS_Bore2
707 | 11164 | Inspection Passed and Images Saved
707 | 11165 | Select PC
707 | 11165 | Ensure size of hole is 40um (+/- 2um)
707 | 11165 | Use LR conditions (1024x768, N=1)
707 | 11165 | Save to TIFF as Anode_TS_40um
707 | 11165 | Inspection Passed and Images Saved
707 | 11166 | Select PC
707 | 11166 | Ensure size of hole is 55um (+/- 2um
707 | 11166 | Use LR conditions (1024x768, N=1)
707 | 11166 | Save to TIFF as Anode_TS_55um
707 | 11166 | Inspection Passed and Images Saved
707 | 11167 | Select PC
707 | 11167 | Ensure size of hole is 90um (+/- 5um)
707 | 11167 | Use LR conditions (1024x768, N=1)
707 | 11167 | Save to TIFF as Anode_TS_90um
707 | 11167 | Inspection Passed and Images Saved
707 | 11168 | Select PC
707 | 11168 | Ensure size of hole is 150um (+/- 5um)
707 | 11168 | Use LR conditions (1024x768, N=1)
707 | 11168 | Save to TIFF as Anode_TS_150um
707 | 11168 | Inspection Passed and Images Saved
707 | 11169 | Select PC
707 | 11169 | (There is no specific checks not regarding the search for particles)
707 | 11169 | Use HR conditions (6144x4608, N=5)
707 | 11169 | Save to TIFF as Anode_US_HRQA
707 | 11169 | Inspection Passed and Images Saved
707 | 11170 | Select PC
707 | 11170 | (Check surface finish on the inner bore edge is free from marks)
707 | 11170 | Use HR conditions (6144x4608, N=5)
707 | 11170 | Save to TIFF as VP_TS_HRQA
707 | 11170 | Inspection Passed and Images Saved
707 | 11171 | Select PC
707 | 11171 | (Check surface finish on the inner bore edge is free from marks)
707 | 11171 | Use LR conditions (1024x768, N=1))
707 | 11171 | Save to TIFF as VP_TS_Bore
707 | 11171 | Inspection Passed and Images Saved
707 | 11172 | Select PC
707 | 11172 | (Check there is no engraved ring around the aperture, this should be on the underside.
707 | 11172 | Use HR conditions (6144x4608, N=5)
707 | 11172 | Save to TIFF as SF_TS_HRQA
707 | 11172 | Inspection Passed and Images Saved
707 | 11173 | Select PC
707 | 11173 | Ensure size of hole is 35um (+/- 2 um)
707 | 11173 | Use HR conditions (6144x4608, N=5)
707 | 11173 | Save to TIFF as SF_TS_35um
707 | 11173 | Inspection Passed and Images Saved
707 | 11174 | Select PC
707 | 11174 | Ensure size of hole is 750um (+/- 15um)
707 | 11174 | Use HR conditions (6144x4608, N=5)
707 | 11174 | Save to TIFF as Plat_TS_750um_HRQA
707 | 11174 | Inspection Passed and Images Saved
26 | 280 | Use Insight Column Tool. Check for cut off and charging particles.
707 | 11186 | Select PC
707 | 11186 | Hold aperture holder, look for holes on the sides.No particles are to be observed
707 | 11262 | Select PC
707 | 11262 | (There is no specific checks not regarding the search for particles)
707 | 11262 | Use HR conditions (6144x4608, N=5)
707 | 11262 | Save to TIFF as Anode_TS_HRQA
707 | 11262 | Inspection Passed and Images Saved
708 | 11236 | Select PC
708 | 11236 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed.
708 | 11237 | Select PC
708 | 11237 | Hold Anode, look for a ring of Holes between the screw holes and aperture. No particles are to be observed
25 | 278 | HV = Condenser Off 220pA
25 | 278 | VP = Condenser On 300Pa
25 | 278 | Actual Spec I
44 | 354 | Check Installed Licences Against Sales Order
708 | 11238 | Select PC
708 | 11238 | Make sure there are no cut-outs on the edge of the anode. No particles are to be observed.
708 | 11239 | Select PC
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
708 | 11241 | Select PC
708 | 11241 | Make sure there are no cut-outs on the edge of the anode. No particles are to be observed.
708 | 11263 | Select PC
708 | 11263 | Make sure there are 3 cut-outs on the edge of the anode. No particles are to be observed
708 | 11264 | Select PC
708 | 11264 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be observed.
708 | 11265 | Select PC
708 | 11265 | Make sure there are 3 cut-outs on the edge of the anode. No particles are to be observed.
707 | 11156 | Ensure size of hole is 20um (+/- 2um)
707 | 11156 | Use LR conditions (1024x768, N=1)
707 | 11156 | Save to TIFF as Multi_TS_20um
707 | 11156 | Inspection Passed and Images Saved
44 | 355 | Attach label to chamber
707 | 11159 | Ensure size of hole is 300um (+/- 10um)
707 | 11159 | Use LR conditions (1024x768, N=1)
707 | 11159 | Save to TIFF as Multi_TS_300um
707 | 11159 | Inspection Passed and Images Saved
707 | 11160 | Ensure size of hole is 120um (+/- 3um)
707 | 11160 | Use LR conditions (1024x768, N=1)
707 | 11160 | Save to TIFF as Multi_TS_120um
707 | 11160 | Inspection Passed and Images Saved
707 | 11161 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11161 | Use HR conditions (6144x4608, N=5)
707 | 11161 | Save to TIFF as Multi_US_HRQA
707 | 11161 | Inspection Passed and Images Saved
707 | 11187 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11187 | Use HR conditions (6144x4608, N=5)
707 | 11187 | Save to TIFF as Multi_TS_HRQA
707 | 11187 | Inspection Passed and Images Saved
707 | 11252 | Ensure size of hole is 30um (+/- 2um)
707 | 11252 | Use LR conditions (1024x768, N=1)
707 | 11252 | Use LR conditions (1024x768, N=1)
707 | 11252 | Inspection Passed and Images Saved
707 | 11277 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11277 | Use HR conditions (6144x4608, N=5)
707 | 11277 | Save to TIFF as Multi_TS_HRQA
707 | 11277 | Inspection Passed and Images Saved
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
23 | 269 | Run macro 'ESB Runup 3h pumpen & 9h hochfahren.MLF'
707 | 11283 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11283 | Use HR conditions (6144x4608, N=5)
707 | 11283 | Save to TIFF as Multi_TS_HRQA
707 | 11283 | Inspection Passed and Images Saved
707 | 11284 | Check there is no engraved ring around the cluster of apertures, this should be on the underside.
707 | 11284 | Use HR conditions (6144x4608, N=5)
707 | 11284 | Save to TIFF as Multi_TS_HRQA
707 | 11284 | Inspection Passed and Images Saved
707 | 11285 | Ensure size of hole is 30um (+/- 2um)
707 | 11285 | Use LR conditions (1024x768, N=1)
707 | 11285 | Use LR conditions (1024x768, N=1)
707 | 11285 | Inspection Passed and Images Saved
707 | 11286 | Ensure size of hole is 20um (+/- 2um)
707 | 11286 | Use LR conditions (1024x768, N=1)
707 | 11286 | Save to TIFF as Multi_TS_20um
707 | 11286 | Inspection Passed and Images Saved
707 | 11287 | Ensure size of hole is 120um (+/- 3um)
707 | 11287 | Use LR conditions (1024x768, N=1)
707 | 11287 | Save to TIFF as Multi_TS_120um
707 | 11287 | Inspection Passed and Images Saved
707 | 11288 | Ensure size of hole is 300um (+/- 10um)
707 | 11288 | Use LR conditions (1024x768, N=1)
707 | 11288 | Save to TIFF as Multi_TS_300um
707 | 11288 | Inspection Passed and Images Saved
29 | 295 | Check for Drift, Charging, ESB issues.
707 | 11289 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11289 | Use HR conditions (6144x4608, N=5)
707 | 11289 | Save to TIFF as Multi_US_HRQA
707 | 11289 | Inspection Passed and Images Saved
707 | 11290 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11290 | Use HR conditions (6144x4608, N=5)
707 | 11290 | Save to TIFF as Multi_US_HRQA
707 | 11290 | Inspection Passed and Images Saved
707 | 11308 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11308 | Use HR conditions (6144x4608, N=5)
707 | 11308 | Save to TIFF as Multi_US_HRQA
707 | 11308 | Inspection Passed and Images Saved
26 | 281 | Type in measured value (Limit <=30um)
26 | 282 | Type in measured value (limit: <=15um)
707 | 11185 | Hold aperture holder, look for holes on the sides.No particles are to be observed
26 | 396 | Type in measured value (Limit: <= 50um)
27 | 284 | 20nA-Emmision rotates CW with I increase
27 | 285 | No image rotation with WD change
27 | 287 | Offset Angle
27 | 289 | C2 Calibration Value
27 | 289 | C3 Calibration Value
20 | 252 | PC S/N
20 | 252 | Hard disk Image
20 | 252 | PC/Model Number
20 | 252 | Record Operating System
20 | 252 | Activate Windows
20 | 252 | Record Windows Product Key
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Install SmartSEM Service Pack (If Applicable)
20 | 252 | Install SmartSEM Hotfix (If Applicable)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Copy Macro Files to C:/support tools
20 | 252 | Approved Datazone Loaded
20 | 254 | Language
21 | 255 | Air (Pressure = 6 Bar)
21 | 255 | Water
21 | 255 | Mains
21 | 255 | Safety Earth
21 | 255 | N2
21 | 256 | Save Full System Checker report
21 | 257 | In Factory Tab: Omit Dongle Check
21 | 257 | In Factory Tab: Enter S/N
21 | 257 | In License Tab: Set All Licenses
21 | 258 | Confirm there are no Errors in EMserver(Apart from Dongle check and Stage Init)
21 | 259 | Check camera Focus and Rotation
21 | 259 | Eliminate camera image tilt
21 | 385
21 | 385 | Select correct System type (HV / VP)
21 | 385 | If Pressure Stage is fitted Tick box
21 | 387
21 | 389 | Tick option 'Low KV Objective Cap'
22 | 260 | Column traveller completed?
22 | 260 | Plinth traveller completed?
22 | 260 | System Configuration correct?
22 | 260 | Build quality satisfactory?
22 | 264 | Tip No.
22 | 265 | Set +/-21V Supplies to +/- 22.5V
28 | 290 | Gun running to Datasheet
28 | 290 | Gun Vacuum acceptable (EHT on/off)
28 | 290 | Nominal Ext V acceptable
28 | 291 | Mechanical Alignment ok
28 | 291 | Column clamping secure
28 | 292 | Cut off ok
28 | 292 | No Charging particles
28 | 293 | ESB concentricity in spec
28 | 293 | ESB-Inlens concenrticity in spec
28 | 294 | Run short image jump macro and analyse images
28 | 409 | Gun X
28 | 409 | Gun Y
28 | 409 | Aper X
28 | 409 | Aper Y
30 | 305 | Holder Height
30 | 305 | Holder Rotation
30 | 305 | Select Customer Sample Holders
30 | 305 | Confirm Holder touch alarm works (Alarm tone and error message)
32 | 311 | Actual:
32 | 312 | Actual @ 1kV
32 | 312 | Actual @ 20kV
32 | 314 | Normal mode
32 | 314 | High Current mode
33 | 316 | Actual for 1kV
33 | 316 | Actual for 15kV
33 | 317 | No excessive stigs etc.
34 | 318 | Actual
37 | 329 | Stig: 30μm, All EHT’s ≤ 20%
37 | 329 | Stig: Off Axis Aper, All EHT’s ≤ 30%
37 | 329 | Aper/Gun: 30μm, All EHT’s ≤ 60%
37 | 331 | Stig: 30μm, All EHT’s ≤ 20%
37 | 331 | Stig: Off Axis Aper, All EHT’s ≤ 30%
37 | 331 | Aper/Gun: 30μm, All EHT’s ≤ 50%
37 | 415 | Operation 5310 submitted
37 | 415 | Operation 5350 submitted
38 | 333 | 30kV Max Ext Voltage =
38 | 333 | Complete X-Ray certificate
38 | 334 | Fil I:
38 | 334 | Ext V:
38 | 335 | Specimen current at 20kV
38 | 337 | Actual (Limit <= 300nm):
38 | 422 | Specimen current at 20kV
39 | 339 | Actual:
39 | 339 | After conditioning confirm there are no EMServer error messages
39 | 339 | Save Gun Monitor as Probestability.MN2
40 | 340 | 5 – 30Pa Variable Pressure
40 | 340 | VP image taken (30kV, 200kX)
40 | 340 | Ensure Penning fitted is Type ‘PKR’
40 | 340 | No arcing of collector up to 290v @ 30Pa
40 | 341 | Test all accessory options as per customer order and record under options fitted
40 | 342 | Carry out any relevant additional tests – Refer to Man Eng if unsure
46 | 369 | Create Directory C:\Support\Manuals
46 | 369 | Create Shortcut to above Directory on Desktops, called ‘Carl Zeiss User Documentation’.
41 | 343 | Before Cladding the system check that all column connections aresecurely fastened.
41 | 343 | Strip down Aluminium support frame and clean tapped holes. Re-assemble frame and secure screws with Loctite 222.
41 | 343 | Ensure Cladding geometry is correct and all parts aligned.
42 | 344 | Actual X-Axis
42 | 344 | Actual Y-Axis
43 | 346 | 1kV
43 | 346 | 15kV
43 | 347 | 1kV
43 | 347 | 15kV
43 | 410 | 1kV ESB
43 | 417 | Operation 7000 sumitted
43 | 417 | Operation 8000 submitted
44 | 348 | Chamber Door Interlock
44 | 348 | Service Switch
44 | 348 | Lens De-activation
44 | 348 | Detector De-activation
44 | 348 | EHT + IPPSU De-activation
44 | 348 | A/L I/L test (if Fitted)
44 | 348 | EMO I/L test (if Fitted)
44 | 349 | Carry out safety tests using the Secutest S111 pat Tester.
44 | 349 | Fill out completely the Product Safety Tests form in Section 13.
44 | 350 | Actual
44 | 351 | Actual
44 | 352 | Limit < 5 min.
44 | 352 | Actual:
44 | 353 | EMserver.mdb
44 | 353 | Customer Images
44 | 353 | License File
44 | 356 | Check System And Dongle S/No. Match In Service Panel
44 | 356 | Un-tick Omit Dongle In Factory Panel
44 | 358 | All verification & Safety tests done and within threshold limits defined
44 | 358 | Sales Order is latest version
44 | 358 | All items on Sales order accounted for
44 | 359 | Remove all Test Macro Shortcuts
44 | 359 | Remove all Test Program Shortcuts
44 | 359 | Set up Desktop icons (both accounts)
44 | 359 | Remove Proxy Server
44 | 359 | Reset WSUS Settings
45 | 362 | Em Server.mdb
45 | 362 | Em Server.txt Log
45 | 362 | Optiprobe.txt Log
45 | 362 | Smart Backup.log
45 | 362 | Export Tablefiller.MERLINCDB
45 | 362 | L500.ini
45 | 362 | License file.xml
45 | 362 | Registry file.reg
45 | 362 | All Images
45 | 362 | Probe Stability.MN2
45 | 362 | CalibSc.dat
45 | 362 | System Checker Full report
45 | 363 | Em Server.mdb
45 | 363 | License File
45 | 363 | Saved Registry
45 | 363 | L500.ini
45 | 365 | Language:
45 | 366 | Nominal Fil I:
45 | 366 | Nominal extractor V:
45 | 367 | High Res Fil I:
45 | 367 | High Res Extractor V:
46 | 371 | X-Ray Certificate
46 | 371 | Quality Certificate
46 | 371 | Quality Feedback Form
46 | 371 | DENKA Data Sheet
46 | 371 | Options Removed packing List
46 | 371 | Options Retained packing List
46 | 371 | Standard Packing List
46 | 371 | PC Inferface assignment
46 | 371 | Declaration of conformity
46 | 371 | Installation report
46 | 371 | Acceptance test record
46 | 371 | Licence Generator report
46 | 419 | C:\Support\Manuals
46 | 419 | Customer USB documentation folder
20 | 253 | Enter the FIB Extension number using EEPromUpdater.exe
36 | 325 | Configuration dependent:
36 | 325 | 40 – 48nA
36 | 325 | 100 – 120nA
36 | 325 | 300 – 360nA
36 | 327 | Extractor current should not change by > 4ua between 10kV & 30kV.
26 | 280 | Use Insight Column Tool. Check for cut off and charging particles.
40 | 342 | Carry out any relevant additional tests – Refer to Man Eng if unsure.
22 | 11032 | Joystick Angle Set
| 11391 | In-Lens and ESB
| 11391 | EO Dynamic PCB
| 11391 | EO Static PCB
| 11391 | Scan Gen PCB
718 | 11392 | In-Lens and ESB
718 | 11392 | EO Dynamic PCB
718 | 11396 | Heater & Thermocouple cables plugged into AC unit
718 | 11396 | Column mounting bolts all secured
718 | 11396 | IGP bracket screws all secured
718 | 11396 | All column cabling cable tied neatly
21 | 12567 | Enter the FIB Extension number using EEPromUpdater.exe
25 | 11410 | With Cond. Off Max Aper. X & Y < 45
25 | 11410 | Align offset ange set to -110 degrees.
25 | 11410 | Align offset angle set to -110 degrees.
656 | 10670 | EO PCB S/No
656 | 10670 | Confirm Material No. is $[SigmaEOMaterial]
25 | 11410 | Align sdkgfjh ds gfd offset sdf jkfsd kjfkjangle set to -110 degrees.
389 | 3314 | Fit Detector to Chamber
389 | 3314 | Install EDS PCBs into PC
389 | 3314 | Install Cables
389 | 3314 | Install EDAX Software
389 | 3314 | Disable Fast Startup
389 | 3314 | Install Smart API Client from SmartSEM USB stick, setup, configure and Run API test
389 | 3314 | Configure imaging – Run SGShell
389 | 3314 | Calibrate Detector – Run eDPPShell
389 | 3314 | Save calibration data on exit
389 | 3314 | Run SmartEDX software – check Spectrum, Mapping and Line Scan
389 | 3314 | Save aquired Mapping and Spectrum Images to Image Folder
389 | 3314 | Save Calibration Data to SmartSem USB stick and in the Image folder.
389 | 3314
388 | 11426 | 15kv 0.7nm
388 | 11426 | 15KV result
388 | 11426 | 1kv 1.2nm
388 | 11426 | 1KV result
388 | 11426 | 500V 1.9nm
388 | 11426 | 500V result
285 | 1222 | Installation of Zen licenece
388 | 11426 | 15kv 0.7nm
388 | 11426 | 15KV result
388 | 11426 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 1kv 1.2nm
388 | 11426 | 1KV result
388 | 11426 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 500V 1.9nm
388 | 11426 | 500V result
388 | 11426 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
| 12636 | Wert =
| 12666 | Wert =
| 13507 | Wert =
860 | 13707 | Wert C1 Cal =
860 | 13707 | Wert C2 Cal =
| 12986 | Wert =
| 12987 | Wert =
| 12983 | Wert =
25 | 11410 | Align offset sdf jkfsd kjfkjangle set to -110 degrees.
25 | 11410 | Align sjfdghkj dfh kfoffset angle set to -110 degrees.
5 | 98 | Configuration Check
5 | 98 | Configuration Check
5 | 98 | Torque done up to 25Nm
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
21 | 255 | Air (Pressure = 6 Bar)
21 | 255 | Water
21 | 255 | Confirm that the Water Flow Switch is set to 0.9 LPM
21 | 255 | Mains
21 | 255 | Safety Earth
21 | 255 | N2
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detecto
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346
389 | 115
389 | 115 | Setup Ethernet Connection
389 | 115 | Install Aztec Software
389 | 115 | Aztec Software Version
389 | 115 | Licence Installation
389 | 115 | Run Aztec Assist+
389 | 115 | Install Detector
389 | 115 | Imaging Setup
389 | 115 | Stage Calibration Setup
389 | 115 | Calibrate Detector
389 | 115 | Setup Ethernet Connection
389 | 115 | Install Aztec Software
389 | 115 | Aztec Software Version
389 | 115 | Licence Installation
389 | 115 | Run Aztec Assist+
389 | 115 | Install Detector
389 | 115 | Imaging Setup
389 | 115 | Stage Calibration Setup
389 | 115 | Calibrate Detector
389 | 115 | Save Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Store C2D Functional Image
21 | 11427 | Air (Pressure = 6 Bar)
21 | 11427 | Water
21 | 11427 | Check Water Flow switch set to 0.9 LPM
21 | 11427 | Mains
21 | 11427 | Safety Earth
21 | 11427 | N2
21 | 11427 | Remove Blank and fit Test Tool Stage
21 | 11427 | Fit Test Tool Stage Controller and cable up
21 | 11427 | Mount Test Tool SE2 Detector onto chamber
21 | 11427 | Mount PM Tube (with Immersion oil) onto SE2 Detector
21 | 11427 | Mount SCM to port on Stage Door
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Setup Ethernet Connection
646 | 11346 | Install Aztec Software
646 | 11346 | Aztec Software Version
646 | 11346 | Licence Installation
646 | 11346 | Run Aztec Assist+
646 | 11346 | Install Detector
646 | 11346 | Imaging Setup
646 | 11346 | Stage Calibration Setup
646 | 11346 | Calibrate Detector
646 | 11346 | Save Aztec Sign Off PDF and Calibration Folder
646 | 11346 | Store Aztec Sign Off PDF and Calibration Folder
2 | 13 | Confirm EO PCB PN is '346001-9110-315'
2 | 1153 | Confirm EO PCB PN is '346001-9110-305'
2 | 1153 | Fit EO Dynamic PCB
2 | 1154 | Confirm EO Static PCB PN is '346001-9100-400'
2 | 1154 | Fit EO Static PCB
648 | 10875 | 2KV Res - 750um Easy Align Aperture- ≤16nm
648 | 10875 | Store Easy Align Aperture (750u) Resolution Images
12 | 1166 | Anode + Extractor measuring jig S/N =
12 | 1166 | Gun head measuring jig S/N =
12 | 1166 | Depth gauge S/N =
9 | 102 | Aperture cleaned using IPA and 1mm brush
9 | 102 | Aperture run through ultra-sonic tank (in upright position)
9 | 102 | Aperture inspected and confimred clean. Eng 1 Initials:
9 | 102 | Aperture inspected and confimred clean. Eng 2 Initials:
9 | 1652 | S/N =
6 | 1545 | Confirm test unit fails using the High voltage probe, Hi-Pot Test Set (356180-9217-000) and shorting test box connected (354933-9205-000).
6 | 1547 | PAT Checkbox Calibration Date
6 | 1547 | Earth Bond Test TP1 (0.1Ω - 0.09Ω)
6 | 1547 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
6 | 1547 | Insulation Test (2.0MΩ - 1.9MΩ)
6 | 1547 | Earth Leakage Test (0.9mA - 0.7mA)
6 | 1545 | Confirm test unit fails using the High voltage probe, Hi-Pot Test Set (356180-9217-000) and shorting test box connected (354933-9205-000).
6 | 1547 | PAT Checkbox Calibration Date
6 | 1547 | Earth Bond Test TP1 (0.1Ω - 0.09Ω)
6 | 1547 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
6 | 1547 | Insulation Test (2.0MΩ - 1.9MΩ)
6 | 1547 | Earth Leakage Test (0.9mA - 0.7mA)
6 | 1253 | Confirm all vacuum modules are Green in Vac Test
10 | 1165 | S/N =
6 | 1547 | PAT Checkbox Calibration Date
6 | 1547 | Earth Bond Test TP1 (0.1Ω - 0.09Ω)
6 | 1547 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
6 | 1547 | Insulation Test (2.0MΩ - 1.9MΩ)
6 | 1547 | Earth Leakage Test (0.9mA - 0.7mA)
6 | 1545 | Confirm test unit fails using the High voltage probe, Hi-Pot Test Set (356180-9217-000) and shorting test box connected (354933-9205-000).
6 | 1545 | Confirm test unit fails using the High voltage probe, Hi-Pot Test Set (356180-9217-000) and shorting test box connected (354933-9205-000).
6 | 1547 | PAT Checkbox Calibration Date
6 | 1547 | Earth Bond Test TP1 (0.1Ω - 0.09Ω)
6 | 1547 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
6 | 1547 | Insulation Test (2.0MΩ - 1.9MΩ)
6 | 1547 | Earth Leakage Test (0.9mA - 0.7mA)
6 | 1545 | Confirm test unit fails using the High voltage probe, Hi-Pot Test Set (356180-9217-000) and shorting test box connected (354933-9205-000).
6 | 1547 | PAT Checkbox Calibration Date
6 | 1547 | Earth Bond Test TP1 (0.1Ω - 0.09Ω)
6 | 1547 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
6 | 1547 | Insulation Test (2.0MΩ - 1.9MΩ)
6 | 1547 | Earth Leakage Test (0.9mA - 0.7mA)
667 | 11411 | Confirm test unit fails with shorting test box connected (354933-9205-000).
20 | 952 | Please print 'Declaration of Conformity' from URL link
6 | 1547 | PAT Checkbox Calibration Date
6 | 1547 | Earth Bond Test TP1 (0.1Ω - 0.09Ω)
6 | 1547 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
6 | 1547 | Insulation Test (2.0MΩ - 1.9MΩ)
6 | 1547 | Earth Leakage Test (0.9mA - 0.7mA)
6 | 1547 | PAT Checkbox Calibration Date
6 | 1547 | Earth Bond Test TP1 (0.1Ω - 0.09Ω)
6 | 1547 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
6 | 1547 | Insulation Test (2.0MΩ - 1.9MΩ)
6 | 1547 | Earth Leakage Test (0.9mA - 0.7mA)
647 | 10551 | Macro Passed
647 | 10551 | Gun Monitor Saved to Product Folder
647 | 10551 | Store CIV Gun Monitor Files (Parram. Mon file & Mon Export File)
44 | 354 | Check Installed Licences Against Sales Order
24 | 273 | Level Frame
24 | 273 | Check pendulum height (with spacer)
24 | 273 | Set up ‘float’ height
24 | 273 | Check inlet / outlet valves
24 | 273 | Set up docking pins
20 | 11255 | Check SCS Folder/Sales Order
20 | 11255 | SCS Number (If applicable)
667 | 11412 | PAT Checkbox Calibration Date
667 | 11412 | Earth Bond Test TP1 (0.1Ω - 0.07Ω)
667 | 11412 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
667 | 11412 | Insulation Test (2.0MΩ - 1.9MΩ)
667 | 11412 | Earth Leakage Test (0.9mA - 0.7mA)
667 | 11411 | Confirm test unit fails with shorting test box connected (354933-9205-000).
667 | 11412 | PAT Checkbox Calibration Date
667 | 11412 | Earth Bond Test TP1 (0.1Ω - 0.07Ω)
667 | 11412 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
667 | 11412 | Insulation Test (2.0MΩ - 1.9MΩ)
667 | 11412 | Earth Leakage Test (0.9mA - 0.7mA)
647 | 10551 | Macro Passed
647 | 10551 | Gun Monitor Saved to Product Folder
647 | 10551 | Store CIV Gun Monitor Files (Parram. Mon file & Mon Export File)
645 | 10761 | Functional Check
644 | 10510 | 500um Lower Aperture Fitted
644 | 10510 | 2nd Rotary Pump is ON in EP Mode
644 | 10510 | 10Pa Chamber Vacuum Achieved
644 | 10510 | 3000Pa Vacuum Achieved (Air)
644 | 10510 | Water Function Checked
644 | 10510 | Approx. 2000Pa Chamber Vacuum Achieved with Water
644 | 10510 | Go To HV, 2nd Rotary Pump is OFF
642 | 10485 | Approx. No. of Hours
650 | 10567 | Obtain the customer Entitlement Certificate
650 | 10567 | Download into the Engineer folder
650 | 10567 | Install the Customer Licences
650 | 10567 | Install Customer Licences
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer.db to the Product Folder
650 | 10567 | Store EM Server.db
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
650 | 10567 | Add Licence to the SmartSEM USB Stick
650 | 10567 | Add EM Server.db to the SmartSEM USB Stick
650 | 10567 | Delete the Production Test Licences
650 | 10567 | Obtain the customer Entitlement Certificate
650 | 10567 | Download into the Engineer folder
650 | 10567 | Install the Customer Licences
650 | 10567 | Install Customer Licences
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer.db to the Product Folder
650 | 10567 | Store EM Server.db
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
650 | 10567 | Add Licence to the SmartSEM USB Stick
650 | 10567 | Add EM Server.db to the SmartSEM USB Stick
650 | 10567 | Delete the Production Test Licences
650 | 10567 | Obtain the customer Entitlement Certificate
650 | 10567 | Download into the Engineer folder
650 | 10567 | Install the Customer Licences
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer.db to the Product Folder
650 | 10567 | Store EM Server.db
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
45 | 365 | Language
390 | 3321 | Delete the Production Test Licences
390 | 3321 | Obtain the customer Entitlement Certificate
390 | 3321 | Download into the Engineer folder
390 | 3321 | Load customer licences and check against order
390 | 3321 | Test customer licenses
704 | 11041 | Save An Entry For Each In Table Editor (Four entries)
37 | 461 | 21k - All Apertures at Normal and High current WD5
37 | 461 | 30kv - Aperture 1 and 7 at Normal and High current at WD5
25 | 704 | Gun X
25 | 704 | Gun Y
25 | 704 | Aper X
25 | 704 | Aper Y
25 | 704 | Gun X
25 | 704 | Gun Y
25 | 704 | Aper X
25 | 704 | Aper Y
45 | 11059 | Copy- 1KV.Tif (Tiff customer image)
45 | 11059 | Copy- 15KV.Tif (Tiff customer image)
45 | 11059 | Copy- 30KV.Tif (Tiff customer image)
45 | 11059 | Copy- 200V.Tif (Tiff customer image)
45 | 11059 | Copy- All Ripple images (Tiff Files)
45 | 11059 | Copy- All Ripple images (JPG Files)
45 | 11059 | Copy- 9876543.xml (XML Smartsem Licence)
45 | 11059 | Copy- Merlin Compact XX-XX S/O (Licence Rep File)
45 | 11059 | Copy- EmServer.accdb (Database)
45 | 11059 | Copy- EmServer.txt (Server Log File)
45 | 11059 | Copy- L500.ini (SmartSEM Config)
45 | 11059 | Copy- xx-xx.reg (Export of Registry File)
45 | 11059 | Copy- Mybackup.smartbackup (SmartSEM Backup)
45 | 11059 | Copy- Logbook (scanned)-PDF
45 | 11059 | Copy- DynamicEHT (mn2 logfile)
45 | 11059 | Copy- Probe Conditioning Stability (mn2 logfile)
45 | 11059 | Copy- 1KV Res.jpg (PATE Result)
45 | 11059 | Copy- 1KV Res.Tif (Tiff Res Image)
45 | 11059 | Copy- 15KV.jpg (PATE Results)
45 | 11059 | Copy- 15KV.Tif (Tiff Res Image)
45 | 11059 | Emission-Condenser OFF image from QA (Tiff Image)
45 | 11059 | Emission-Condenser ON image from QA (Tiff Image)
45 | 11059 | Merlin Compact User Manual (English) (PDF Manual)
45 | 11059 | Merlin Compact User Manual (German) (PDF Manual)
45 | 11059 | Merlin Compact Assembly (PDF Manual)
645 | 10761 | Functional Check
400 | 3403 | Fit foam between column cladding and chamber (356581-5085-000 can be used and modified as needed)
400 | 3403 | Fit LH Bracket
400 | 3403 | Fit RH Bracket
400 | 3403 | Fit Rear Bracket
400 | 3403 | Fit Front Bracket
400 | 3403 | Checked By Name:
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448 , then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Install Smart SEM (See WI 9693) Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Run ASAP Screening Program
20 | 252 | System Specific Datazone Loaded
82 | 1591 | Take and save images at Mag 200, 2kX and 35kX.
82 | 1591 | Mag is within +/- 1% for each.
82 | 1591 | Orthognality is within +/- 0.1 deg.
81 | 11112 | Measured value =
81 | 1169 | Measured value =
81 | 1169 | Measured value =
81 | 1169 | Nominal Extractor Voltage (If >4.2kv please consult MEng)
81 | 1169 | Nominal Extractor Voltage above Datasheet value (If >500v please consult MEng)
650 | 10567 | Delete the Production Test Licences
650 | 10567 | Obtain the customer Entitlement Certificate
650 | 10567 | Download into the Engineer folder
650 | 10567 | Install the Customer Licences
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer.db to the Product Folder
650 | 10567 | Store EM Server.db
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
403 | 1167 | Download Licence from SharePoint
403 | 1167 | Delete the Licence in SharePoint
403 | 1167 | Install the licence using the Zeiss License Management Tool.
403 | 1167 | Download Licence from SharePoint
403 | 1167 | Delete the Licence in SharePoint
403 | 1167 | Install the licence using the Zeiss License Management Tool.
285 | 1168 | Download Licence from SharePoint
285 | 1168 | Delete the Licence in SharePoint
285 | 1168 | Install the licence using the Zeiss License Management Tool.
390 | 3321 | Delete the Production Test Licences
390 | 3321 | Obtain the customer Entitlement Certificate
390 | 3321 | Download into the Engineer folder
390 | 3321 | Load customer licences and check against order
390 | 3321 | Test customer licenses
390 | 3321 | Delete the Production Test Licences
390 | 3321 | Obtain the customer Entitlement Certificate
390 | 3321 | Download into the Engineer folder
390 | 3321 | Load customer licences and check against order
390 | 3321 | Test customer licenses
651 | 669 | The Licence file is present inthe Engineer folder
651 | 669 | The EM Server.db is present in the Engineer folder
285 | 1222 | Installation of Zen licence
403 | 10467 | Install SmartSEM + ASAP Engine
403 | 10467 | Copy the SmartSEM sotware into a Folder on the Desktop. Name it SmaerSEM 8.0 Software
403 | 10467 | SmartSEM Version
403 | 10467 | SmartSEM Service Pack
403 | 10467 | SmartSEM Hotfix
403 | 10467 | ASAP Engine Version
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Set the Time and Date on the PC to current UK
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
285 | 3242 | Select PC type
285 | 3242 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
285 | 3242 | PC S/No.
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Upgrade ASAP Firmware (check Service Pack for latest version) and API
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Set the Time and Date on the PC to current UK
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
403 | 10463 | Set the Time and Date on the PC to current UK
37 | 1170 | Tablefiller successfully completed.
37 | 1170 | Tablefiller successfully completed.
37 | 328 | Tablefiller successfully completed.
285 | 3242 | Select PC type
285 | 3242 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
285 | 3242 | PC S/No.
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Upgrade ASAP Firmware (check Service Pack for latest version) and API
285 | 3242 | Copy the SmartSEM sotware into a Folder on the Desktop. Name it SmartSEM 8.0 Software
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
403 | 10464 | Check That AC Loss Policy Depend On, is Set To Power On
403 | 10464 | Disable Windows Fast Start UP
403 | 10464 | Set the Time and Date on the PC to current UK
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
285 | 3242 | Select PC type
285 | 3242 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
285 | 3242 | PC S/No.
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Upgrade ASAP Firmware (check Service Pack for latest version) and API
285 | 3242 | Copy the SmartSEM sotware into a Folder on the Desktop. Name it SmartSEM 8.0 Software
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Set the Time and Date on the PC to current UK
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Select PC type
285 | 3242 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
285 | 3242 | PC S/No.
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Upgrade ASAP Firmware (check Service Pack for latest version) and API
285 | 3242 | Copy the SmartSEM sotware into a Folder on the Desktop. Name it SmartSEM 8.0 Software
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Set the Time and Date on the PC to current UK
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Select PC type
285 | 3242 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
285 | 3242 | PC S/No.
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Upgrade ASAP Firmware (check Service Pack for latest version) and API
285 | 3242 | Copy the SmartSEM sotware into a Folder on the Desktop. Name it SmartSEM 8.0 Software
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Set the Time and Date on the PC to current UK
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 1168 | Download Licence from SharePoint
285 | 1168 | Delete the Licence in SharePoint
285 | 1168 | Install the licence using the Zeiss License Management Tool.
403 | 10467 | Install SmartSEM + ASAP Engine
403 | 10467 | Copy the SmartSEM sotware into a Folder on the Desktop. Name it SmaerSEM 8.0 Software
403 | 10467 | SmartSEM Version
403 | 10467 | SmartSEM Service Pack
403 | 10467 | SmartSEM Hotfix
403 | 10467 | ASAP Engine Version
403 | 10467 | Install SmartSEM + ASAP Engine
403 | 10467 | Copy the SmartSEM sotware into a Folder on the Desktop. Name it SmaerSEM 8.0 Software
403 | 10467 | SmartSEM Version
403 | 10467 | SmartSEM Service Pack
403 | 10467 | SmartSEM Hotfix
403 | 10467 | ASAP Engine Version
403 | 1167 | Download Licence from SharePoint
403 | 1167 | Delete the Licence in SharePoint
403 | 1167 | Install the licence using the Zeiss License Management Tool.
403 | 1167 | Download Licence from SharePoint
403 | 1167 | Delete the Licence in SharePoint
403 | 1167 | Install the licence using the Zeiss License Management Tool.
403 | 1167 | Download Licence from SharePoint
403 | 1167 | Delete the Licence in SharePoint
403 | 1167 | Install the licence using the Zeiss License Management Tool.
403 | 10467 | Install SmartSEM + ASAP Engine
403 | 10467 | Copy the SmartSEM sotware into a Folder on the Desktop. Name it SmaerSEM 8.0 Software
403 | 10467 | SmartSEM Version
403 | 10467 | SmartSEM Service Pack
403 | 10467 | SmartSEM Hotfix
403 | 10467 | ASAP Engine Version
285 | 3242 | Select PC type
285 | 3242 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
285 | 3242 | PC S/No.
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
285 | 3242 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Upgrade ASAP Firmware (check Service Pack for latest version) and API
285 | 3242 | Copy the SmartSEM sotware into a Folder on the Desktop. Name it SmartSEM 8.0 Software
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Set the Time and Date on the PC to current UK
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
102 | 1171 | Nano VP funtions
102 | 1171 | Collector Bias functions
387 | 11010 | Ensure EHT Earth Connection is Torqued to 3Nm
285 | 1168 | Download Licence from SharePoint
285 | 1168 | Delete the Licence in SharePoint
285 | 1168 | Install the licence using the Zeiss License Management Tool.
645 | 10761 | Functional Check
644 | 10510 | 500um Lower Aperture Fitted
644 | 10510 | 2nd Rotary Pump is ON in EP Mode
644 | 10510 | 10Pa Chamber Vacuum Achieved
644 | 10510 | 3000Pa Vacuum Achieved (Air)
644 | 10510 | Water Function Checked
644 | 10510 | Approx. 2000Pa Chamber Vacuum Achieved with Water
644 | 10510 | Go To HV, 2nd Rotary Pump is OFF
102 | 1171 | Nano VP funtions
102 | 1171 | Collector Bias functions
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448 , then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Install Smart SEM. Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Run ASAP Screening Program
20 | 252 | System Specific Datazone Loaded
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448 , then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Install Smart SEM (See WI 9693) Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Run ASAP Screening Program
20 | 252 | System Specific Datazone Loaded
21 | 387 | Order revision in use or 'n/a'
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448 , then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Install Smart SEM (See WI 9693) Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Run ASAP Screening Program
20 | 252 | System Specific Datazone Loaded
20 | 11255 | Check SCS Folder/Sales Order
20 | 11255 | SCS Number (If applicable)
20 | 254 | Language
25 | 11401 | 30um aperture entered on central aperture
27 | 286 | Stage Angle =
30 | 297 | Check Mag cal errors are <1% in X & Y axes in all 3 Mag Ranges. Orthoganality ± 0.1deg.
30 | 297 | Store accuracy images (Mag 1, Mag 2, Mag 3 and Orthog)
39 | 339 | After conditioning confirm there are no EMServer error messages
39 | 339 | Store Gun monitor trace as 30kV Cycle Conditioning.mn2
44 | 348 | Chamber Door Interlock
44 | 348 | Service Switch
44 | 348 | Lens De-activation
44 | 348 | Detector De-activation
44 | 348 | EHT + IPPSU De-activation
44 | 11256 | Check SCS Folder/Sales Order
44 | 349 | Record SEAWARD PRIMETEST 100 SN:
44 | 349 | Calibration Due Date:
44 | 349 | Tester verification test; Confirm earth short circuit = (<0.03Ω)
44 | 349 | Test all earth points as shown in procedure. (Max = 0.1Ω) Record max resistance shown
44 | 349 | Confirm correct SN label applied to product: (CSA Compliant / Non-CSA compliant)
44 | 353 | CalibSC.dat
44 | 353 | Customer Images
44 | 353 | License File
44 | 354 | Check Installed Licences Against Sales Order
44 | 11090 | All Warning labels Fitted.
44 | 356 | Check System And Dongle S/No. Match In Service Panel
44 | 356 | Un-tick Omit Dongle In Factory Panel
44 | 356 | Check Dongle fits in USB HUB (“ICY BOX”) and replace if needed.
44 | 356 | Check System And Dongle S/No. Match In Service Panel
44 | 356 | Un-tick Omit Dongle In Factory Panel
44 | 356 | Check Dongle fits in USB HUB (“ICY BOX”) and replace if needed.
82 | 973 | (10kV, Central Aperture) - Aperture and Guns <60%
82 | 973 | (10kV, Central Aperture) - Stigs <35%
82 | 973 | (10kV, Outer Apertures) - Stigs <45%
82 | 973 | (1kV, Central Aperture) - Aperture and Guns <60%
82 | 973 | (1kV, Central Aperture) - Stigs <20%
82 | 973 | (1kV, Outer Apertures) - Stigs <30%
82 | 973 | <10% user align wobble required on all apertures
388 | 1176
388 | 1176
388 | 1176
44 | 356 | Check System And Dongle S/No. Match In Service Panel
44 | 356 | Un-tick Omit Dongle In Factory Panel
44 | 356 | Check Dongle fits in USB HUB (“ICY BOX”) and replace if needed.
82 | 1593 | (10kV, Central Aperture) - Aperture and Guns <60%
82 | 1593 | (10kV, Central Aperture) - Stigs <20%
82 | 1593 | (10kV, Outer Apertures) - Stigs <30%
82 | 1593 | (1kV, Central Aperture) - Aperture and Guns <60%
82 | 1593 | (1kV, Central Aperture) - Stigs <20%
82 | 1593 | (1kV, Outer Apertures) - Stigs <30%
82 | 1593 | <10% user align wobble required on all apertures
45 | 10779 | Special Fitted
45 | 10779 | If Fitted, Checked by SCS Team Member Name
46 | 1174 | X-Ray Certificate
46 | 1174 | DENKA Data Sheet
46 | 1174 | Options Removed packing List
46 | 1174 | Options Retained packing List
46 | 1174 | Standard Packing List
46 | 1174 | X-Ray Certificate
46 | 1174 | DENKA Data Sheet
46 | 1174 | Options Removed packing List
46 | 1174 | Options Retained packing List
46 | 1174 | Standard Packing List
46 | 371 | X-Ray Certificate
46 | 371 | Quality Certificate
46 | 371 | Quality Feedback Form
46 | 371 | DENKA Data Sheet
46 | 371 | Options Removed packing List
46 | 371 | Options Retained packing List
46 | 371 | Standard Packing List
46 | 371 | PC Inferface assignment
46 | 371 | Declaration of conformity
46 | 371 | Installation report
46 | 371 | Acceptance test record
46 | 371 | Licence Generator report
388 | 11425 | 15kv 0.9nm
388 | 11425 | 15KV result
388 | 11425 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 1kv 1.3nm
388 | 11425 | 1KV result
388 | 11425 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 500V 2nm
388 | 11425 | 500V result
388 | 11425 | Store 500V Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 15kv 0.9nm
388 | 11425 | 15KV result
388 | 11425 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 1kv 1.3nm
388 | 11425 | 1KV result
388 | 11425 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 500V 2nm
388 | 11425 | 500V result
388 | 11425 | Store 500V Resolution Images and Pate image for 2 x consecutive images
285 | 3243 | Language
388 | 11426 | 15kv 0.7nm
388 | 11426 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 1kv 1.2nm
388 | 11426 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 500V 1.9nm
388 | 11426 | 500V result
388 | 11426 | Store 500V Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 15kv 0.9nm
388 | 11425 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 1kv 1.3nm
388 | 11425 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 500V 2nm
388 | 11425 | 500V result
388 | 11425 | Store 500V Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 15kv 0.9nm
388 | 11425 | 15KV result
388 | 11425 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 1kv 1.3nm
388 | 11425 | 1KV result
388 | 11425 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 500V 2nm
388 | 11425 | 500V result
388 | 11425 | Store 500V Resolution Images and Pate image for 2 x consecutive images
388 | 3292 | 1kV (Spec: $[Resolution1kV]nm)
388 | 3292 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 3292 | 15kV (Spec: $[Resolution15kV]nm)
388 | 3292 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 15kv 0.7nm
388 | 11426 | 15KV result
388 | 11426 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 1kv 1.2nm
388 | 11426 | 1KV result
388 | 11426 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 500V 1.9nm
388 | 11426 | 500V result
388 | 11426 | Store 500V Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 15kv 0.9nm
388 | 11425 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 1kv 1.3nm
388 | 11425 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 500V 2nm
388 | 11425 | 500V result
388 | 11425 | Store 500V Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 15kv 0.9nm
388 | 11425 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 1kv 1.3nm
388 | 11425 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 500V 2nm
388 | 11425 | Store 500V Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 15kv 0.7nm
388 | 11426 | 15KV result
388 | 11426 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 1kv 1.2nm
388 | 11426 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 500V 1.9nm
388 | 11426 | Store 500V Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 15kv 0.7nm
388 | 11426 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 1kv 1.2nm
388 | 11426 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 500V 1.9nm
388 | 11426 | Store 500V Resolution Images and Pate image for 2 x consecutive images
388 | 3292 | 1kV (Spec: $[Resolution1kV]nm)
388 | 3292 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 3292 | 15kV (Spec: $[Resolution15kV]nm)
388 | 3292 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
376 | 3250 | Sales order is latest revision
376 | 3250 | All accessories are accounted for and are fitted to correct ports
376 | 3250 | Identify any un-described items on sales order
388 | 1176 | 15kv 0.9nm
388 | 1176 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 1176 | 1kv 1.3nm
388 | 1176 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 1176 | 500V 2nm
388 | 1176 | Store 500V Resolution Images and Pate image for 2 x consecutive images
397 | 1177 | Fit beamsleeve funnel
397 | 1177 | Pneumatics and segments functional
397 | 1177 | SmartSEM Control tested
397 | 1177 | Images saved
285 | 11101 | Check SCS Folder/Sales Order
285 | 11101 | SCS Number (If applicable)
388 | 1176 | 15kv 0.9nm
388 | 1176 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 1176 | 1kv 1.3nm
388 | 1176 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 1176 | 500V 2nm
388 | 1176 | Store 500V Resolution Images and Pate image for 2 x consecutive images
388 | 1176 | 15kv 0.9nm
388 | 1176 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 1176 | 1kv 1.3nm
388 | 1176 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 1176 | 500V 2nm
388 | 1176 | Store 500V Resolution Images and Pate image for 2 x consecutive images
399 | 3398 | Uninstall MS Security Essentials
399 | 3398 | Remove test macro files
399 | 3398 | Delete Contents of Log Folder
399 | 3398 | Download Folder Images Deleted
399 | 3398 | Remove images from image gallery
399 | 3398 | Empty Recycle Bin
399 | 3398 | Clear Internet History
399 | 931 | PAT Tester Serial Number
399 | 931 | PAT Tester Cal Expiry Date
399 | 931 | Earth resistance – measured value must be < 0.1 Ω
399 | 931 | Insulation resistance – measured value must be > 1.0M Ω
399 | 931 | Test Certificate Completed & Signed
399 | 3400 | Em Server.accdb
399 | 3400 | Licence File
399 | 969 | Clean Front Panel
399 | 969 | Prepare Logo
399 | 969 | Fitting Logo
40 | 85 | Install ZEN core EM.Exe
40 | 85 | ZEN core Software Version
40 | 85 | Setup Exchange Position
40 | 85 | Capture Image
40 | 85 | Install ZEN core EM.Exe (This is NOT ZEN Core)
40 | 85 | ZEN core Software Version
40 | 85 | Setup Exchange Position
40 | 85 | Capture Image
399 | 3398 | Remove test macro files
399 | 3398 | Delete Contents of Log Folder
399 | 3398 | Download Folder Images Deleted
399 | 3398 | Remove images from image gallery
399 | 3398 | Empty Recycle Bin
399 | 3398 | Clear Internet History
646 | 10544 | Installation
646 | 10544 | Functional Test
646 | 10544 | Store SmartStitch Functional Image
646 | 10544 | Installation
646 | 10544 | Functional Test
646 | 10544 | Store SmartStitch Functional Image
646 | 10544 | SmartStitch Folder is present on Desktop
651 | 1181 | Check SmartStitch can be run
651 | 1181 | SmartStitch folder is present on the Desktop
651 | 1181
389 | 3313 | Installation
389 | 3313 | Functional Test
389 | 3313 | Store SmartStitch Functional Image
389 | 3313 | SmartStitch Folder is present on Desktop
40 | 11325 | Installation
40 | 11325 | Functional Test
40 | 11325 | Installation
40 | 11325 | Functional Test
40 | 11325 | Store SmartStitch Functional Image
40 | 11325 | SmartStitch Folder is present on Desktop
651 | 1179 | Download the CEU folder and copy to the Desktop.
651 | 1179 | Download the CEU folder and copy to the Desktop.
651 | 1179 | Install RClientinstaller.exe (if not already installed)
651 | 1179 | Run SEM_Calibration Extraction Tool.
651 | 1179 | Download the CEU folder and copy to the Desktop.
651 | 1179 | Install RClientinstaller.exe (if not already installed)
651 | 1179 | Run SEM_Calibration Extraction Tool.
651 | 1179 | Store the SEM_CEU.txt and EVO**-**-**.csv
651 | 1179 | Download the CEU folder and copy to the Desktop.
651 | 1179 | Install RClientinstaller.exe (if not already installed)
651 | 1179 | Run SEM_Calibration Extraction Tool.
651 | 1179 | Store the SEM_CEU.txt and EVO**-**-**.csv
651 | 1179 | If RClientinstaller.exe was just installed for the CEU, Uninstall.
651 | 1179 | Download the CEU folder and copy to the Desktop.
651 | 1179 | Install RClientinstaller.exe (if not already installed)
651 | 1179 | Run SEM_Calibration Extraction Tool.
651 | 1179 | Store the SEM_CEU.txt and EVO**-**-**.csv
651 | 1179 | If RClientinstaller.exe was just installed for the CEU, Uninstall.
651 | 1179 | Download the CEU folder and copy to the Desktop.
651 | 1179 | Install RClientinstaller.exe (if not already installed)
651 | 1179 | Run SEM_Calibration Extraction Tool.
651 | 1181 | Check SmartStitch can be run
651 | 1181 | SmartStitch folder is present on the Desktop
651 | 1181
30 | 906 | Store Images in QA folder
651 | 1179 | Download the CEU folder and copy to the Desktop.
651 | 1179 | Install RClientinstaller.exe (if not already installed)
651 | 1179 | Run SEM_Calibration Extraction Tool.
651 | 1179 | Store Standard Resolution Images + Oscillation Check
651 | 1179 | Download the CEU folder and copy to the Desktop.
651 | 1179 | Install RClientinstaller.exe (if not already installed)
651 | 1179 | Run SEM_Calibration Extraction Tool.
651 | 1179 | Store the SEM_CEU.txt and EVO**-**-**.csv files
398 | 1180 | Download the CEU folder and copy to the Desktop.
398 | 1180 | Install RClientinstaller.exe (if not already installed)
398 | 1180 | Run SEM_Calibration Extraction Tool.
398 | 1180 | Store the SEM_CEU.txt and Sigma***-**-**.csv files
398 | 1180 | Download the CEU folder and copy to the Desktop.
398 | 1180 | Install RClientinstaller.exe (if not already installed)
398 | 1180 | Run SEM_Calibration Extraction Tool.
398 | 1180 | Store the SEM_CEU.txt and Sigma***-**-**.csv files
398 | 1180 | If the RClientinstaller has been installed for the the CEU only, then uninstall
651 | 1179 | Download the CEU folder and copy to the Desktop.
651 | 1179 | Install RClientinstaller.exe (if not already installed)
651 | 1179 | Run SEM_Calibration Extraction Tool.
651 | 1179 | Store the SEM_CEU.txt and EVO**-**-**.csv files
651 | 1179 | If the RClientinstaller has been installed for the the CEU only, then uninstall
651 | 10580 | SmartSEM Version
651 | 10580 | Check SmartSEM 8 folder is present on the Desktop
393 | 3346 | Ensure all customer licences are loaded in the Zeiss Licence Management Tool and functions correctly
393 | 3346 | Check thst the SmartSEM 8 folder is present on the Desktop
661 | 1184 | Turbo cable fitted Correctly
661 | 1184 | Turbo cable fitted Correctly
403 | 10467 | Install SmartSEM + ASAP Engine
403 | 10467 | Copy the SmartSEM sotware into a Folder on the Desktop. Name it SmartSEM 8.* Software
403 | 10467 | SmartSEM Version
403 | 10467 | SmartSEM Service Pack
403 | 10467 | SmartSEM Hotfix
403 | 10467 | ASAP Engine Version
651 | 1179 | Download the CEU folder and copy to the Desktop.
651 | 1179 | Install RClientinstaller.exe (if not already installed)
651 | 1179 | Run SEM_Calibration Extraction Tool.
651 | 1179 | Store the SEM_CEU.txt and EVO**-**-**.csv files
651 | 1179 | If the RClientinstaller has been installed for the the CEU only, then uninstall
43 | 346 | 1kV =
43 | 346 | Store 1KV Resolution Images and Pate image
43 | 346 | 15kV =
43 | 346 | Store 1KV Resolution Images and Pate image
42 | 344 | Actual X-Axis
42 | 344 | Store X ripple shots and results
42 | 344 | Actual Y-Axis
42 | 344 | Store Y ripple shots and results
42 | 11013 | Actual X-Axis
42 | 11013 | Store X ripple shots and results
42 | 11013 | Actual Y-Axis
42 | 11013 | Store Y ripple shots and results
42 | 11047 | Actual X-Axis
42 | 11047 | Store X ripple shots and results
42 | 11047 | Actual Y-Axis
42 | 11047 | Store Y ripple shots and results
650 | 10569 | Check thet the following coditions are set= I Beam=20uA, Fil I=1.95A, Long Fil. Life selected
650 | 10569 | Secure Gun Head with 3 screws
33 | 316 | Actual for 1kV =
33 | 316 | Store 1KV Resolution Images and Pate image
33 | 316 | Actual for 15kV =
33 | 316 | Store 15KV Resolution Images and Pate image
33 | 11062 | 1kV =
33 | 11062 | Store 1KV Resolution Images and Pate image
33 | 11062 | 15kV =
33 | 11062 | Store 15KV Resolution Images and Pate image
388 | 1176 | 15kv 0.9nm
388 | 1176 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 1176 | 1kv 1.3nm
388 | 1176 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 1176 | 500V 2nm
388 | 1176 | Store 500V Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 15kv 0.9nm
388 | 11425 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 1kv 1.3nm
388 | 11425 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 500V 2nm
388 | 11425 | Store 500V Resolution Images and Pate image for 2 x consecutive images
650 | 10569 | Check thet the following coditions are set= I Beam=20uA, Fil I=1.95A, Long Fil. Life selected
650 | 10569 | Secure Gun Head with 3 M6x18(000000-0004-529) screws
40 | 902 | Dowel pins have been mounted correctly to the mounting plate
388 | 11426 | 15kv 0.7nm
388 | 11426 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 1kv 1.2nm
388 | 11426 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 500V 1.9nm
388 | 11426 | Store 500V Resolution Images and Pate image for 2 x consecutive images
388 | 3292 | 1kV (Spec: $[Resolution1kV]nm)
388 | 3292 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 3292 | 15kV (Spec: $[Resolution15kV]nm)
388 | 3292 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
40 | 11081 | Fit Test tool BB Electronics
40 | 11081 | Function Test
40 | 11081 | Remove Test tool BB Electronics
40 | 1559 | Macros loaded/tested
40 | 1559 | $exchange set up
40 | 1559 | Rod Position Correctly Adjusted
40 | 1559 | Front Panel Control all Functional
644 | 10497 | System Vacuum & Gun Vacuum is within Spec
644 | 10497 | Gun Chamber has been cleaned and inspected using UV light
644 | 10497 | EHT set socket well has been cleaned and inspected using UV light
644 | 10497 | Clean Anode in Plasma Oven for 20 Minutes
644 | 10497 | Check Cathode is aligned before fitting
644 | 10497 | Lab6 Firing Unit Fitted
644 | 10497 | Overnight Bakeout
652 | 10618 | Radiation Cert
652 | 10618 | Customer Entitlement Certificate
652 | 10618 | Itemes removed
652 | 10618 | Items retained
652 | 10618 | System and accessories packing list
5 | 1257 | Fit SCM
5 | 1257 | Fit Nav Cam
646 | 10526 | Number of Chamberscope's on Order
646 | 10526 | Chamberscope 1 - Select Port
646 | 10526 | Check Camera for focus
646 | 10526 | Functional Check-Black/White + Colour Mode
646 | 10526 | Chamberscope 2 - Select Port or Not Fitted
646 | 10526 | Chamberscope 2 - Check Camera 2 for Focus
646 | 10526 | Chamberscope 2 - Functional Check-Black/White + Colour Mode
646 | 10526 | Chamberscope 2 - Run wxPropview to Remove Defective Pixels
646 | 10526 | Chamberscope 2 - Add Icon to Toolbar
646 | 11100 | Select Port
646 | 11100 | Check Camera for focus
646 | 11100 | Functional Check-Black/White + Colour Mode
646 | 11100 | Add Icon to Toolbar
650 | 10567 | Delete the Production Test Licences
650 | 10567 | Obtain the customer Entitlement Certificate
650 | 10567 | Download into the Engineer folder
650 | 10567 | Install the Customer Licences
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer.db to the Product Folder
650 | 10567 | Print out the Customer Entitlement Certicate - To be sent to the customer
650 | 10567 | Store EM Server.db
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
403 | 10463 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
403 | 10463 | Hard Drive 1 Size
403 | 10463 | Hard Drive 2 Size
403 | 10463 | RAM
403 | 10463 | Processor
403 | 10463 | Motherboard Model
651 | 669 | The Customer Entitlement Certificate is present in the Engineer folder
651 | 669 | The EM Server.db is present in the Engineer folder
644 | 10498 | Gun Monitor Setup
644 | 10498 | Test No. 96 Selected
644 | 10498 | Fit Hitek EHT set
644 | 10498 | Run up to 30KV and leave for 15 minute
644 | 10498 | Shutdown and run up to 30KV, Repeat 5 times
644 | 10498 | Run up to 30KV and leave for 30 minutes (There must be no spikes)
644 | 10498 | 30KV Gun Monitor Traced Saved
644 | 10760 | No X-Ray Leakage Detected
652 | 10618 | Radiation Cert
652 | 10618 | Customer Entitlement Certificate
652 | 10618 | Itemes removed
652 | 10618 | Items retained
652 | 10618 | System and accessories packing list
390 | 3321 | Delete the Production Test Licences
390 | 3321 | Obtain the customer Entitlement Certificate
390 | 3321 | Download into the Engineer folder
390 | 3321 | Print out the Customer Entitlement Certificate
390 | 3321 | Load customer licences and check against order
390 | 3321 | Test customer licenses
390 | 3321 | Delete the Production Test Licences
390 | 3321 | Obtain the customer Entitlement Certificate
390 | 3321 | Download into the Engineer folder
390 | 3321 | Print out the Customer Entitlement Certificate - To be shipped to the customer
390 | 3321 | Load customer licences and check against order
390 | 3321 | Test customer licenses
390 | 3321 | Delete the Production Test Licences
390 | 3321 | Obtain the customer Entitlement Certificate
390 | 3321 | Download into the Engineer folder
390 | 3321 | Print out the Customer Entitlement Certificate - To be sent to the customer
390 | 3321 | Load customer licences and check against order
390 | 3321 | Test customer licenses
642 | 10476 | ASAP Mainboard Revision
642 | 10476 | ASAP Mainboard Hardware
642 | 10476 | ASAP VBE 1 Revision
642 | 10476 | ASAP VBE 1 Firmware Version
642 | 10476 | ASAP VBE 2 (if fitted) Revision
642 | 10476 | ASAP VBE 2 (if fitted) Firmware Version
642 | 10476 | ASAP DFE 1 Revision
642 | 10476 | ASAP DFE 1 Firmware Version
642 | 10476 | ASAP DFE 2 (if fitted) Revision
642 | 10476 | ASAP DFE 2 (if fitted) Firmware Version
642 | 10476 | ASAP DFE 3 (if fitted) Revision
642 | 10476 | ASAP DFE 3 (if fitted) Firmware Version
642 | 10476 | ASAP DFE 4 (if fitted) Revision
642 | 10476 | ASAP DFE 4 (if fitted) Firmware Version
403 | 10463 | Select PC
403 | 10463 | PC Serial No
643 | 10487 | Connect the Hitek Test Tool EHT Set
643 | 10487 | Gun Monitor Setup
643 | 10487 | Test No. 96 Selected
643 | 10487 | EHT Stable at 30KV - ≥15 minutes
643 | 10487 | 30KV Gun Monitor Traced Saved
5 | 1257 | Fit SCM
5 | 1257 | Fit Nav Cam
400 | 3405 | X-Ray Leakage Check
400 | 3405 | Quality Certificate
400 | 3405 | Customer Entitlement Certificate
400 | 3405 | Gun/Aperture Align Sheet
400 | 3405 | Denka Data Sheet
400 | 3405 | Test Accessories Packing list
400 | 3405 | Standard Packing List
400 | 3405 | Items retained on Microscope Packing List
400 | 3405 | Items removed from Microscope Packing List
400 | 3405 | China RoHS Certificate
642 | 10476 | ASAP has connected
642 | 10476 | ASAP Mainboard Hardware
642 | 10476 | ASAP VBE 1 Revision
642 | 10476 | ASAP VBE 1 Firmware Version
642 | 10476 | ASAP VBE 2 (if fitted) Revision
642 | 10476 | ASAP VBE 2 (if fitted) Firmware Version
642 | 10476 | ASAP DFE 1 Revision
642 | 10476 | ASAP DFE 1 Firmware Version
642 | 10476 | ASAP DFE 2 (if fitted) Revision
642 | 10476 | ASAP DFE 2 (if fitted) Firmware Version
642 | 10476 | ASAP DFE 3 (if fitted) Revision
642 | 10476 | ASAP DFE 3 (if fitted) Firmware Version
642 | 10476 | ASAP DFE 4 (if fitted) Revision
642 | 10476 | ASAP DFE 4 (if fitted) Firmware Version
642 | 10476 | ASAP has connected
642 | 10476 | ASAP Mainboard Hardware
642 | 10476 | ASAP VBE 1 Revision
642 | 10476 | ASAP VBE 1 Firmware Version
642 | 10476 | ASAP VBE 2 (if fitted) Revision
642 | 10476 | ASAP VBE 2 (if fitted) Firmware Version
642 | 10476 | ASAP DFE 1 Revision
642 | 10476 | ASAP DFE 1 Firmware Version
642 | 10476 | ASAP DFE 2 (if fitted) Revision
642 | 10476 | ASAP DFE 2 (if fitted) Firmware Version
642 | 10476 | ASAP DFE 3 (if fitted) Revision
642 | 10476 | ASAP DFE 3 (if fitted) Firmware Version
642 | 10476 | ASAP DFE 4 (if fitted) Revision
642 | 10476 | ASAP DFE 4 (if fitted) Firmware Version
642 | 10476 | ASAP has connected
642 | 10476 | ASAP Mainboard Hardware
642 | 10476 | ASAP VBE 1 Revision
642 | 10476 | ASAP VBE 1 Firmware Version
642 | 10476 | ASAP VBE 2 (if fitted) Revision
642 | 10476 | ASAP VBE 2 (if fitted) Firmware Version
642 | 10476 | ASAP DFE 1 Revision
642 | 10476 | ASAP DFE 1 Firmware Version
642 | 10476 | ASAP DFE 2 (if fitted) Revision
642 | 10476 | ASAP DFE 2 (if fitted) Firmware Version
642 | 10476 | ASAP DFE 3 (if fitted) Revision
642 | 10476 | ASAP DFE 3 (if fitted) Firmware Version
642 | 10476 | ASAP DFE 4 (if fitted) Revision
642 | 10476 | ASAP DFE 4 (if fitted) Firmware Version
642 | 10476 | ASAP has connected
642 | 10476 | Stage PCB has connected
642 | 10476 | EHT Set has connected
642 | 10476 | Vacuum PCB has connected
642 | 10476 | EO PCB has connected
651 | 1179 | Download the CEU folder and copy to the Desktop.
651 | 1179 | Install RClientinstaller.exe (if not already installed)
651 | 1179 | Run SEM_Calibration Extraction Tool.
651 | 1179 | Store the SEM_CEU.txt and EVO**-**-**.csv files
651 | 1179 | If the RClientinstaller has been installed for the the CEU only, then uninstall
38 | 1187 | Configuration dependent:
38 | 1187 | Actual Value
401 | 3411 | Licensed software on Sales Order matches Customer Entitlement Certificate
44 | 354 | Delete the Production Test Licences
44 | 354 | Obtain the customer Entitlement Certificate
44 | 354 | Download into the Engineer folder
44 | 354 | Install the Customer Licences
44 | 354 | Copy Contents of Image Folder to the Product Folder
44 | 354 | Copy EMServer.db to the Product Folder
44 | 354 | Print out the Customer Entitlement Certicate - To be sent to the customer
83 | 1596 | Check that the licences installed in the Zeiss Licence Management Tool are correct to the Sales Order and that they functio
46 | 371 | X-Ray Certificate
46 | 371 | Quality Certificate
46 | 371 | Customer Entitlement Certificate
46 | 371 | Quality Feedback Form
46 | 371 | DENKA Data Sheet
46 | 371 | Options Removed packing List
46 | 371 | Options Retained packing List
46 | 371 | Standard Packing List
46 | 371 | PC Inferface assignment
46 | 371 | Declaration of conformity
46 | 371 | Installation report
46 | 371 | Acceptance test record
46 | 371 | Licence Generator report
651 | 11397 | Change in Working Distance ≤0.2mm
643 | 10488 | Ensure Gun Monitor is Running and Test 96 Selected.
643 | 10488 | Carry Out Initial Gun Run Up
643 | 10488 | Connect Customers Spellman EHT set
643 | 10488 | Carry out Gun Run Up
643 | 10488 | X Gun Tilt ≤50%
643 | 10488 | Y Gun Tilt ≤50%
643 | 10488 | X Gun Shift ≤50%
643 | 10488 | Y Gun Shift ≤50%
642 | 10476 | Check ASAP connection
642 | 10476 | Check Stage Connection
642 | 10476 | Check EHT Set connection
642 | 10476 | Check Vacuum PCB connection
642 | 10476 | Check EO PCB connection
648 | 10553 | 30KV Res- ≤2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 1KV Res- ≤15nm
648 | 10553 | 200V Res- Not measured
648 | 10553 | Store Standard Resolution Images
648 | 10553 | 30KV Res- ≤2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 1KV Res- ≤15nm
648 | 10553 | 200V Res- Not measured
648 | 10553 | Store Standard Resolution Images
643 | 10488 | Ensure Gun Monitor is Running and Test 96 Selected.
643 | 10488 | Carry Out Initial Gun Run Up
643 | 10488 | Connect the Customers Spellman EHT Set
643 | 10488 | Carry Out Gun Run Up
643 | 10488 | X Gun Tilt ≤50%
643 | 10488 | Y Gun Tilt ≤50%
643 | 10488 | X Gun Shift ≤50%
643 | 10488 | Y Gun Shift ≤50%
398 | 1180 | Download the CEU folder and copy to the Desktop.
398 | 1180 | Install RClientinstaller.exe (if not already installed)
398 | 1180 | Run SEM_Calibration Extraction Tool.
398 | 1180 | Store the SEM_CEU.txt and Sigma***-**-**.csv files
398 | 1180 | If the RClientinstaller has been installed for the the CEU only, then uninstall
45 | 362 | EM Server.db
45 | 362 | Em Server.Log
45 | 362 | MyBackup.smartbackup
45 | 362 | L500.ini
45 | 362 | License file.xml (N/A for Crossbeam)
45 | 362 | Registry file.reg
45 | 362 | CalibSC.dat
45 | 11024 | EM Server.db
45 | 11024 | Em Server.Log
45 | 11024 | MyBackup.smartbackup
45 | 11024 | L500.ini
45 | 11024 | License file.xml (N/A for Crossbeam)
45 | 11024 | Registry file.reg
45 | 11024 | CalibDC.dat
45 | 11024 | Export Tablefiller
20 | 11038 | PC S/N
20 | 11038 | If SEM Enhanced PC(349535-9001-020) has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Hard disk Image (Version)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Copy Macro Files to C:/support tools
20 | 11038 | Approved Datazone Loaded
403 | 1134 | Print Shipment Labels
403 | 1134 | Label Material Boxes
403 | 1134 | Place left over stickes in Document Holder
195 | 11244 | No X-Ray Leakage detected.
644 | 10498 | Conect the Hitek EHT Set
644 | 10498 | Gun Monitor Setup
644 | 10498 | Test No. 96 Selected
644 | 10498 | Run up to 30KV and leave for 15 minute
644 | 10498 | Shutdown and run up to 30KV, Repeat 5 times
644 | 10498 | Run up to 30KV and leave for 30 minutes (There must be no spikes)
644 | 10498 | 30KV Gun Monitor Traced Saved
644 | 10500 | Gun Monitor Setup
644 | 10500 | Test No. 96 Selected
644 | 10500 | New Filament Selected
644 | 10500 | Long Fil. Life Selected
644 | 10500 | Lab6 Firing Unit Run Up to 30KV
644 | 10500 | Connect the Customers Spellman EHT Set
644 | 10500 | Run the LaB6 to 30KV
644 | 10500 | X Gun Tilt ≤50%
644 | 10500 | Y Gun Tilt ≤50%
644 | 10500 | X Gun Shift ≤50%
644 | 10500 | Y Gun Shift ≤50%
644 | 10500 | Gun Monitor Setup
644 | 10500 | Test No. 96 Selected
644 | 10500 | New Filament Selected
644 | 10500 | Long Fil. Life Selected
644 | 10500 | Carry Out Initial LaB6 Gun Run Up to 30KV
644 | 10500 | Connect the Customers Spellman EHT Set
644 | 10500 | Run Up KaB6 to 30KV
644 | 10500 | Run the LaB6 to 30KV
644 | 10500 | X Gun Tilt ≤50%
644 | 10500 | Y Gun Tilt ≤50%
644 | 10500 | X Gun Shift ≤50%
644 | 10500 | Y Gun Shift ≤50%
648 | 10611 | 30KV Res- ≤2nm
648 | 10611 | 3KV Res- ≤6nm
648 | 10611 | 1KV Res- ≤9nm
648 | 10611 | 200V Res- Not Measured
648 | 10611 | Store Standard LaB6 Resolution
644 | 10500 | Gun Monitor Setup
644 | 10500 | Test No. 96 Selected
644 | 10500 | New Filament Selected
644 | 10500 | Long Fil. Life Selected
644 | 10500 | Carry Out Initial LaB6 Gun Run Up to 30KV
644 | 10500 | Connect the Customers Spellman EHT Set
644 | 10500 | Run the LaB6 to 30KV
644 | 10500 | X Gun Tilt ≤50%
644 | 10500 | Y Gun Tilt ≤50%
644 | 10500 | X Gun Shift ≤50%
644 | 10500 | Y Gun Shift ≤50%
403 | 10465 | Windows 10 Version
403 | 10465 | Windows 10 Version
403 | 10465 | Disable Drivers and Windows Updates
403 | 10463 | Select PC
403 | 10463 | PC Serial No
403 | 1167 | Check that Time & Date on PC are set to current UK
403 | 1167 | Download Licence from SharePoint
403 | 1167 | Delete the Licence in SharePoint
403 | 1167 | Install the licence using the Zeiss License Management Tool.
403 | 1134 | Print Shipment Labels
403 | 1134 | Label Material Boxes
403 | 1134 | Place left over stickes in Document Holder
403 | 1134 | Print Shipment Labels
403 | 1134 | Label Material Boxes
403 | 1134 | Place left over stickes in Document Holder
403 | 1134 | Print Shipment Labels
403 | 1134 | Label Material Boxes
403 | 1134 | Place left over stickes in Document Holder
650 | 10569 | Check thet the following coditions are set= I Beam=20uA, Fil I=1.60A, Long Fil. Life selected
650 | 10569 | Secure Gun Head with 3 M6x18(000000-0004-529) screws
651 | 669 | The Customer Entitlement Certificate is present in the Engineer folder
651 | 669 | The EM Server.db is present in the Engineer folder
4 | 1190 | Was the Stage baked independently
20 | 11038 | PC S/N
20 | 11038 | If SEM Enhanced PC(349535-9001-020) has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Hard disk Image (Version)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Approved Datazone Loaded
386 | 3415 | No X-Ray Leakage detected.
665 | 11193 | Serial Number
665 | 11193 | Check that the Assembly is the Correct Part Number
665 | 11193 | Correct Collector Beak Fitted
665 | 11193 | Fit Detector Assembly
665 | 11193 | Apply Emersion Oil to PM Tube Face
665 | 11193 | Fit PM Head Amp Assembly
30 | 907 | XB Tilt Test carried out
665 | 11193 | Serial Number (or NA)
665 | 11193 | Check that the Assembly is the Correct Part Number
665 | 11193 | Correct Collector Beak Fitted
665 | 11193 | Fit Detector Assembly
665 | 11193 | Apply Emersion Oil to PM Tube Face
665 | 11193 | Fit PM Head Amp Assembly
676 | 11194 | Check that the Assembly is the Correct Part Number
676 | 11194 | Serial Number (or NA)
676 | 11194 | Correct Collector Beak Fitted
676 | 11194 | Fit Detector Assembly
676 | 11194 | Apply Emersion Oil to PM Tube Face
676 | 11194 | Fit PM Head Amp Assembly
659 | 10685 | Earth Plate
659 | 10685 | Cable Plates
659 | 10685 | Bellows
659 | 10685 | Bellows is concentric to hole in the desktop
661 | 10710 | Torque Chamber Bolts to 35Nm
661 | 10710 | Record Chamber SN
661 | 10710 | Check Bellows is concentric to the hole in the desktop.
4 | 1190 | Was the Stage baked independently
4 | 1190 | Was the Stage baked independently
285 | 1135 | Print shipment labels
285 | 1135 | Label material boxes
285 | 1135 | Place left over stickers in document holder
4 | 37 | Shim Stage
4 | 37 | Check Mag Clamp (adjust position if necessary, see WI)
4 | 37 | Fit Gun Interlock Key to Stage Door
4 | 37 | Record Stage SN (Can be found on right hand side of XY platform)
4 | 37
4 | 37 | Shim Stage
4 | 37 | Check Mag Clamp (adjust position if necessary, see WI)
4 | 37 | Fit Gun Interlock Key to Stage Door
4 | 37 | Record Stage SN (Can be found on right hand side of XY platform)
4 | 1190 | Was the Stage baked independently? (Chamber not baked)
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
665 | 10737 | Diode Holder Assembly aligned using jig.
665 | 10737 | All screw fully tightened
665 | 10737 | Check alignment is Ok again using jig
707 | 11290 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11290 | Multi_US_HRQA - Passed inspection criteria, Image Saved
389 | 3303 | Function Test
389 | 3303 | Disable Confirmation Check
389 | 3303 | Turbo Load Check
389 | 3303 | Save .CSV of quite mode cycle time
707 | 11308 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11308 | Multi_US_HRQA - Passed inspection criteria, Image Saved
707 | 11308 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11308 | Multi_US_HRQA - Passed inspection criteria, Image Saved
646 | 10541 | Installation of Quiet Mode Vessel
646 | 10541 | Quiet Mode Fitted Selected in SEM Status
646 | 10541 | Compare On/Off Times With Graph
646 | 10541 | Pump Starts When Quiet Mode is Deselected
646 | 10541 | Run Gun Monitor-Turbo Power and System Vacuum
708 | 11239 | Hold Anode, look for NO Holes between the screw holes and aperture. No particles are to be obsered
707 | 11290 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11290 | Multi_US_HRQA - Passed inspection criteria, Image Saved
707 | 11308 | Check there is an engraved ring around the cluster of apertures, this should be on the underside
707 | 11308 | Multi_US_HRQA - Passed inspection criteria, Image Saved
44 | 355 | Ensure the S/N & P/N match the License generator report / Sales order.
44 | 353 | CalibSC.dat
44 | 353 | Customer Images
44 | 353 | License File
44 | 11064 | CalibDC.dat
44 | 11064 | Customer Images
44 | 11064 | License File
44 | 11090 | All Warning labels Fitted.
44 | 358 | Confirm all deviations have been cleared on OST - The Andon lamp should be green
44 | 358 | All verification & Safety tests done and within threshold limits defined
44 | 358 | Sales Order is latest version
44 | 358 | All items on Sales order accounted for
26 | 281 | Check concentricity of ESB wrt scanning system. (Limit <= 30um)
26 | 281 | Measure the ESB hole diametre (Limit <= 180um)
26 | 281 | Store ESB Concentricity Image
26 | 396 | Type in measured value (Limit: <= 50um)
26 | 396 | Type in measured value (Limit: <= 50um)
26 | 396 | Store ESB Concentricity image
26 | 396 | Type in measured value (Limit: <= 50um)
26 | 396 | Measure the ESB hole diametre (Limit <= 150um)
26 | 396 | Store ESB Concentricity image
26 | 396 | Check concentricity of ESB wrt scanning system. (Limit: <= 50um)
26 | 396 | Measure the ESB hole diametre (Limit <= 150um)
26 | 396 | Store ESB Concentricity image
26 | 396 | Check concentricity of ESB wrt scanning system. (Limit: <= 50um)
26 | 396 | Measure the ESB hole diametre (Limit <= 150um)
26 | 396 | Store ESB Concentricity image
26 | 281 | Check concentricity of ESB wrt scanning system. (Limit <= 30um)
26 | 281 | Measure the ESB hole diametre (Limit <= 180um)
26 | 281 | Store ESB Concentricity Image
26 | 396 | Check concentricity of ESB wrt scanning system. (Limit: <= 50um)
26 | 396 | Measure the ESB hole diametre (Limit <= 150um)
26 | 396 | Store ESB Concentricity image
26 | 281 | Check concentricity of ESB wrt scanning system. (Limit <= 30um)
26 | 281 | Measure the ESB hole diametre (Limit <= 180um)
26 | 281 | Store ESB Concentricity Image
26 | 396 | Check concentricity of ESB wrt scanning system. (Limit: <= 50um)
26 | 396 | Measure the ESB hole diametre (Limit <= 150um)
26 | 396 | Store ESB Concentricity image
26 | 281 | Check concentricity of ESB wrt scanning system. (Limit <= 30um)
26 | 281 | Measure the ESB hole diametre (Limit <= 180um)
26 | 281 | Store ESB Concentricity Image with annotations
5 | 1256 | Fit SCM
26 | 396 | Check concentricity of ESB wrt scanning system. (Limit: <= 50um)
26 | 396 | Measure the ESB hole diametre (Limit <= 150um)
26 | 396 | Store ESB Concentricity image with annotations
26 | 396 | Check concentricity of ESB wrt scanning system. (Limit: <= 50um)
26 | 396 | Measure the ESB hole diameter (Limit <= 150um)
26 | 396 | Store ESB Concentricity image with annotations
26 | 281 | Check concentricity of ESB wrt scanning system. (Limit <= 30um)
26 | 281 | Measure the ESB hole diameter (Limit <= 180um)
26 | 281 | Store ESB Concentricity Image with annotations
26 | 281 | Check concentricity of ESB wrt scanning system. (Limit <= 30um)
26 | 281 | Measure the ESB hole diameter (Limit <= 180um)
26 | 281 | Store ESB Concentricity Image with annotations
26 | 396 | Check concentricity of ESB wrt scanning system. (Limit: <= 50um)
26 | 396 | Measure the ESB hole diameter (Limit <= 150um)
26 | 396 | Store ESB Concentricity image with annotations
678 | 621 | Confirm PAT test passed and results stored
678 | 621 | Confirm PAT test passed and results stored
21 | 256 | Correct Firmware installed
21 | 256 | Store System checker full report
21 | 256 | Correct Firmware installed
21 | 256 | Store System checker full report
678 | 621 | Confirm PAT test passed and results stored
21 | 256 | Please confirm stage controller version and Axes version are compatible (see 'Stage FW Compatability' docuement attached)
21 | 256 | Store System checker full report
21 | 256 | Ensure the Stage Controller version is compatible with the Stage Axes version (see 'Stage FW Compatability' document attached)
21 | 256 | Store System checker full report
21 | 256 | Ensure the Stage Controller version is compatible with the Stage Axes version (See compatible verions in drop down list)
21 | 256 | Store System checker full report
21 | 256 | Ensure the Stage Controller version is compatible with the Stage Axes version (See compatible verions in drop down list)
21 | 256 | Store System checker full report
21 | 256 | Ensure the Stage Controller version is compatible with the Stage Axes version (See compatible verions in drop down list)
21 | 256 | Store System checker full report
26 | 281 | Check concentricity of ESB wrt scanning system.
26 | 281 | Measure the ESB hole diameter
26 | 281 | Store ESB Concentricity Image with annotations
26 | 396 | Check concentricity of ESB wrt scanning system. (Limit: <= 50um)
26 | 396 | Measure the ESB hole diameter (Limit <= 150um)
26 | 396 | Store ESB Concentricity image with annotations
26 | 396 | Check concentricity of ESB wrt scanning system. (Limit: <= 50um)
26 | 396 | Measure the ESB hole diameter
26 | 396 | Store ESB Concentricity image with annotations
649 | 10555 | BSD Resolution ≤4nm.
649 | 10555 | Store BSD Resolution Image
678 | 621 | PAT Checkbox Calibration Date
678 | 621 | Earth Bond Test TP1 (0.1Ω - 0.07Ω)
678 | 621 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
678 | 621 | Insulation Test (2.0MΩ - 1.9MΩ)
678 | 621 | Earth Leakage Test (0.9mA - 0.7mA)
678 | 951 | Secutest SN:
678 | 951 | Secutest Calibration Date:
678 | 951 | PE Test - Confirm all test points (<0.1Ω)
678 | 951 | R INS - ( > 0.5MΩ)
678 | 951 | I HLNC - ( < 500µA)
678 | 951 | I HLSF - ( < 3500µA)
678 | 951 | Confirm EN 61010 Passed
678 | 951 | Confirm Visual Inspection Criterea Passed
678 | 951 | Record maximum leakage current
678 | 1203 | IGP Heaters (IGP#1) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 1203 | IGP Heaters (IGP#2) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 1203 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 1203 | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 1203 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 1203 | IGP Heaters (IGP#1) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 1203 | IGP Heaters (IGP#2) (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 1203 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 1203 | IGP FLange Heater (#2) (346022-8124-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 1203 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 1204 | HAL SN:
678 | 1204 | HAL Calibration Date:
678 | 1204 | IGP Heaters (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 1204 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 1204 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 1204 | IGP Heaters (000000-0513-455). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 1204 | IGP Flange Heater (#1) (346022-8122-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
678 | 1204 | Cathode Head Heater (346022-8120-000). . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
285 | 3247 | Set Air Suspension Height
285 | 3247 | Check Correct Operation
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448 , then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Check the Entitlement certificate to determine which version of Smartsem to install
20 | 252 | Install Smart SEM (See WI 9693) Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Run ASAP Screening Program
20 | 252 | System Specific Datazone Loaded
389 | 11074 | Check each individual segment in normal and inverse modes
389 | 11074 | Analyze Segments
389 | 11074 | Gain Ranges
389 | 11074 | Pneumatics
389 | 11074 | Alignment
389 | 11074 | Vac Degradation < 3x10-5 mbar
389 | 11074 | CZ Detector Cal
389 | 11074 | Z Adjustment
26 | 281 | Check concentricity of ESB wrt scanning system.
26 | 281 | Measure the ESB hole diameter
26 | 281 | Store ESB Concentricity Image with annotations
26 | 281 | Check concentricity of ESB wrt scanning system.
26 | 281 | Measure the ESB hole diameter
26 | 281 | Store ESB Concentricity Image with annotations
26 | 396 | Check concentricity of ESB wrt scanning system. (Limit: <= 50um)
26 | 396 | Measure the ESB hole diameter.
26 | 396 | Store ESB Concentricity image with annotations
643 | 1273 | Mains Sync set to OFF
8 | 1274 | Run a 20min plasma cylce
8 | 1274 | Run a 20min plasma cycle
26 | 281 | Check concentricity of ESB wrt scanning system.
26 | 281 | Measure the ESB hole diameter.
26 | 281 | Store ESB Concentricity Image with annotations
648 | 10553 | 30KV Res- ≤2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 1KV Res- ≤15nm
648 | 10553 | 200V Res- Not measured
648 | 10553 | Store Standard Resolution Images
648 | 10554 | 2KV Res - 100um VP Aperture- ≤16nm
648 | 10554 | 2KV Res - 400um Easy VP Aperture- ≤16nm
648 | 10554 | Store VP & Easy VP Aperture Resolution Images
648 | 10611 | 30KV Res- ≤2nm
648 | 10611 | 3KV Res- ≤6nm
648 | 10611 | 1KV Res- ≤9nm
648 | 10611 | 200V Res- Not Measured
648 | 10611 | Store Standard LaB6 Resolution
648 | 10612 | 2KV Res - 100um VP Apertutre- ≤12nm
648 | 10612 | Store VP Aperture (100u) Resolution Images
648 | 10881 | 2KV Res - 750um Easy VP Aperture- ≤12nm
648 | 10881 | Store Easy VP Aperture (750u) Resolution Images
678 | 1232 | System Vacuum Reading OK
678 | 1232 | Gun Vacuum Reading OK
678 | 1232 | Initialise Stage
648 | 10614 | 2KV Res - 750um Easy Align Aperture- ≤12nm
648 | 10614 | Store Easy Align Aperture (750u) Resolution Images
648 | 10875 | 2KV Res - 750um Easy Align Aperture- ≤16nm
648 | 10875 | Store Easy Align Aperture (750u) Resolution Images
648 | 10658 | 2KV - 100um VP Aperture- ≤12nm
648 | 10658 | 2KV - 140um VP Aperture- ≤12nm
648 | 10658 | Store 2KV Aperture Resolution Images
649 | 230 | BSD Resolution ≤4nm.
649 | 230 | Store BSD Resolution Image
389 | 950 | Fit the Insitu substage
389 | 950 | Test the Insitu substage
389 | 950 | System vac <2.0e-06mbar
389 | 950 | X-ray test
389 | 950 | Repackage all hardware excluding the chamber feed through in Deben packaging
40 | 948 | Fit the Insitu substage
40 | 948 | Test the Insitu substage
40 | 948 | System vac <2.0e-06mbar
40 | 948 | X-ray test
40 | 948 | Repackage all hardware excluding the chamber feed through in Deben packaging
663 | 10731 | EHT Set Type Fitted
663 | 10731 | EHT Serial Number
663 | 10731 | Revision/Mod Level
663 | 10731 | Torque Earth Connection to 3Nm
30 | 906 | Store Images in QA folder
649 | 203 | BSD Resolution 3.8nm
649 | 203 | Store BSD Resolution Image
649 | 10556 | 30KV Resolution - Easy VP- ≤3.4nm
649 | 10556 | Store VPSE Resolution Image
649 | 204 | VPSE Resolution 3.0nm
649 | 204 | Store VPSE Resolution Image
649 | 11206 | 30KV Res- ≤2nm
649 | 11206 | Store 30KV ( Large ZTR)Resolution Image
649 | 10659 | 30KV Resolution- ≤5.5nm
649 | 10659 | Store VPSE Resolution Image
649 | 10557 | Resolution @ 7KV- ≤9nm
649 | 10557 | Store Beam Deceleration Resolution Image
81 | 1588 | Measured value =
81 | 1588 | Save annotated image of Max Current
22 | 264 | Tip No.
22 | 264 | Confirm that the Tip S/N on the Denka datasheet is the same as the Tip S/N on the column sticker.
22 | 264 | Filled out Normal Nominal Values in Gun Parameters.
646 | 10535 | Black Level Calibration
646 | 10535 | Functional Check
646 | 10535 | Function Image
646 | 10535 | Store SCD Functional Image
646 | 10535 | Black Level Calibration
646 | 10535 | Functional Check
646 | 10535 | Function Image
646 | 10535 | Store SCD Functional Image
40 | 186 | Install 3DSM Software and Record Software Version
40 | 186 | Set Up Image Using One segment of the BSD
40 | 186 | Run 3DSM and Setup
40 | 186 | Save 3DSM Project
40 | 186 | Store 3DSM Project
40 | 11080 | Vacuum Stability Test
40 | 11080 | Slew Rate <2.5us
40 | 11080 | Measure Specimen Current into a FC and then blank the beam with the EBBU. Confirm that measured Specimen current is zero.
40 | 11081 | Fit Test tool BB Electronics
40 | 11081 | Vacuum Stability Test
40 | 11081 | Slew Rate <2.5us
40 | 11081 | Measure Specimen Current into a FC and then blank the beam with the EBBU. Confirm that measured Specimen current is zero.
40 | 11081 | Remove Test tool BB Electronics
30 | 906 | Store Images in Image gallery Stage Folder
30 | 906 | Compare photographs to confirm that position of specimen holder has not changed
646 | 10540 | Fit 300um Aperture - Required for Tunsten Systems only
646 | 10540 | Installation
646 | 10540 | Cable Open Circuit to Ground
646 | 10540 | Blanking Tests Using 20um Aperture
646 | 10540 | Blanking Tests Using 30um Aperture
646 | 10540 | Blank/UnBlank Ratio is better than 0.1% for all settings
646 | 10540 | Low Mag Blanking Check
4 | 668 | Fit chamber to damper
4 | 668 | Fit chamber bolts item 10 and 12 and Torque to 42Nm.
4 | 668 | Fit chamber bolts item 13 and 14 and Torque to 21Nm.
4 | 668 | Fit chamber pumping tube and Toruqe bolts item 7 to 12Nm.
4 | 668 | Fit Turbo guard with non silver vent screw 000000-0539-927.
4 | 668 | Record Chamber SN
644 | 10760 | Serial No. (or N/A)
644 | 10760 | No X-Ray Leakage Detected
195 | 11244 | Serial No. (or N/A)
195 | 11244 | No X-Ray Leakage detected.
386 | 3415 | Serial No. (or N/A)
386 | 3415 | No X-Ray Leakage detected.
14 | 1649 | IGP 1 Pressure (≤ 9.0e-10 mbar)
14 | 1649 | IGP 2 Pressure
21 | 895 | Record Water Flow Temperature (20ºC +/- 4ºC)
21 | 895 | Set Water Flow Meter to 0.9 lpm
30 | 303 | Actual (Spec: 6.7nm / min)
30 | 303 | Store stage drift macro results image
42 | 11013 | Actual X-Axis
42 | 11013 | Store X ripple shots and results
42 | 11013 | Actual Y-Axis
42 | 11013 | Store Y ripple shots and results
661 | 10710 | Check Stage Rail Grub Screws are Tight
661 | 10710 | Torque Chamber Bolts to 35Nm
661 | 10710 | Record Chamber SN
661 | 10710 | Check Bellows is concentric to the hole in the desktop.
671 | 10797 | Check Stage Rail Grub Screws are Tight
671 | 10797 | Torque Chamber Bolts To 10Nm
671 | 10797 | Record Chamber SN
2 | 1218 | Check if ASAP or h-ASAP
2 | 1218 | h-ASAP Serial No (if applicable, if not applicable enter NA)
661 | 10720 | Turbo Pump Serial Number
661 | 10720 | Qualification Marks Checked
656 | 1219 | Check if ASAP or h-ASAP
656 | 1219 | h-ASAP Serial No (if applicable, if not applicable enter NA)
30 | 303 | Actual (Spec: 6.7nm / min)
30 | 303 | Store stage drift macro results image
40 | 452 | Check each individual segment in normal and inverse modes (Save all images to aSTEM folder)
40 | 452 | Analyze Segments
40 | 452 | Gain Ranges
40 | 452 | Pneumatics
40 | 452 | Alignment
40 | 452 | Stage Safe Position
40 | 452 | Vac Degradation < 3x10-5 mbar
40 | 452 | CZ Detector Cal
40 | 452 | Store Holder Position 1
40 | 453 | Check each individual segment in normal and inverse modes (Save all images to aSTEM folder)
40 | 453 | Analyze Segments
40 | 453 | Gain Ranges
40 | 453 | Pneumatics
40 | 453 | Alignment
40 | 453 | Stage Safe Position
40 | 453 | Vac Degradation < 3x10-5 mbar
40 | 453 | CZ Detector Cal
40 | 453 | Store Holder Position 1
17 | 1229 | Create new folder (pole piece serial number) within Low kV Isolator Concentricity Tracking folder. Upload concentricity photo taken with Dino-lite into this folder
30 | 299 | Record Kompass results (µm)
30 | 299 | Store Kompass results window
678 | 10851 | ASAP FW/HW Present
678 | 10851 | Stage FW/HW Present
678 | 10851 | Vacuum FW/HW Present
678 | 10851 | EO PCB FW/HW Present
659 | 1220 | Check if ASAP or h-ASAP
659 | 1220 | h-ASAP Serial No (if applicable, if not applicable enter NA)
663 | 10731 | EHT Set Type Fitted
663 | 10731 | EHT Serial Number
663 | 10731 | Revision/Mod Level
663 | 10731 | Qualification Marks Checked
663 | 10731 | Torque Earth Connection to 3Nm
663 | 10730 | Valves Cables
663 | 10730 | Heater Cables
663 | 10730 | Ion Pump Cable
663 | 10730 | Ion Pump PSU Serial Number
663 | 10730 | Fit Ion Pump PSU Bracket
663 | 10730 | Fit Ion Pump PSU
663 | 10730 | PSU Qualification Marks Checked
665 | 11260 | Detector Switch Box Qualification Marks Checked
665 | 11260 | Detector Switch Box SN
665 | 11260 | Detector Switch Box Qualification Marks Checked
285 | 1168 | Download Licence from SharePoint
285 | 1168 | Delete the Licence in SharePoint
285 | 1168 | Install the licence using the Zeiss License Management Tool.
285 | 1168 | Installation of Zen licence if Hasap is fitted using the Zeiss activation tool
403 | 1167 | Check that Time & Date on PC are set to current UK
403 | 1167 | Download Licence from SharePoint
403 | 1167 | Delete the Licence in SharePoint
403 | 1167 | Install the licence using the Zeiss License Management Tool.
403 | 1167 | Installation of Zen licence
678 | 1232 | System Vacuum Reading OK
678 | 1232 | Gun Vacuum Reading OK - Confirm gun vacuum is improving (Not saturated)
678 | 1232 | Initialise Stage
8 | 1645 | Serial:
8 | 1645 | Remove metal protective cover from detector and inspect detector head
30 | 1224 | carry out scanning calibration (Hasap only)
30 | 1230 | Stage mapping completed successfuly
678 | 1232 | System Vacuum Reading OK
678 | 1232 | Gun Vacuum Reading OK - Confirm gun vacuum is improving (Not saturated)
678 | 1232 | Initialise Stage
678 | 1232 | System Vacuum Reading OK
678 | 1232 | Gun Vacuum Reading OK - Confirm gun vacuum <3x10-6 mbar (If vacuum showing no improvment please contact ME)
678 | 1232 | Initialise Stage - Confirm stage initialises without errors
46 | 11099 | X-Ray Certificate
46 | 11099 | Quality Certificate
46 | 11099 | SVSEM FIB EXT Datasheet
46 | 11099 | DENKA Data Sheet
46 | 11099 | Options Removed packing List
46 | 11099 | Options Retained packing List
46 | 11099 | Standard Packing List
46 | 11099 | Send all supplier paperwork to Oberkochen
678 | 1232 | System Vacuum Reading - Confirm system vacuum reads < 1x10-5mbar
678 | 1232 | Gun Vacuum Reading OK - Confirm gun vacuum <3x10-6 mbar (If vacuum showing no improvment please contact ME)
678 | 1232 | Initialise Stage - Confirm stage initialises without errors
46 | 11123 | Remove docking brackets and fit transit brackets
46 | 11123 | Power system down.
46 | 11123 | Blow water out of the cooling system.
46 | 11123 | Put thread lock on pendulum screws and transit block Screws.
46 | 11123 | Chamber secured to vibration plate (N/A to XB 550L)
46 | 11123 | Chamber securing screws tightened to 25nm (N/A to XB 550L)
46 | 11123 | Chamber securing screws tightness checked. Type your initials (For XB 550L type N/A)
46 | 11123 | Secure Airfeet with foam blocks.
46 | 11123 | Secure Turbo pump with foam blocks.
46 | 11123 | Fit Covers to all unused connectors.
46 | 11123 | Mains and Earth cable stowed on Mains unit.
46 | 11123 | Fit shipping Eyes to transit blocks.
46 | 11123 | All doors closed and secured.
46 | 11123 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 11123 | Print shipping stickers.
46 | 11123 | Fit 6x M8x16mm cap head screws to stage door
46 | 1233 | Remove docking brackets and fit transit brackets
46 | 1233 | Power system down.
46 | 1233 | Blow water out of the cooling system.
46 | 1233 | Put thread lock on pendulum screws and transit block Screws.
46 | 1233 | Chamber secured to vibration plate (N/A to XB 550L)
46 | 1233 | Chamber securing screws tightened to 25nm (N/A to XB 550L)
46 | 1233 | Chamber securing screws tightness checked. Type your initials (For XB 550L type N/A)
46 | 1233 | Secure Airfeet with foam blocks.
46 | 1233 | Secure Turbo pump with foam blocks.
46 | 1233 | Fit Covers to all unused connectors.
46 | 1233 | Mains and Earth cable stowed on Mains unit.
46 | 1233 | Fit shipping Eyes to transit blocks.
46 | 1233 | All doors closed and secured.
46 | 1233 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 1233 | Print shipping stickers.
46 | 1233 | Fit 8x M6x55mm cap head screws to stage door
8 | 11336 | Serial:
8 | 11336 | Remove metal protective cover from detector and inspect detector head
5 | 1256 | Fit SCM
382 | 1225 | carry out scanning calibration (Hasap only)
645 | 1226 | carry out scanning calibration (Hasap only)
661 | 10720 | Turbo Pump Serial Number
661 | 10720 | Qualification Marks Checked
663 | 10731 | EHT Set Type Fitted
663 | 10731 | EHT Serial Number
663 | 10731 | Revision/Mod Level
663 | 10731 | Qualification Marks Checked
663 | 10731 | Torque Earth Connection to 3Nm
663 | 10730 | Valves Cables
663 | 10730 | Heater Cables
663 | 10730 | Ion Pump Cable
663 | 10730 | Ion Pump PSU Serial Number
663 | 10730 | Fit Ion Pump PSU Bracket
663 | 10730 | Fit Ion Pump PSU
663 | 10730 | PSU Qualification Marks Checked
665 | 10743 | Fit Stage Elecrtical Feedthrough
665 | 10743 | PCB Fitted
665 | 10743 | PSU Fitted
665 | 10743 | PSU Qualification Marks Checked
665 | 10743 | Fit Stage Elecrtical Feedthrough
665 | 10743 | PCB Fitted
665 | 10743 | PSU Qualification Marks Checked
665 | 10743 | PSU Fitted
665 | 10743 | Cables Fitted
8 | 94 | Serial:
8 | 94 | Remove metal protective cover from detector and inspect detector head
46 | 1233 | Remove docking brackets and fit transit brackets
46 | 1233 | Power system down.
46 | 1233 | Blow water out of the cooling system.
46 | 1233 | Put thread lock on pendulum screws and transit block Screws.
46 | 1233 | Chamber secured to vibration plate (N/A to XB 550L)
46 | 1233 | Chamber securing screws tightened to 25nm (N/A to XB 550L)
46 | 1233 | Chamber securing screws tightness checked. Type your initials (For XB 550L type N/A)
46 | 1233 | Secure Airfeet with foam blocks.
46 | 1233 | Secure Turbo pump with foam blocks.
46 | 1233 | Fit Covers to all unused connectors.
46 | 1233 | Mains and Earth cable stowed on Mains unit.
46 | 1233 | Fit shipping Eyes to transit blocks.
46 | 1233 | All doors closed and secured.
46 | 1233 | Strap fitted to secure pendulum airfeet (bubble wrap protecting switch panel).
46 | 1233 | Print shipping stickers.
46 | 1233 | Fit 8x M6x55mm cap head screws to stage door
6 | 1236 | Print and fit the relevant S/N label to the system. See WI for details.
6 | 1236 | Print and fit the relevant S/N label to the system. See WI for details.
6 | 1236 | Confirm SN label has NO CSA or SEMI markings.
6 | 1236 | Print and fit the relevant S/N label to the system. See WI for details.
6 | 1236 | Confirm SN label does not have any CSA or SEMI markings.
6 | 11345 | Print and fit the relevant S/N label to the system. See WI for details.
6 | 11345 | Confirm SN label does not have any CSA or SEMI markings.
6 | 1253 | Confirm all FW is correct (Update any FW not latest version)
8 | 3295 | Serial:
8 | 3295 | Remove metal protective cover from detector and inspect detector head
5 | 10959 | Warning - ONLY VENT using the electronic vent valve in the chamber
5 | 10959 | No Leaks Detected (Prior to running up Turbo)
5 | 10959 | No Leaks Detected (With Turbo running) when in high sensitivity
5 | 10959 | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng
663 | 414 | No Leaks Detected (Prior to running up Turbo)
663 | 414 | No Leaks Detected (With Turbo running) when in high sensitivity
663 | 414 | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng
674 | 416 | No Leaks Detected (Prior to running up Turbo)
674 | 416 | No Leaks Detected (With Turbo running) when in high sensitivity
674 | 416 | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng
27 | 431 | Nominal Extractor Voltage =
27 | 431 | Anode Diameter =
6 | 1236 | Print and fit the relevant S/N label to the system. See WI for details.
6 | 1236 | Confirm SN label does not have any CSA markings.
6 | 1236 | Confirm SN label does not have any SEMI markings.
6 | 1236 | Confirm SN label does not have any CSA markings.
6 | 1236 | Confirm SN label does not have any SEMI markings.
6 | 11345 | Confirm SN label does not have any CSA or SEMI markings.
6 | 11345 | Confirm SN label does not have any CSA or SEMI markings.
40 | 1560 | Disable/Enable
40 | 1560 | Save Vac Trace
6 | 11345 | Confirm SN label does not have any CSA markings.
6 | 11345 | Confirm SN label does have the correct SEMI marking.
6 | 1234 | Confirm SN label does not have any CSA markings.
6 | 1234 | Confirm SN label does have the correct SEMI marking.
6 | 1234 | Confirm SN label does not have any CSA markings.
6 | 1234 | Confirm SN label does not have any SEMI marking.
6 | 1234 | Confirm SN label does have the correct CSA markings.
6 | 1234 | Confirm SN label does not have any SEMI marking.
6 | 1234 | Confirm SN label does have the correct CSA marking.
6 | 1234 | Confirm SN label does not have any SEMI marking.
6 | 11345 | Confirm SN label does not have any CSA markings.
6 | 11345 | Confirm SN label does have the correct SEMI marking.
6 | 1234 | Confirm SN label does have the correct CSA marking.
6 | 1234 | Confirm SN label does not have any SEMI marking.
6 | 1235 | Confirm SN label does have the correct CSA marking.
6 | 1235 | Confirm SN label does not have any SEMI marking.
6 | 1235 | Confirm SN label does have the correct CSA marking.
6 | 1235 | Confirm SN label does have the correct SEMI marking.
6 | 1234 | Confirm SN label does have the correct CSA marking.
6 | 1234 | Confirm SN label does not have any SEMI marking.
6 | 11345 | Confirm SN label does not have any CSA markings.
6 | 11345 | Confirm SN label does have the correct SEMI marking.
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448 , then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Check the Entitlement certificate to determine which version of Smartsem to install
20 | 252 | Install Smart SEM (See WI 9693) Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Run ASAP Screening Program
20 | 252 | System Specific Datazone Loaded
661 | 10720 | Turbo Pump Serial Number
661 | 10720 | Fit Turbo Earth Cable
661 | 10720 | Qualification Marks Checked
389 | 89 | Test Valves Using Raman Test Interlock Box
389 | 89 | Install SemRamanStage Software
389 | 89 | Run SEM Raman Correlation
389 | 89 | Save Calibration File as Raman
389 | 89 | second chamberscope fitted
389 | 89 | Run Smart Stage Mapping
6 | 1236 | Confirm SN label does not have any CSA markings.
6 | 1236 | Confirm SN label does not have any SEMI markings.
| 1239 | S/N =
6 | 1236 | Confirm SN label does not have any CSA markings.
6 | 1236 | Confirm SN label does not have any SEMI markings.
656 | 10663 | Damper Plate is Sigma Version (Additional airlock hole)
656 | 10663 | Damper Plate is Located Correctly
6 | 11345 | Confirm SN label does not have any CSA markings.
6 | 11345 | Confirm SN label does have the correct SEMI marking.
6 | 1234 | Confirm SN label does have the correct CSA marking.
6 | 1234 | Confirm SN label does not have any SEMI marking.
6 | 1235 | Confirm SN label does have the correct CSA marking.
6 | 1235 | Confirm SN label does have the correct SEMI marking.
6 | 1235 | Confirm SN label does have the correct CSA marking.
6 | 1235 | Confirm SN label does have the correct SEMI marking.
40 | 1241 | Fit the HT-BSD
40 | 1241 | Software setup of the HTBSD
40 | 1241 | Test the HTBSD
40 | 1241 | X-ray check HT-BSD
40 | 1241 | System vac in spec (<2.0e-06mbar)
40 | 1241 | Remove detector before shipping
40 | 1241 | Fit the Rapid BSD
40 | 1241 | Software setup of the Rapid BSD
40 | 1241 | Test the Rapid BSD
40 | 1241 | X-ray check Rapid BSD
40 | 1241 | System vac in spec (<2.0e-06mbar)
40 | 1241 | Remove detector before shipping
| 1242 | Create new folder (pole piece serial number) within Low kV Isolator Concentricity Tracking folder. Upload concentricity photo taken with Dino-lite into this folder
22 | 266 | Test Power Supply Serial Number
22 | 266 | Use the 35kV External supply for 30mins.
22 | 266 | Save Gun Moitor Log as GunConditioning_TIP SN
22 | 266 | Tighten gun cable screws and confirm they are a minimum torque of 1Nm with torque driver.
25 | 11035 | With Cond. Off Max Aper. X & Y < 30%
25 | 11035 | Gun Head Clamping Screws Secure
25 | 11035 | Tighten gun head clamping screws and confirm they are a minimum torque of 1Nm with torque driver.
25 | 11035 | Condenser Clamping Screws Secure
25 | 11035 | Tighten condensor clamping screws and confirm they are a minimum torque of 1Nm with torque driver.
25 | 275 | Are Gun Head clamping screws secure?
25 | 275 | Tighten gun head clamping screws and confirm they are a minimum torque of 1Nm with torque driver.
25 | 275 | Are Condensor clamping screws secure?
25 | 275 | Tighten condensor clamping screws and confirm they are a minimum torque of 1Nm with torque driver.
25 | 275 | Store 1kV emission image (Condensor off)
25 | 275 | Check emission image against 'emissions distortion criteria'
25 | 275 | Store 10kV emission image (Condensor off)
25 | 11313 | Are Gun Head clamping screws secure?
25 | 11313 | Tighten gun head clamping screws and confirm they are a minimum torque of 1Nm with torque driver.
25 | 11313 | Are Condensor clamping screws secure?
25 | 11313 | Tighten condensor clamping screws and confirm they are a minimum torque of 1Nm with torque driver.
25 | 11313 | Store 1kV emission image (Condensor off)
25 | 11313 | Store 10kV emission image (Condensor off)
6 | 628 | Secutest SN:
6 | 628 | Secutest Calibration Date:
6 | 628 | PE Test - Confirm all test points (<0.1Ω)
6 | 628 | R INS - ( > 0.5MΩ)
6 | 628 | I HLSF - ( < 8000µA)
6 | 628 | Confirm EN 61010 Passed
6 | 628 | Confirm Visual Inspection Criterea Passed
678 | 621 | PAT Checkbox Calibration Date
678 | 621 | Earth Bond Test TP1 (0.1Ω - 0.07Ω)
678 | 621 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
678 | 621 | Insulation Test (2.0MΩ - 1.9MΩ)
678 | 621 | Earth Leakage Test (0.9mA - 0.7mA)
678 | 951 | Seaward Supernova SN:
678 | 951 | Seaward Calibration Date:
678 | 951 | PE Test - Confirm all test points (<0.1Ω)
678 | 951 | R INS - ( > 0.5MΩ)
678 | 951 | I HLSF - ( < 6000µA)
678 | 951 | Confirm EN 61010 Passed
678 | 951 | Confirm Visual Inspection Criterea Passed
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | R INS - ( > 0.5MΩ)
667 | 11413 | I HLSF - ( < 6000µA)
667 | 11413 | Confirm EN 61010 Passed
667 | 11413 | Confirm Visual Inspection Criterea Passed
25 | 427 | Confirm that the contents of Tablefiller are deleted before starting the procedure. Stored values can affect the alignment routine.
25 | 427 | Mechanically align the column using the Merlin Alignment Wizard
25 | 427 | Once adjusters are backed off, tighten C1 + C2 clamping screws and confirm they are a minimum torque of 1Nm with torque driver.
25 | 427 | Confirm that the contents of Tablefiller are deleted before starting the procedure. Stored values can affect the alignment routine.
25 | 427 | Mechanically align the column using the Merlin Alignment Wizard
25 | 427 | Tighten Gun Head clamping screws and confirm they are a minimum torque of 1Nm with torque driver.
25 | 427 | Once adjusters are backed off, tighten C1 + C2 clamping screws and confirm they are a minimum torque of 1Nm with torque driver.
25 | 427 | Confirm that the contents of Tablefiller are deleted before starting the procedure. Stored values can affect the alignment routine.
25 | 427 | Mechanically align the column using the Merlin Alignment Wizard
25 | 427 | Tighten Gun Head clamping screws and confirm the 3 out of 4 accessible are a minimum torque of 1Nm with torque driver.
25 | 427 | Once adjusters are backed off, tighten C1 + C2 clamping screws and confirm they are a minimum torque of 1Nm with torque driver.
25 | 275 | Are Gun Head clamping screws secure?
25 | 275 | Tighten Gun Head clamping screws and confirm the 3 out of 4 accessible are a minimum torque of 1Nm with torque driver.
25 | 275 | Are Condensor clamping screws secure?
25 | 275 | Tighten condensor clamping screws and confirm they are a minimum torque of 1Nm with torque driver.
25 | 275 | Store 1kV emission image (Condensor off)
25 | 275 | Check emission image against 'emissions distortion criteria'
25 | 275 | Store 10kV emission image (Condensor off)
25 | 11313 | Are Gun Head clamping screws secure?
25 | 11313 | Tighten Gun Head clamping screws and confirm the 3 out of 4 accessible are a minimum torque of 1Nm with torque driver.
25 | 11313 | Are Condensor clamping screws secure?
25 | 11313 | Tighten condensor clamping screws and confirm they are a minimum torque of 1Nm with torque driver.
25 | 11313 | Store 1kV emission image (Condensor off)
25 | 11313 | Store 10kV emission image (Condensor off)
25 | 11035 | With Cond. Off Max Aper. X & Y < 30%
25 | 11035 | Gun Head Clamping Screws Secure
25 | 11035 | Tighten Gun Head clamping screws and confirm the 3 out of 4 accessible are a minimum torque of 1Nm with torque driver.
25 | 11035 | Condenser Clamping Screws Secure
25 | 11035 | Tighten condensor clamping screws and confirm they are a minimum torque of 1Nm with torque driver.
678 | 951 | Secutest SN:
678 | 951 | Secutest Calibration Date:
678 | 951 | PE Test - Confirm all test points (<0.1Ω)
678 | 951 | R INS - ( > 0.5MΩ)
678 | 951 | I HLNC - ( < 500µA)
678 | 951 | I HLSF - ( < 3500µA)
678 | 951 | Confirm EN 61010 Passed
678 | 951 | Confirm Visual Inspection Criterea Passed
678 | 951 | Record maximum leakage current
40 | 916 | Unpack Volutome package and check it's completeness
40 | 916 | Ensure Arivis Software Licence is available (410145-1031-100)
40 | 916 | Remove transit clamps from Volutome stage
40 | 916 | Ensure hardware has been configured according to work instructions
40 | 916 | Fit Volutome stage
40 | 916 | Connect Volutome stage and controller to the SEM
40 | 916 | Install Volutome software installation package
40 | 916 | Configure Volutome stage motors
40 | 916 | Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator
40 | 916 | Set up the Volutome workbench in Zen
40 | 916 | Record VolumeBSD serial number
40 | 916 | Record Volutome stage serial number
40 | 916 | Record VolumeBSD amplifier serial number
40 | 916 | Record VolumeBSD mechanics serial number
40 | 916 | Record C-885 PI controller serial number
40 | 916 | Record E-727 PI controller serial number
40 | 916 | Function check the Volutome stage light
40 | 916 | Adjust Volutome stage centre (Crosshairs should be within 4 central squares)
40 | 916 | Install Nav cam in-line connections (If applicable)
40 | 916 | Perform the microtome and stage functionality checks
40 | 916 | Fit and align the VolumeBSD diode
40 | 916 | Check VolumeBSD alignment repeatability
40 | 916 | Check all gain ranges of the VolumeBSD (normal/inverted)
40 | 916 | Perform Signal/noise ratio check
40 | 916 | Ensure BSD diode protector is fitted and adjusted in chamber
40 | 916 | Install cable tie to loop BSD signal cable
40 | 916 | Check SCM with Volutome stage
40 | 916 | Oscillation check - Measure the amplitude of the oscillation
40 | 916 | Perform X-ray check
40 | 916 | Pump Down Time
40 | 916 | Ultimate Vacuum
40 | 916 | Focal CC fitted to Volutome stage
40 | 916 | Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar
40 | 916 | Generate live image from Axiocam
40 | 916 | Re-package Volutome and check it's completeness
40 | 916 | ManEng/ Senior Eng/ Production Manger Sign off
389 | 917 | Unpack Volutome package and check it's completeness
389 | 917 | Ensure Arivis Software Licence is available (410145-1031-100)
389 | 917 | Remove transit clamps from Volutome stage
389 | 917 | Ensure hardware has been configured according to work instructions
389 | 917 | Fit Volutome stage
389 | 917 | Connect Volutome stage and controller to the SEM
389 | 917 | Install Volutome software installation package
389 | 917 | Configure Volutome stage motors
389 | 917 | Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator
389 | 917 | Set up the Volutome workbench in Zen
389 | 917 | Record VolumeBSD serial number
389 | 917 | Record Volutome stage serial number
389 | 917 | Record VolumeBSD amplifier serial number
389 | 917 | Record VolumeBSD mechanics serial number
389 | 917 | Record C-885 PI controller serial number
389 | 917 | Record E-727 PI controller serial number
389 | 917 | Function check the Volutome stage light
389 | 917 | Adjust Volutome stage centre (Crosshairs should be within 4 central squares)
389 | 917 | Install Nav cam in-line connections (If applicable)
389 | 917 | Perform the microtome and stage functionality checks
389 | 917 | Fit and align the VolumeBSD diode
389 | 917 | Check VolumeBSD alignment repeatability
389 | 917 | Check all gain ranges of the VolumeBSD (normal/inverted)
389 | 917 | Perform Signal/noise ratio check
389 | 917 | Ensure BSD diode protector is fitted and adjusted in chamber
389 | 917 | Install cable tie to loop BSD signal cable
389 | 917 | Check SCM with Volutome stage
389 | 917 | Oscillation check - Measure the amplitude of the oscillation
389 | 917 | Perform X-ray check
389 | 917 | Pump Down Time
389 | 917 | Ultimate Vacuum
389 | 917 | Focal CC fitted to Volutome stage
389 | 917 | Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar
389 | 917 | Generate live image from Axiocam
389 | 917 | Re-package Volutome and check it's completeness
389 | 917 | ManEng/ Senior Eng/ Production Manger Sign off
6 | 628 | Secutest SN:
6 | 628 | Secutest Calibration Date:
6 | 628 | PE Test - Confirm all test points (<0.1Ω)
6 | 628 | R INS - ( > 0.5MΩ)
6 | 628 | I HLSF - ( < 8000µA)
6 | 628 | Confirm EN 61010 Passed
6 | 628 | Confirm Visual Inspection Criterea Passed
6 | 628 | Secutest SN:
6 | 628 | Secutest Calibration Date:
6 | 628 | PE Test - Confirm all test points (<0.1Ω)
6 | 628 | R INS - ( > 0.5MΩ)
6 | 628 | I HLSF - ( < 8000µA)
6 | 628 | Confirm EN 61010 Passed
6 | 628 | Confirm Visual Inspection Criterea Passed
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | R INS - ( > 0.5MΩ)
667 | 11413 | I HLSF - ( < 6000µA)
667 | 11413 | Confirm EN 61010 Passed
667 | 11413 | Confirm Visual Inspection Criterea Passed
644 | 10506 | Vp Functional Check
644 | 10506 | Fit Label to the Black Plastic Aperture Box
678 | 621 | PAT Checkbox Calibration Date
678 | 621 | Earth Bond Test TP1 (0.1Ω - 0.07Ω)
678 | 621 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
678 | 621 | Insulation Test (2.0MΩ - 1.9MΩ)
678 | 621 | Earth Leakage Test (0.9mA - 0.7mA)
678 | 951 | Seaward Supernova SN:
678 | 951 | Secutest Calibration Date:
678 | 951 | PE Test - Confirm all test points (<0.1Ω)
678 | 951 | R INS - ( > 0.5MΩ)
678 | 951 | I HLNC - ( < 500µA)
678 | 951 | I HLSF - ( < 3500µA)
678 | 951 | Confirm EN 61010 Passed
678 | 951 | Confirm Visual Inspection Criterea Passed
678 | 951 | Record maximum leakage current
678 | 951 | Seaward Supernova SN:
678 | 951 | Seaward Calibration Date:
678 | 951 | PE Test - Confirm all test points (<0.1Ω)
678 | 951 | R INS - ( > 0.5MΩ)
678 | 951 | I HLSF - ( < 3500µA)
678 | 951 | Confirm EN 61010 Passed
678 | 951 | Confirm Visual Inspection Criterea Passed
678 | 951 | Record maximum leakage current
678 | 951 | Seaward Supernova SN:
678 | 951 | Seaward Calibration Date:
678 | 951 | PE Test - Confirm all test points (<0.1Ω)
678 | 951 | R INS - ( > 0.5MΩ)
678 | 951 | I HLSF - ( < 6000µA)
678 | 951 | Confirm EN 61010 Passed
678 | 951 | Confirm Visual Inspection Criterea Passed
676 | 10829 | Fit Adaptor Plate to Chamber
676 | 10829 | Fit Amplifier
676 | 10829 | Check and Setup Push Pull Assembly
676 | 10829 | Fit Push Pull Assembly
663 | 10728 | Scan Cables
663 | 10728 | Lens Cables
663 | 10728 | EHT Cables
663 | 10728 | Penning Cable
25 | 198 | Gun X
25 | 198 | Gun Y
25 | 198 | Aper X
25 | 198 | Aper Y
663 | 10728 | Scan Cables
663 | 10728 | Lens Cables
663 | 10728 | EHT Cables
663 | 10728 | Penning Cable
6 | 1547 | PAT Checkbox Calibration Date
6 | 1547 | Earth Bond Test TP1 (0.1Ω - 0.09Ω)
6 | 1547 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
6 | 1547 | Insulation Test (2.0MΩ - 1.9MΩ)
6 | 1547 | Earth Leakage Test (0.9mA - 0.7mA)
667 | 11412 | PAT Checkbox Calibration Date
667 | 11412 | Earth Bond Test TP1 (0.1Ω - 0.07Ω)
667 | 11412 | Earth Bond Test TP2 (0.55Ω - 0.52Ω)
667 | 11412 | Insulation Test (2.0MΩ - 1.9MΩ)
667 | 11412 | Earth Leakage Test (0.9mA - 0.7mA)
667 | 11413 | Secutest SN:
667 | 11413 | Secutest Calibration Date:
667 | 11413 | PE Test - Confirm all test points (<0.1Ω)
667 | 11413 | R INS - ( > 0.5MΩ)
667 | 11413 | I HLSF - ( < 6000µA)
667 | 11413 | Confirm EN 61010 Passed
667 | 11413 | Confirm Visual Inspection Criterea Passed
26 | 396 | Check concentricity of ESB wrt scanning system. (Limit: <= 50um)
26 | 396 | Measure the ESB hole diameter.
26 | 396 | Record the ESB serial number.
26 | 396 | Store ESB Concentricity image. Add ESB serial number as a text annotation in SmartSEM. (filename must be exactly esbdiameter.tif)
27 | 284 | Confirm the 60um aperture emission image still visible in High current mode
382 | 6931 | Drift
26 | 281 | Check concentricity of ESB wrt scanning system.
26 | 281 | Record the ESB serial number.
26 | 281 | Measure the ESB hole diameter.
26 | 281 | Store ESB Concentricity image. Add ESB serial number as a text annotation in SmartSEM. (filename must be exactly esbdiameter.tif)
26 | 281 | Check concentricity of ESB wrt scanning system.
26 | 281 | Record the ESB serial number - This is the S/N that is shown on the ESB label (1026-XXX or 1125-XXX).
26 | 281 | Measure the ESB hole diameter.
26 | 281 | Store ESB Concentricity image. Add ESB serial number as a text annotation in SmartSEM. (filename must be exactly esbdiameter.tif)
26 | 396 | Check concentricity of ESB wrt scanning system. (Limit: <= 50um)
26 | 396 | Measure the ESB hole diameter.
26 | 396 | Record the ESB serial number - This is the S/N that is shown on the ESB label (1026-XXX or 1125-XXX).
26 | 396 | Store ESB Concentricity image. Add ESB serial number as a text annotation in SmartSEM. (filename must be exactly esbdiameter.tif)
43 | 345 | Refer to Work Instruction for column conditions
43 | 345 | Store contamination shot
40 | 916 | Unpack Volutome package and check it's completeness
40 | 916 | Ensure Arivis Software Licence is available (410145-1031-100)
40 | 916 | Upload data from USB stick to the O drive product folder
40 | 916 | Remove transit clamps from Volutome stage
40 | 916 | Ensure hardware has been configured according to work instructions
40 | 916 | Fit Volutome stage
40 | 916 | Connect Volutome stage and controller to the SEM
40 | 916 | Install Volutome software installation package
40 | 916 | Configure Volutome stage motors
40 | 916 | Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator
40 | 916 | Set up the Volutome workbench in Zen
40 | 916 | Record VolumeBSD serial number
40 | 916 | Record Volutome stage serial number
40 | 916 | Record VolumeBSD amplifier serial number
40 | 916 | Record VolumeBSD mechanics serial number
40 | 916 | Record C-885 PI controller serial number
40 | 916 | Record E-727 PI controller serial number
40 | 916 | Function check the Volutome stage light
40 | 916 | Adjust Volutome stage centre (Crosshairs should be within 4 central squares)
40 | 916 | Install Nav cam in-line connections (If applicable)
40 | 916 | Perform the microtome and stage functionality checks
40 | 916 | Fit and align the VolumeBSD diode
40 | 916 | Check VolumeBSD alignment repeatability
40 | 916 | Check all gain ranges of the VolumeBSD (normal/inverted)
40 | 916 | Perform Signal/noise ratio check
40 | 916 | Ensure BSD diode protector is fitted and adjusted in chamber
40 | 916 | Install cable tie to loop BSD signal cable
40 | 916 | Check SCM with Volutome stage
40 | 916 | Oscillation check - Measure the amplitude of the oscillation
40 | 916 | Perform X-ray check
40 | 916 | Pump Down Time
40 | 916 | Ultimate Vacuum
40 | 916 | Focal CC fitted to Volutome stage
40 | 916 | Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar
40 | 916 | Generate live image from Axiocam
40 | 916 | Re-package Volutome and check it's completeness
40 | 916 | ManEng/ Senior Eng/ Production Manger Sign off
2 | 10 | Record AC unit S/N
2 | 10 | Qualification Marks Checked
2 | 10 | Fit mounting bracket 'Complete Holding Angle'
382 | 6928 | Macro Result 1 ($[StageRepeatability])
382 | 6928 | Macro Result 2 ($[StageRepeatability])
382 | 6928 | Macro Result 3 ($[StageRepeatability])
382 | 6928 | Macro Result 1 ($[StageRepeatability])
382 | 6928 | Macro Result 2 ($[StageRepeatability])
382 | 6928 | Macro Result 3 ($[StageRepeatability])
382 | 6928 | Image of Macro Results
2 | 10 | Record AC unit S/N
2 | 10 | Qualification Marks Checked
2 | 10 | Fit mounting bracket to AC unit 'Angle Bracket' - 349500-8040-000
2 | 10 | Set up mounting bracket height using Jig 356580-0174-000
382 | 6928 | Macro Result 1 ($[StageRepeatability])
382 | 6928 | Macro Result 2 ($[StageRepeatability])
382 | 6928 | Macro Result 3 ($[StageRepeatability])
382 | 6928 | Image of Macro Results
382 | 6928 | Macro Result 1 ($[StageRepeatability])
382 | 6928 | Macro Result 2 ($[StageRepeatability])
382 | 6928 | Macro Result 3 ($[StageRepeatability])
382 | 6928 | Image of Macro Results
30 | 299 | Macro Result 1 ($[StageRepeatability])
30 | 299 | Macro Result 2 ($[StageRepeatability])
30 | 299 | Macro Result 3 ($[StageRepeatability])
30 | 299 | Image of Macro Results
12 | 162 | Tip Re-measured Height Avg: Z =
12 | 162 | Actual Tip Height: Y - Z =
12 | 162 | Tip Gap: Actual Tip Height – 0.75 =
12 | 162 | Final Shim Value (Please ensure to always round up the shim value) =
12 | 162 | Denka Torque Setting =
40 | 916 | Unpack Volutome and check all parts on the packing list are present (packing list link can be found in attachments)
40 | 916 | Ensure Arivis Software Licence is available (410145-1031-100)
40 | 916 | Upload data from USB stick to the O drive product folder
40 | 916 | Remove transit clamps from Volutome stage
40 | 916 | Ensure hardware has been configured according to work instructions
40 | 916 | Fit Volutome stage
40 | 916 | Connect Volutome stage and controller to the SEM
40 | 916 | Install Volutome software installation package
40 | 916 | Configure Volutome stage motors
40 | 916 | Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator
40 | 916 | Set up the Volutome workbench in Zen
40 | 916 | Record VolumeBSD serial number
40 | 916 | Record Volutome stage serial number
40 | 916 | Record VolumeBSD amplifier serial number
40 | 916 | Record VolumeBSD mechanics serial number
40 | 916 | Record C-885 PI controller serial number
40 | 916 | Record E-727 PI controller serial number
40 | 916 | Function check the Volutome stage light
40 | 916 | Adjust Volutome stage centre (Crosshairs should be within 4 central squares)
40 | 916 | Install Nav cam in-line connections (If applicable)
40 | 916 | Perform the microtome and stage functionality checks
40 | 916 | Fit and align the VolumeBSD diode
40 | 916 | Check VolumeBSD alignment repeatability
40 | 916 | Check all gain ranges of the VolumeBSD (normal/inverted)
40 | 916 | Perform Signal/noise ratio check
40 | 916 | Ensure BSD diode protector is fitted and adjusted in chamber
40 | 916 | Install cable tie to loop BSD signal cable
40 | 916 | Check SCM with Volutome stage
40 | 916 | Oscillation check - Measure the amplitude of the oscillation
40 | 916 | Perform X-ray check
40 | 916 | Pump Down Time
40 | 916 | Ultimate Vacuum
40 | 916 | Focal CC fitted to Volutome stage
40 | 916 | Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar
40 | 916 | Generate live image from Axiocam
40 | 916 | Re-package Volutome and check it's completeness
40 | 916 | ManEng/ Senior Eng/ Production Manger Sign off
26 | 281 | Check concentricity of ESB wrt scanning system.
26 | 281 | Record the ESB serial number - This is the S/N that is shown on the ESB label (1026-XXX or 1125-XXX).
26 | 281 | Measure the ESB hole diameter.
26 | 281 | Store ESB Concentricity image. Add ESB serial number as a text annotation in SmartSEM. (filename must be exactly esbdiameter.tif)
26 | 396 | Check concentricity of ESB wrt scanning system. (Limit: <= 50um)
26 | 396 | Measure the ESB hole diameter.
26 | 396 | Record the ESB serial number - This is the S/N that is shown on the ESB label (1026-XXX or 1125-XXX).
26 | 396 | Store ESB Concentricity image. Add ESB serial number as a text annotation in SmartSEM. (filename must be exactly esbdiameter.tif)
44 | 11090 | All Warning labels Fitted.
382 | 6928 | Macro Result 1 ($[StageRepeatability])
382 | 6928 | Macro Result 2 ($[StageRepeatability])
382 | 6928 | Macro Result 3 ($[StageRepeatability])
382 | 6928 | Image of Macro Results
30 | 299 | Macro Result 1 ($[StageRepeatability])
30 | 299 | Macro Result 2 ($[StageRepeatability])
30 | 299 | Macro Result 3 ($[StageRepeatability])
30 | 299 | Image of Macro Results
41 | 1244 | All Warning labels Fitted.
389 | 917 | Unpack Volutome and check all parts on the packing list are present (packing list link can be found in attachments)
389 | 917 | Ensure Arivis Software Licence is available (410145-1031-100)
389 | 917 | Remove transit clamps from Volutome stage
389 | 917 | Ensure hardware has been configured according to work instructions
389 | 917 | Fit Volutome stage
389 | 917 | Connect Volutome stage and controller to the SEM
389 | 917 | Install Volutome software installation package
389 | 917 | Configure Volutome stage motors
389 | 917 | Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator
389 | 917 | Set up the Volutome workbench in Zen
389 | 917 | Record VolumeBSD serial number
389 | 917 | Record Volutome stage serial number
389 | 917 | Record VolumeBSD amplifier serial number
389 | 917 | Record VolumeBSD mechanics serial number
389 | 917 | Record C-885 PI controller serial number
389 | 917 | Record E-727 PI controller serial number
389 | 917 | Function check the Volutome stage light
389 | 917 | Adjust Volutome stage centre (Crosshairs should be within 4 central squares)
389 | 917 | Install Nav cam in-line connections (If applicable)
389 | 917 | Perform the microtome and stage functionality checks
389 | 917 | Sigma cartesian stage angle correction completed
389 | 917 | Fit and align the VolumeBSD diode
389 | 917 | Check VolumeBSD alignment repeatability
389 | 917 | Check all gain ranges of the VolumeBSD (normal/inverted)
389 | 917 | Perform Signal/noise ratio check
389 | 917 | Ensure BSD diode protector is fitted and adjusted in chamber
389 | 917 | Install cable tie to loop BSD signal cable
389 | 917 | Check SCM with Volutome stage
389 | 917 | Oscillation check - Measure the amplitude of the oscillation
389 | 917 | Perform X-ray check
389 | 917 | Pump Down Time
389 | 917 | Ultimate Vacuum
389 | 917 | Focal CC fitted to Volutome stage
389 | 917 | Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar
389 | 917 | Generate live image from Axiocam
389 | 917 | Re-package Volutome and check it's completeness
389 | 917 | ManEng/ Senior Eng/ Production Manger Sign off
4 | 43 | Fit the Nitrogen vent pipe to TIV vent valve
388 | 11425 | 15kv 0.9nm
388 | 11425 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 1kv 1.3nm
388 | 11425 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11425 | 500V 2nm
388 | 11425 | Store 500V Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 15kv 0.7nm
388 | 11426 | Store 15KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 1kv 1.2nm
388 | 11426 | Store 1KV Resolution Images and Pate image for 2 x consecutive images
388 | 11426 | 500V 1.9nm
388 | 11426 | Store 500V Resolution Images and Pate image for 2 x consecutive images
30 | 299 | Macro Result 1 ($[StageRepeatability])
30 | 299 | Macro Result 2 ($[StageRepeatability])
30 | 299 | Macro Result 3 ($[StageRepeatability])
30 | 299 | Image of Macro Results
389 | 1245 | Stage mapping completed successfuly
389 | 1245 | Stage mapping completed successfuly
30 | 1230 | Stage mapping completed successfuly
4 | 43 | Fit the Nitrogen vent pipe to TIV vent valve
382 | 1225 | carry out scanning calibration (Hasap only)
30 | 1224 | carry out scanning calibration (Hasap only)
30 | 303 | Actual (Spec: 6.7nm / min)
30 | 303 | Store stage drift macro results image
645 | 1226 | carry out scanning calibration (Hasap only)
645 | 1226 | carry out scanning calibration (Hasap only)
30 | 303 | Actual (Spec: 6.7nm / min)
30 | 303 | Store stage drift macro results image
30 | 303 | Actual (Spec: 6.7nm / min)
30 | 303 | Store stage drift macro results image
30 | 1230 | Stage mapping completed successfuly
2 | 5 | Fit the pre-assembled Air and Water assembly
30 | 299 | Macro Result 1 ($[StageRepeatability])
30 | 299 | Macro Result 2 ($[StageRepeatability])
30 | 299 | Macro Result 3 ($[StageRepeatability])
30 | 299 | Image of Macro Results
382 | 6931 | Drift
382 | 6928 | Macro Result 1 ($[StageRepeatability])
382 | 6928 | Macro Result 2 ($[StageRepeatability])
382 | 6928 | Macro Result 3 ($[StageRepeatability])
382 | 6928 | Image of Macro Results
389 | 89 | Test Valves Using Raman Test Interlock Box
389 | 89 | Install SemRamanStage Software
389 | 89 | Run SEM Raman Correlation
389 | 89 | Save Calibration File as Raman
389 | 89 | second chamberscope fitted
389 | 89 | Run Smart Stage Mapping
2 | 9 | Set water flow pressure to 0.9L / min
2 | 6 | Fit hose
2 | 6 | Set regulator pressure to 0.03MPa
2 | 6 | Fit hose
2 | 6 | Set regulator pressure to 0.03MPa
5 | 1255 | Apply label to In-Lens And ESB cable
2 | 6 | Fit hoses # 1,2,3,4,5,7,9
2 | 6 | Set regulator pressure to 0.03MPa
2 | 6 | Fit nitrogen hoses # 1,2,3,4,5,7,9
2 | 6 | Set regulator pressure to 0.03MPa
2 | 1247 | Fit Stage PCB
2 | 1247 | Fit IPPSU 1, Record S/N
2 | 1247 | Set DIP switches
2 | 1247 | Qualification Marks Checked
2 | 7 | Fit Stage PCB
2 | 7 | Fit IPPSU 1, Record S/N
2 | 7 | Set DIP switches
2 | 7 | Qualification Marks Checked
2 | 1247 | Fit Stage PCB
2 | 1247 | Fit IPPSU 1, Record S/N
2 | 1247 | Set DIP switches
2 | 1247 | Qualification Marks Checked
2 | 1247 | Fit Stage PCB
2 | 1247 | Fit IPPSU 1, Record S/N
2 | 1247 | Set DIP switches
2 | 1247 | Qualification Marks Checked
2 | 1247 | Fit Stage PCB
2 | 1247 | Fit IPPSU 1, Record S/N
2 | 1247 | Set DIP switches
2 | 1247 | Qualification Marks Checked
2 | 7 | Fit Stage PCB
2 | 7 | Fit IPPSU 1, Record S/N
2 | 7 | Set DIP switches
2 | 7 | Qualification Marks Checked
2 | 7 | Fit Stage PCB
2 | 7 | Fit IPPSU 1, Record S/N
2 | 7 | Set DIP switches
2 | 7 | Qualification Marks Checked
2 | 7 | Fit Stage PCB
2 | 7 | Fit IPPSU 1, Record S/N
2 | 7 | Set DIP switches
2 | 7 | Qualification Marks Checked
2 | 8 | Record IPPSU S/N
2 | 8 | Set DIP switches
2 | 8 | Qualification Marks Checked
2 | 1218 | Check if ASAP or h-ASAP
2 | 1218 | h-ASAP Serial No (if applicable, if not applicable enter NA)
2 | 1218 | Check if ASAP or H-ASAP
2 | 1218 | h-ASAP Serial No (if applicable, if not applicable enter NA)
2 | 1218 | Check if ASAP or H-ASAP
2 | 1218 | h-ASAP Serial No (if applicable, if not applicable enter NA)
2 | 1218 | Check if ASAP or H-ASAP
2 | 1218 | h-ASAP Serial No (if applicable, if not applicable enter NA)
2 | 1218 | Check if ASAP or H-ASAP
2 | 1218 | h-ASAP Serial No (if applicable, if not applicable enter NA)
2 | 1218 | If present fit
2 | 1218 | Check if ASAP or H-ASAP
2 | 1218 | h-ASAP Serial No (if applicable, if not applicable enter NA)
2 | 1218 | Fit additional ASAP cards if provided.
2 | 1218 | Check if ASAP or H-ASAP
2 | 1218 | h-ASAP Serial No (if applicable, if not applicable enter NA)
2 | 1218 | Fit additional ASAP cards if provided.
2 | 1218 | Check if ASAP or H-ASAP
2 | 1218 | h-ASAP Serial No (if applicable, if not applicable enter NA)
2 | 1218 | Fit additional ASAP cards if provided (DFE / VBE Cards)
2 | 1218 | Check if ASAP or H-ASAP
2 | 1218 | h-ASAP Serial No (if applicable, if not applicable enter NA)
2 | 1218 | Fit additional ASAP cards if provided (DFE / VBE Cards)
2 | 1218 | Check if ASAP / H-ASAP / SVSEM
2 | 1218 | h-ASAP Serial No (if applicable, if not applicable enter NA)
2 | 1218 | Fit additional ASAP cards if provided (DFE / VBE Cards)
2 | 1218 | Check if ASAP / H-ASAP / SVSEM
2 | 1218 | H-ASAP Serial No (if applicable, if not applicable enter NA)
2 | 1218 | Fit additional ASAP cards if provided (DFE / VBE Cards)
2 | 1218 | Check if ASAP / H-ASAP / SVSEM
2 | 1218 | Serial No
2 | 1218 | Fit additional ASAP cards if provided (DFE / VBE Cards)
2 | 1218 | Check if ASAP / H-ASAP / SVSEM
2 | 1218 | Record Serial No
2 | 1218 | Fit additional ASAP cards if provided (DFE / VBE Cards)
2 | 17 | Fit X4 braided earth straps to rear PCB's. (Vac, EOS, EOD, ASAP /SVSEM)
2 | 17 | Fit X4 braided earth straps to rear PCB's. (Vac, EOS, EOD, ASAP /SVSEM)
2 | 17 | Fit earth cable to DC unit.
2 | 17 | Fit X4 braided earth straps to rear PCB's. (Vac, EOS, EOD, ASAP /SVSEM)
2 | 17 | Fit earth cable to DC unit.
2 | 21 | Set Valve Screw torque IN (3 turns)
2 | 21 | Set Valve Screw torque OUT (3 turns)
2 | 21 | Set foot height
2 | 21 | Install Compressed Air Line
2 | 21 | Install Control Cable
2 | 22 | Set Valve Screw torque IN (1.5 turns)
2 | 22 | Set Valve Screw torque OUT(3 turns)
2 | 22 | Set foot height
2 | 22 | Install Compressed Air Line
2 | 22 | Install Control Cable
2 | 23 | Set Valve Screw torque IN (1.5 turns)
2 | 23 | Set Valve Screw torque OUT(3 turn)
2 | 23 | Set foot height
2 | 23 | Install Compressed Air Line
2 | 23 | Install Control Cable
2 | 20 | Install Compressed Air Line
2 | 20 | Set foot height
2 | 20 | Install Compressed Air Line
2 | 20 | Set foot height
2 | 21 | Set Valve Screw torque IN (3 turns)
2 | 21 | Set Valve Screw torque OUT (3 turns)
2 | 21 | Set foot height
2 | 21 | Install Compressed Air Line
2 | 21 | Install Control Cable
2 | 21 | Set Valve Screw torque IN (3 turns)
2 | 21 | Set Valve Screw torque OUT (3 turns)
2 | 21 | Set foot height
2 | 21 | Install Compressed Air Line
2 | 21 | Install Control Cable
2 | 22 | Set Valve Screw torque IN (1.5 turns)
2 | 22 | Set Valve Screw torque OUT(3 turns)
2 | 22 | Set foot height
2 | 22 | Install Compressed Air Line
2 | 22 | Install Control Cable
2 | 22 | Set Valve Screw torque IN (1.5 turns)
2 | 22 | Set Valve Screw torque OUT(3 turns)
2 | 22 | Set foot height
2 | 22 | Install Compressed Air Line
2 | 22 | Install Control Cable
2 | 23 | Set Valve Screw torque IN (1.5 turns)
2 | 23 | Set Valve Screw torque OUT(3 turn)
2 | 23 | Set foot height
2 | 23 | Install Compressed Air Line
2 | 23 | Install Control Cable
2 | 21 | Set Valve Screw torque IN (3 turns)
2 | 21 | Set Valve Screw torque OUT (3 turns)
2 | 21 | Set foot height
2 | 21 | Install Compressed Air Line
2 | 21 | Install Control Cable
2 | 22 | Set Valve Screw torque IN (1.5 turns)
2 | 22 | Set Valve Screw torque OUT(3 turns)
2 | 22 | Set foot height
2 | 22 | Install Compressed Air Line
2 | 22 | Install Control Cable
2 | 23 | Set Valve Screw torque IN (1.5 turns)
2 | 23 | Set Valve Screw torque OUT(3 turn)
2 | 23 | Set foot height
2 | 23 | Install Compressed Air Line
2 | 23 | Install Control Cable
2 | 23 | Set Valve Screw torque IN (1.5 turns)
2 | 23 | Set Valve Screw torque OUT(3 turn)
2 | 23 | Set foot height
2 | 23 | Install Compressed Air Line
2 | 23 | Install Control Cable
4 | 43 | Fit the Nitrogen vent pipe to TIV vent valve
2 | 22 | Set Valve Screw torque IN (1.5 turns)
2 | 22 | Set Valve Screw torque OUT(3 turns)
2 | 22 | Set foot height
2 | 22 | Install Compressed Air Line
2 | 22 | Install Control Cable
4 | 39 | Turbo Pump S/No.
4 | 39 | Qualification Marks Checked
84 | 1601 | Ventilation of chamber possible. Confirm Gun does not switch off.
84 | 1601 | While system is vented check touch alarm works.
84 | 1601 | Inspect inside of chamber using UV light and goggles (ME Trolley). If contamination found please raise deviation and clean.
46 | 1567 | Remove the sample holder from the system.
46 | 1567 | Remove Docking Brackets and fit Transit clamps.
46 | 1567 | Airfeet deflated and free to move.
46 | 1567 | Backing lines re-configured so that PVSV valve is attached to flexi hose. Secure valve to frame.
46 | 1567 | Confirm that the 3 'o' rings are fitted to the flexi hose backing line
46 | 1567 | If applicable, ensure that the A/L push rod is removed and blank fitted in its place. Ensure A/L is pumped down before powering down system.
46 | 1567 | If applicable, ensure that the Beamsleeve funnel has been removed from the detector
46 | 1567 | Select 'Partial vent on standby' in Smartsem.
30 | 299 | Macro Result 1 ($[StageRepeatability])
30 | 299 | Macro Result 2 ($[StageRepeatability])
30 | 299 | Macro Result 3 ($[StageRepeatability])
30 | 299 | Image of Macro Results
30 | 299 | Macro Result 1 ($[StageRepeatability])
30 | 299 | Macro Result 2 ($[StageRepeatability])
30 | 299 | Macro Result 3 ($[StageRepeatability])
30 | 299 | Image of Macro Results
30 | 1230 | Stage mapping completed successfuly
4 | 42 | Fit flexi pipe to turbo
4 | 42 | Fit 3X isolation O-rings to flexi pipe
4 | 39 | Turbo Pump S/No.
4 | 39 | Qualification Marks Checked
703 | 11017 | Time =
703 | 11017 | Date =
703 | 11017 | Time =
703 | 11017 | Date =
703 | 11016 | Time =
703 | 11016 | Date =
5 | 10992 | Add immersion oil to PM glass
5 | 10993 | Add immersion oil to PM glass
5 | 10992 | Add 1 drop of immersion oil to PM glass
5 | 10993 | Apply 1 drop immersion oil to PM glass
5 | 10992 | Apply 1 drop of immersion oil to PM glass
5 | 1544 | Fit A/L Key to Stage Door
5 | 1544 | Check distance of 200mm door extended = 228mm
5 | 10959 | Warning - ONLY VENT using the electronic vent valve in the chamber
5 | 10959 | No Leaks Detected (Prior to running up Turbo)
5 | 10959 | No Leaks Detected (With Turbo running) when in high sensitivity
5 | 10959 | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng
5 | 10959
5 | 10959 | No Leaks Detected (Prior to running up Turbo)
5 | 10959 | No Leaks Detected (With Turbo running) when in high sensitivity
5 | 10959 | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng
5 | 10959 | No Leaks Detected (Prior to running up Turbo)
5 | 10959 | No Leaks Detected (With Turbo running) when in high sensitivity
5 | 10959 | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng
5 | 10959 | No Leaks Detected (Prior to running up Turbo)
5 | 10959 | No Leaks Detected (With Turbo running) when in high sensitivity
5 | 10959 | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng
5 | 10959 | No Leaks Detected (With Turbo running) when in high sensitivity
5 | 10959 | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng
5 | 52 | Fit Ducting Plate & Earth Cable
5 | 52 | Fit Manifold
5 | 52 | Fit Air and Nitrogen Pipes
5 | 11274 | Fit and route all column cables
5 | 11274 | Fit and route all column cables
5 | 11274 | Fit and route all column cables
5 | 11274 | Fit and route all column cables
5 | 11274 | Safety – Torque the earth stud to 15 N.m
5 | 11274 | Torque both earth zinc plated M8 nuts to 15 N.m
5 | 11274 | Fit and route all column cables
5 | 11274 | Safety – Torque the earth stud to 15 N.m
5 | 11274 | Torque both earth zinc plated M8 nuts to 15 N.m
5 | 52 | Fit Ducting Plate & Earth Cable
5 | 52 | Fit Manifold
5 | 52 | Fit Air and Nitrogen Pipes
5 | 11274 | Fit and route all column cables
5 | 11274 | Safety – Torque the earth stud to 15 N.m
5 | 11274 | Torque both earth zinc plated M8 nuts to 15 N.m
5 | 11274 | Fit and route all column cables
5 | 11274 | Safety – Torque the earth stud to 15 N.m
5 | 11274 | Torque both earth zinc plated M8 nuts to 15 N.m
5 | 52 | Fit Ducting Plate & Earth Cables
5 | 52 | Fit Manifold
5 | 52 | Fit Air and Nitrogen Pipes
5 | 11274 | Fit and route all column cables
5 | 11274 | Safety – Torque the earth stud to 15 N.m
5 | 11274 | Torque both earth zinc plated M8 nuts to 15 N.m
12 | 141 | Fit UHV housing to column - M4x16 screws torqued to 2.5Nm
648 | 10553 | 30KV Res- ≤2.2nm
648 | 10553 | 3KV Res- ≤7.5nm - If Unable to Achieve Contact Man Eng.
648 | 10553 | 1KV Res- ≤14nm - If Unable To Achieve Contact Man Eng.
648 | 10553 | 200V Res- Not measured
648 | 10553 | Store Standard Resolution Images
648 | 10553 | 30KV Res- ≤2.2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 3KV Res- ≤7.5nm - If Unable to Achieve Contact Man Eng.
648 | 10553 | 1KV Res- ≤14nm - If Unable To Achieve Contact Man Eng.
648 | 10553 | 200V Res- Not measured
648 | 10553 | Store Standard Resolution Images
648 | 10553 | 30KV Res- ≤2.2nm
648 | 10553 | 3KV Res- ≤7.5nm - If Unable to Achieve Contact Man Eng.
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 1KV Res- ≤14nm - If Unable To Achieve Contact Man Eng.
648 | 10553 | 1KV Res- ≤15nm
648 | 10553 | 200V Res- Not measured
648 | 10553 | Store Standard Resolution Images
648 | 10553 | 30KV Res- ≤2.2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 3KV Res- ≤7.5nm - If Unable to Achieve Contact Man Eng.
648 | 10553 | 1KV Res- ≤15nm
648 | 10553 | 1KV Res- ≤14nm - If Unable To Achieve Contact Man Eng.
648 | 10553 | 200V Res- Not measured
648 | 10553 | Store Standard Resolution Images
648 | 10553 | 30KV Res- ≤2.2nm
648 | 10553 | 3KV Res- ≤8nm
648 | 10553 | 3KV Res- ≤7.5nm - If Unable to Achieve Contact Man Eng.
648 | 10553 | 1KV Res- ≤15nm
648 | 10553 | 1KV Res- ≤14nm - If Unable To Achieve Contact Man Eng.
648 | 10553 | 200V Res- Not measured
648 | 10553 | Store Standard Resolution Images
12 | 141 | Fit UHV housing to column - M4x16 screws torqued to 2.5Nm
12 | 141 | Fit UHV bucket to housing -BN30 grease applied to M4x25 screws and torqued to 2Nm
12 | 141 | Fit UHV housing to column - M4x16 screws torqued to 2.5Nm
12 | 141 | Fit UHV bucket to housing - BN30 grease applied to M4x25 screws and torqued to 2Nm
46 | 1567 | Inspect inside of chamber using UV light and goggles (ME Trolley). If contamination found please raise deviation and clean.
46 | 1567 | Remove the sample holder from the system.
46 | 1567 | Remove Docking Brackets and fit Transit clamps.
46 | 1567 | Airfeet deflated and free to move.
46 | 1567 | Backing lines re-configured so that PVSV valve is attached to flexi hose. Secure valve to frame.
46 | 1567 | Confirm that the 3 'o' rings are fitted to the flexi hose backing line
46 | 1567 | If applicable, ensure that the A/L push rod is removed and blank fitted in its place. Ensure A/L is pumped down before powering down system.
46 | 1567 | If applicable, ensure that the Beamsleeve funnel has been removed from the detector
46 | 1567 | Select 'Partial vent on standby' in Smartsem.
30 | 299 | Macro Result 1 ($[StageRepeatability])
30 | 299 | Macro Result 2 ($[StageRepeatability])
30 | 299 | Macro Result 3 ($[StageRepeatability])
30 | 299 | Image of Macro Results
676 | 10826 | Select Stage Type
676 | 10826 | Serial Number
676 | 10826 | Shim Stage Door
676 | 10826 | Set Stage Clamp Height - Using Jig
382 | 6927 | Select Stage Type
382 | 6927 | Macro run time
382 | 6927 | Macro run time
40 | 1251 | Resolution Result
6 | 60 | No Pressure loss in 2mins
382 | 6928 | Select Stage Type
382 | 6928 | Macro Result 1 ($[StageRepeatability])
382 | 6928 | Macro Result 2 ($[StageRepeatability])
382 | 6928 | Macro Result 3 ($[StageRepeatability])
382 | 6928 | Image of Macro Results
665 | 10734 | Select Stage Type
665 | 10734 | Stage Serial Number
665 | 10734 | Shim Stage Door
665 | 10734 | Set Stage Clamp Height
645 | 1250 | Select Stage Type
645 | 1250 | Macro Result 1 ($[StageRepeatability])
645 | 1250 | Macro Result 2 ($[StageRepeatability])
645 | 1250 | Macro Result 3 ($[StageRepeatability])
645 | 1250 | Image of Macro Results
645 | 1250 | Select Stage Type
645 | 1250 | Macro Result 1 ($[StageRepeatability])
645 | 1250 | Macro Result 2 ($[StageRepeatability])
645 | 1250 | Macro Result 3 ($[StageRepeatability])
645 | 1250 | Image of Macro Results
645 | 10516 | X Axis
645 | 10516 | Y Axis
645 | 10516 | X+Y Axis
645 | 10516 | Y+X Axis
645 | 10516 | 10mm Square
645 | 10517 | Z Axis
645 | 10518 | X Axis
645 | 10518 | Y Axis
21 | 12567 | Enter the FIB Extension number using EEPromUpdater.exe
21 | 12567 | Record FIB Extension Serial No.
30 | 299 | Macro Result 1
30 | 299 | Macro Result 2
30 | 299 | Macro Result 3
30 | 299 | Image of Macro Results
6 | 1252 | Warning -ONLY VENT using the electronic vent valve in the chamber
6 | 1252 | Check interlock keys fitted to door
6 | 1252 | Check vented screws fitted to inside of chamber ceiling
6 | 1252 | Check magnet clamp set correctly
6 | 1252 | Check door alignment
6 | 61 | Safety Earth
6 | 61 | Compressed Air (Set to 0.5MPa)
6 | 61 | PVP Valve
6 | 61 | Power
6 | 61 | Rotary Pump
6 | 61 | Nitrogen (Set to 0.05 MPa)
6 | 61 | Cooling Water
6 | 61 | Set Water Flow switch to 0.9 LPM
6 | 61 | EHT Set Serial Number SN:
6 | 61 | EHT Lens output resistance ( < 2.2kΩ)
6 | 61 | Switch on system breaker and run system to green button
6 | 61 | Safety Earth
6 | 61 | Compressed Air (Set to 0.5MPa)
6 | 61 | PVP Valve
6 | 61 | Power
6 | 61 | Rotary Pump
6 | 61 | Nitrogen (Set to 0.05 MPa)
6 | 61 | Cooling Water
6 | 61 | Set Water Flow switch to 0.9 LPM
6 | 61 | EHT Lens output resistance ( < 2.2kΩ)
6 | 61 | Switch on system breaker and run system to green button
6 | 61 | Safety Earth
6 | 61 | Compressed Air (Set to 0.5MPa)
6 | 61 | PVP Valve.
6 | 61 | Power - 32A Connected to wall and switch ON.
6 | 61 | Rotary Pump IEC Lead
6 | 61 | Nitrogen (Set to 0.05 MPa inside plinth)
6 | 61 | Cooling Water
6 | 61 | Set Water Flow switch to 0.9 LPM
6 | 61 | EHT Lens output resistance ( < 2.2kΩ)
6 | 61 | Switch on system breaker and run system to green button
6 | 61 | Safety Earth
6 | 61 | Compressed Air (Set to 0.5MPa)
6 | 61 | PVP Valve.
6 | 61 | Power - 32A Connected to wall and switch ON.
6 | 61 | Rotary Pump IEC Lead
6 | 61 | Nitrogen (Set to 0.05 MPa inside plinth)
6 | 61 | Cooling Water - Confirm flow is set to 0.9LPM
6 | 61 | EHT Lens output resistance ( < 2.2kΩ)
6 | 61 | Switch on system breaker and run system to green button
6 | 61 | Safety Earth
6 | 61 | Compressed Air (Set to 0.5MPa)
6 | 61 | PVP Valve.
6 | 61 | Power - 32A Connected to wall and switch ON.
6 | 61 | Rotary Pump IEC Lead
6 | 61 | Nitrogen (Set to 0.05 MPa inside plinth)
6 | 61 | Cooling Water - Confirm flow is set to 0.9LPM
6 | 61 | EHT Lens output resistance ( < 2.2kΩ)
6 | 61 | Switch on system breaker and run system to green button
6 | 61 | Safety Earth
6 | 61 | Compressed Air (Set to 0.5MPa)
6 | 61 | PVP Valve.
6 | 61 | Power - 32A Connected to wall and switch ON.
6 | 61 | Rotary Pump IEC Lead
6 | 61 | Nitrogen (Set to 0.05 MPa inside plinth)
6 | 61 | Cooling Water - Confirm flow is set to 0.9LPM
6 | 61 | EHT Lens output resistance (Safety Test) ( < 2.2kΩ)
6 | 61 | Switch on system breaker and run system to green button
2 | 1264 | Fit Stage PCB
2 | 1264 | Set Stage links (348247-9002-120)
2 | 1264 | Fill in Jumper settings label
3 | 1265 | Fit Stage PCB
3 | 1265 | Set Stage links (348247-9001-500)
3 | 1265 | Fill in Jumper settings label
3 | 1266 | Set Stage links (348247-9002-120)
3 | 1266 | Fill in Jumper settings label
2 | 7 | Fit Stage PCB
3 | 1267 | Set Stage links (348247-9002-120)
3 | 1267 | Fill in Jumper settings label
659 | 11399 | Fit Quiet Mode PCB
659 | 11399 | Fit Cable X45
659 | 11399 | Fit Cable X11
21 | 1268 | Vacuum System is Configured correctly
21 | 1268 | Vacuum System is Configured correctly
6 | 63 | Confirm all vacuum modules are Green in Vac Test
6 | 63 | Confirm Turbo reached 1000Hz and penning reading
6 | 63 | Confirm IGP's haev switched on without trip (system vac better than 2x10-5mbar)
2 | 1264 | Fit Stage PCB
2 | 1264 | Set Stage links (348247-9002-120)
2 | 1264 | Fill in Jumper settings label
2 | 7 | Fit Stage PCB
2 | 7 | Set Stage links (348247-9001-500)
2 | 7 | Fill in Jumper settings label
3 | 1266 | Fit Stage PCB
3 | 1266 | Set Stage links (348247-9002-120)
3 | 1266 | Fill in Jumper settings label
2 | 7 | Fit Stage PCB
2 | 1264 | Fit Stage PCB
3 | 1267 | Reconfigure loght barries for XB-550L Stage.
8 | 1274 | 20 Min plasma clean - Ceramic disc, Bean Can, Gehause Mont
6 | 1253 | Run latest system checker and confirm all FW is correct (Update any FW not latest verion)
6 | 1253 | Run latest system checker and confirm all FW is correct (Update any FW not latest version)
3 | 1265 | Fit Stage PCB
3 | 1265 | Set Stage links (348247-9001-500)
3 | 1265 | Fill in Jumper settings label
6 | 63 | Confirm Vacuum System is configured correctly
6 | 63 | Confirm all vacuum modules are Green in Vac Test
6 | 63 | Confirm Turbo reached 1000Hz and penning reading
6 | 63 | Confirm IGP's haev switched on without trip (system vac better than 2x10-5mbar)
32 | 1269 | System is ready for QA Engineer to perform Interim QA
32 | 1269 | System is ready for QA Engineer to perform Interim QA
667 | 10765 | Connect safety check rig (354993-9205-000) and confirm no LED's lit.
667 | 10765 | HAL SN:
667 | 10765 | HAL Calibration Date:
667 | 10765 | DC IR (Direct Current Insulation Resistance @500V) (>10.0MΩ)
667 | 10765 | DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
667 | 10765 | Confirm System Passed Hi-Pot Tests
667 | 1270 | HAL SN:
667 | 1270 | HAL Calibration Date:
667 | 1270 | Column Heater (354920-9179-000). . . . . . . . . . . . .. .. . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . . DC HIPOT (Direct Current High Potential @2000V) (<10.0 mA)
20 | 11021 | PC Serial Number:
20 | 11021 | Connect Up OEM Test Tool PC
20 | 11021 | Unistall SmartSEM
20 | 11021 | Install SmartSEM Version
20 | 11021 | Install Latest Service Pack
20 | 11021 | Install Gemini Support Tools-Version
20 | 11021 | Run ASAP Screening Program
20 | 11021 | System Specific Datazone Loaded
645 | 1226 | carry out scanning calibration (Hasap only)
6 | 69 | Foot Height
6 | 69 | Pipe Configuration
6 | 69 | Sealing Plugs Removed
4 | 1254 | Fit Ducting Plate & Earth Cables
4 | 1254 | Fit Manifold
4 | 1254 | Fit Air and Nitrogen Pipes
5 | 1255 | Apply label to In-Lens And ESB cable
5 | 1255 | Apply label to IGP cables (IGP1 & IGP2)
5 | 1255 | Apply label to SCM cable
5 | 1255 | Apply label to In-Lens And ESB cable
5 | 1255 | Apply label to IGP cables (IGP1 & IGP2)
5 | 1255 | Apply label to SCM cable
5 | 1255 | In-Lens And ESB cable
5 | 1255 | IGP cables (IGP1 & IGP2)
5 | 1255 | SCM cable
5 | 1255 | In-Lens And ESB cable
5 | 1255 | IGP cables (IGP1 & IGP2)
5 | 1255 | SCM cable
5 | 1255 | In-Lens And ESB cable
5 | 1255 | IGP cables (IGP1 & IGP2)
5 | 1255 | SCM cable
5 | 1255 | CCD camera cables
5 | 11274 | Fit and route all column cables
5 | 11274 | Safety – Torque the earth stud to 15 N.m
5 | 11274 | Torque both earth zinc plated M8 nuts to 15 N.m
5 | 11274 | Fit and route all column cables
5 | 11274 | Safety – Torque the earth stud to 15 N.m
5 | 11274 | Torque both earth zinc plated M8 nuts to 15 N.m
5 | 1255 | In-Lens And ESB cable
5 | 1255 | IGP cables (IGP1 & IGP2)
5 | 1255 | SCM cable
5 | 1255 | CCD camera cables
5 | 11274 | Fit and route all column cables
5 | 11274 | Fit and route objective water pipes
5 | 11274 | Safety – Torque the earth stud to 15 N.m
5 | 11274 | Torque both earth zinc plated M8 nuts to 15 N.m
5 | 11274 | Fit and route all column cables
5 | 11274 | Fit and route objective water pipes
5 | 11274 | Safety – Torque the earth stud to 15 N.m
5 | 11274 | Torque both earth zinc plated M8 nuts to 15 N.m
5 | 11274 | Fit and route all column cables
5 | 11274 | Fit and route objective water pipes
5 | 11274 | Safety – Torque the earth stud to 15 N.m
5 | 11274 | Torque both earth zinc plated M8 nuts to 15 N.m
5 | 11274 | Fit and route all column cables
5 | 11274 | Fit and route objective water pipes
5 | 11274 | Safety – Torque the earth stud to 15 N.m
5 | 11274 | Torque both earth zinc plated M8 nuts to 15 N.m
5 | 11274 | Fit and route all column cables
5 | 11274 | Fit and route objective water pipes
5 | 11274 | Safety – Torque the earth stud to 15 N.m
5 | 11274 | Torque both earth zinc plated M8 nuts to 15 N.m
5 | 58 | Configuration Check
5 | 58 | Port configuration Check
5 | 10991 | Fit Indigo Filter
5 | 1259 | Configure ASAP / SVSEM detectors ports as specified in procedure.
5 | 1261 | Fit Nitrogen purge valve to turbo.
5 | 1261 | Fit Nitrogen vent line.
5 | 1261 | Fit vent cable.
2 | 7 | Fit Stage PCB
2 | 7 | Set Stage links
2 | 7 | Fill in Jumper settings label
2 | 7 | Fit IPPSU 1, Record S/N
2 | 7 | Set DIP switches
2 | 7 | Qualification Marks Checked
397 | 11075 | Pneumatics functional
397 | 11075 | SmartSEM Control tested
397 | 11075 | Images saved
2 | 7 | Fit Stage PCB
2 | 7 | Set Stage links
2 | 7 | Fill in Jumper settings label
2 | 7 | Fit IPPSU 1, Record S/N
2 | 7 | Set DIP switches
2 | 7 | IGP Qualification Marks Checked
5 | 11274 | Fit and route all column cables
5 | 11274 | Fit and route objective water pipes
5 | 11274 | Safety – Torque the earth stud to 15 N.m
5 | 11274 | Torque both earth zinc plated M8 nuts to 15 N.m
5 | 11274 | Confirm HV cable is routed though plinth, and secured on top of plinth
6 | 68 | No Leaks Detected (With Turbo running) when in high sensitivity
6 | 68 | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng
6 | 68 | No Leaks Detected (With Turbo running) when in high sensitivity
6 | 68 | Monitor turbo noise level when at full speed. If >62dB with microphone underneath please see Man Eng
6 | 68 | No Leaks Detected (With Turbo running) when in high sensitivity
6 | 710 | Opening of column manual valves required to achieve gun vacuum
6 | 710 | Actual IGP1 vacuum
3 | 1266 | Fit Stage PCB
3 | 1266 | Set Stage links (348247-9002-120)
3 | 1266 | Fill in Jumper settings label
397 | 3377 | Beam is blanked when the EBBU is on
397 | 3377 | Remove EBBU Electronics and fit Blanking Cap before shipping
4 | 1262 | Fit Ducting Plate & Earth Cables
4 | 1262 | Fit Manifold
4 | 1262 | Fit Air and Nitrogen Pipes
5 | 52 | Fit Ducting Plate & Earth Cables
5 | 52 | Fit Manifold
5 | 52 | Fit Air and Nitrogen Pipes
5 | 52 | Fit Ducting Plate & Earth Cables
5 | 52 | Fit Manifold
5 | 52 | Fit Air and Nitrogen Pipes
2 | 1247 | Fit Stage PCB
2 | 1247 | Fit IPPSU 1, Record S/N
2 | 1247 | Set DIP switches
2 | 1247 | Qualification Marks Checked
2 | 7 | Fit Stage PCB
2 | 7 | Set Stage links
2 | 7 | Fill in Jumper settings label
2 | 7 | Fit IPPSU 1, Record S/N
2 | 7 | Set DIP switches
2 | 7 | IGP Qualification Marks Checked
2 | 1247 | Fit IPPSU 1, Record S/N
2 | 1247 | Set DIP switches
2 | 1247 | Qualification Marks Checked
2 | 7 | Fit Stage PCB
2 | 7 | Set Stage links
2 | 7 | Fill in Jumper settings label
2 | 1247 | Fit IPPSU 1, Record S/N
2 | 1247 | Set DIP switches
2 | 1247 | Qualification Marks Checked
2 | 7 | Fit Stage PCB
2 | 7 | Set Stage links
2 | 7 | Fill in Jumper settings label
2 | 7 | Fit Stage PCB
2 | 7 | Set Stage links (348247-9001-500)
2 | 7 | Fill in Jumper settings label
3 | 1266 | Set Stage links (348247-9002-120)
3 | 1266 | Fill in Jumper settings label
3 | 1265 | Set Stage links (348247-9001-500)
3 | 1265 | Fill in Jumper settings label
397 | 11075 | Pneumatics functional
397 | 11075 | SmartSEM Control tested
397 | 11075 | Images saved
30 | 1224 | carry out scanning calibration (Hasap only)
382 | 1225 | carry out scanning calibration (Hasap only)
2 | 1264 | Fit Stage PCB
2 | 1264 | Set Stage links (348247-9001-500)
2 | 1264 | Fill in Jumper settings label
2 | 1264 | Fit Stage PCB
2 | 1264 | Set Stage links (348247-9001-500)
2 | 1264 | Fill in Jumper settings label
2 | 1264 | Fit Stage PCB
2 | 1264 | Set Stage links (348247-9001-500)
2 | 1264 | Fill in Jumper settings label
2 | 7 | Fit Stage PCB
2 | 7 | Set Stage links (348247-9001-500)
2 | 7 | Fill in Jumper settings label
2 | 1264 | Fit Stage PCB
2 | 1264 | Set Stage links (348247-9001-500)
2 | 1264 | Fill in Jumper settings label
2 | 1264 | Fit Stage PCB
2 | 1264 | Set Stage links (348247-9001-500)
2 | 1264 | Fill in Jumper settings label
2 | 1264 | Fit Stage PCB
2 | 1264 | Set Stage links (348247-9001-00)
2 | 1264 | Fill in Jumper settings label
6 | 1547 | PAT Checkbox Calibration Date
6 | 1547 | Earth Bond Test TP1 (0.07Ω - 0.09Ω)
6 | 1547 | Earth Bond Test TP2 (0.50Ω - 0.53Ω)
6 | 1547 | Insulation Test (2.0MΩ - 1.9MΩ)
6 | 1547 | Earth Leakage Test (0.9mA - 0.7mA)
6 | 1547 | PAT Checkbox Calibration Date
6 | 1547 | Earth Bond Test TP1 (0.07Ω - 0.09Ω)
6 | 1547 | Earth Bond Test TP2 (0.50Ω - 0.53Ω)
6 | 1547 | Insulation Test (2.0MΩ - 1.9MΩ)
6 | 1547 | Earth Leakage Test (0.7mA - 0.9mA)
389 | 11074 | Check each individual segment in normal and inverse modes
389 | 11074 | Analyze Segments
389 | 11074 | Gain Ranges
389 | 11074 | Pneumatics
389 | 11074 | Alignment
389 | 11074 | Vac Degradation < 3x10-5 mbar
389 | 11074 | CZ Detector Cal
389 | 11074 | Z Adjustment
20 | 11038 | PC S/N
20 | 11038 | If SEM Enhanced PC(349535-9001-020) has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Hard disk Image (Version)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Check the Entitlement certificate to determine which version of Smartsem to install
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Approved Datazone Loaded
20 | 11038 | PC S/N
20 | 11038 | If SEM Enhanced PC(349535-9001-020) has Serial Number Before 746448, then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 11038 | PC/Model (SAP No.)
20 | 11038 | Hard disk Image (Version)
20 | 11038 | Record Operating System
20 | 11038 | Install Firewire Driver
20 | 11038 | Check the Entitlement certificate to determine which version of Smartsem to install
20 | 11038 | Install Smart SEM. Record Version
20 | 11038 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 11038 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 11038 | Install CZ Support Tools. Ver
20 | 11038 | Disable Automatic Updates
20 | 11038 | Approved Datazone Loaded
703 | 637 | 3 hour plasma clean completed
703 | 637 | 3 hour plasma clean completed
40 | 916 | Unpack Volutome and check all parts on the packing list are present (packing list link can be found in attachments)
40 | 916 | Ensure Arivis Software Licence is available (1 year or permanent)
40 | 916 | Upload data from USB stick to the O drive product folder
40 | 916 | Remove transit clamps from Volutome stage
40 | 916 | Ensure both safety switches are fitted to the Volutome door
40 | 916 | Ensure hardware has been configured according to work instructions
40 | 916 | Fit Volutome stage
40 | 916 | Check protection circuit for Piezo is fitted (349567-9013-000)
40 | 916 | Connect Volutome stage and controller to the SEM
40 | 916 | Fit Stemi holder bracket to Volutome stage door (STEMI adjustment bracket needs to fit / disconnect without force from door)
40 | 916 | Install the latest version of Zen (configured for Volutome)
40 | 916 | Install PI Micromove
40 | 916 | Ensure Athena inspector manager has all green lights, Physik microtome and Physik stage are present - save a screenshot in the Volutome folder
40 | 916 | Run getversions.exe and save text file in the Volutome folder
40 | 916 | Configure Volutome stage motors
40 | 916 | Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator
40 | 916 | Set up the Volutome workbench in Zen
40 | 916 | Record VolumeBSD serial number
40 | 916 | Record Volutome stage serial number
40 | 916 | Record VolumeBSD amplifier serial number
40 | 916 | Record VolumeBSD mechanics serial number
40 | 916 | Record C-885 PI controller serial number
40 | 916 | Record E-727 PI controller serial number
40 | 916 | Volutome 'internal LED light' function check
40 | 916 | Volutome 'External cold light' function check
40 | 916 | Adjust Volutome stage centre (Crosshairs should be within 4 central squares)
40 | 916 | Install Nav cam in-line connections (If applicable)
40 | 916 | Perform the microtome and stage functionality checks
40 | 916 | Fit and align the VolumeBSD diode
40 | 916 | Check VolumeBSD alignment repeatability
40 | 916 | Check all gain ranges of the VolumeBSD (normal/inverted)
40 | 916 | Perform Signal/noise ratio check
40 | 916 | Ensure BSD diode protector is fitted and adjusted in chamber
40 | 916 | Install cable tie to loop BSD signal cable
40 | 916 | Test BSD functionality in Zen - Insert, retract and gain ranges all functional
40 | 916 | Check SCM with Volutome stage
40 | 916 | Oscillation check - Measure the amplitude of the oscillation
40 | 916 | Perform X-ray check
40 | 916 | Pump Down Time
40 | 916 | Ultimate Vacuum
40 | 916 | Focal CC fitted to Volutome stage
40 | 916 | Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar
40 | 916 | Dummy experiment completed with no erros reported in Zen - Results saved in Volutome folder
40 | 916 | Generate live image from Axiocam
40 | 916 | Check functionality of all additional software installed e.g Atlas
40 | 916 | Re-package Volutome and check it's completeness
40 | 916 | ManEng/ Senior Eng/ Production Manger Sign off
389 | 917 | Unpack Volutome and check all parts on the packing list are present (packing list link can be found in attachments)
389 | 917 | Ensure Arivis Software Licence is available (1 year or permanent)
389 | 917 | Upload data from USB stick to the O drive product folder
389 | 917 | Remove transit clamps from Volutome stage
389 | 917 | Ensure both safety switches are fitted to the Volutome door
389 | 917 | Ensure hardware has been configured according to work instructions
389 | 917 | Fit Volutome stage
389 | 917 | Check protection circuit for Piezo is fitted (349567-9013-000)
389 | 917 | Connect Volutome stage and controller to the SEM
389 | 917 | Fit Stemi holder bracket to Volutome stage door (STEMI adjustment bracket needs to fit / disconnect without force from door)
389 | 917 | Install the latest version of Zen (configured for Volutome)
389 | 917 | Install PI Micromove
389 | 917 | Ensure Athena inspector manager has all green lights, Physik microtome and Physik stage are present - save a screenshot in the Volutome folder
389 | 917 | Run getversions.exe and save text file in the Volutome folder
389 | 917 | Configure Volutome stage motors
389 | 917 | Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator
389 | 917 | Set up the Volutome workbench in Zen
389 | 917 | Record VolumeBSD serial number
389 | 917 | Record Volutome stage serial number
389 | 917 | Record VolumeBSD amplifier serial number
389 | 917 | Record VolumeBSD mechanics serial number
389 | 917 | Record C-885 PI controller serial number
389 | 917 | Record E-727 PI controller serial number
389 | 917 | Volutome 'internal LED light' function check
389 | 917 | Volutome 'External cold light' function check
389 | 917 | Adjust Volutome stage centre (Crosshairs should be within 4 central squares)
389 | 917 | Install Nav cam in-line connections (If applicable)
389 | 917 | Sigma cartesian stage angle correction completed
389 | 917 | Perform the microtome and stage functionality checks
389 | 917 | Fit and align the VolumeBSD diode
389 | 917 | Check VolumeBSD alignment repeatability
389 | 917 | Check all gain ranges of the VolumeBSD (normal/inverted)
389 | 917 | Perform Signal/noise ratio check
389 | 917 | Ensure BSD diode protector is fitted and adjusted in chamber
389 | 917 | Install cable tie to loop BSD signal cable
389 | 917 | Test BSD functionality in Zen - Insert, retract and gain ranges all functional
389 | 917 | Check SCM with Volutome stage
389 | 917 | Oscillation check - Measure the amplitude of the oscillation
389 | 917 | Perform X-ray check
389 | 917 | Pump Down Time
389 | 917 | Ultimate Vacuum
389 | 917 | Focal CC fitted to Volutome stage
389 | 917 | Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar
389 | 917 | Dummy experiment completed with no erros reported in Zen - Results saved in Volutome folder
389 | 917 | Generate live image from Axiocam
389 | 917 | Check functionality of all additional software installed e.g Atlas
389 | 917 | Re-package Volutome and check it's completeness
389 | 917 | ManEng/ Senior Eng/ Production Manger Sign off
3 | 25 | Torque all earth cables on central ground panel including supplier fitted earths.
3 | 25 | Torque all earth cables to specified torque value.
3 | 25 | Torque all earth cables to fitted to earth terminal strip.
3 | 25 | Torque the following earth connections; DC Unit, Stage PCB,
3 | 25 | Torque all earth cables to fitted to earth terminal strip.
3 | 25 | Torque the following earth connections; DC Unit, Stage PCB, Turbo pump.
3 | 25 | Torque all earth cables to fitted to earth terminal strip.
3 | 25 | Torque the following earth connections; DC Unit, Stage PCB, Turbo pump.
389 | 3301 | Function Test
389 | 3301 | Edit Exchange/Resume Macros
8 | 92 | Cooling Isolation Check (Open Circuit Coil – GND)
20 | 252 | PC S/N
20 | 252 | PC/Model (SAP No.)
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted, Is Serial Number Before 746448
20 | 252 | If SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Has Serial Number Before 746448 , then the PC needs reimaging – follow the attached WI. The correct image version is V2.
20 | 252 | Hard disk Image (Version)
20 | 252 | Record Operating System
20 | 252 | Check the Entitlement certificate to determine which version of Smartsem to install
20 | 252 | Install Smart SEM (See WI 9693) Record Version
20 | 252 | Check for available SmartSEMService Packs and install them. (N/A if none)
20 | 252 | Check for available SmartSEM hotfixes and install them. (N/A if none)
20 | 252 | Install CZ Support Tools. Ver
20 | 252 | Disable Automatic Updates
20 | 252 | Run ASAP Screening Program
20 | 252 | System Specific Datazone Loaded
20 | 252 | In SEM status set 'Mains Sync' to off
6 | 1236 | Confirm SN label does not have any CSA markings.
6 | 1236 | Confirm SN label does not have any SEMI markings.
6 | 11345 | Confirm SN label does not have any CSA markings.
6 | 11345 | Confirm SN label does have the correct SEMI marking.
6 | 1234 | Confirm SN label does have the correct CSA marking.
6 | 1234 | Confirm SN label does not have any SEMI marking.
6 | 1235 | Confirm SN label does have the correct CSA marking.
6 | 1235 | Confirm SN label does have the correct SEMI marking.
656 | 10663 | Damper Plate is Sigma Version (Additional airlock hole)
656 | 10663 | Damper Plate is Located Correctly (Centred)
403 | 10467 | Install SmartSEM + ASAP Engine
403 | 10467 | Copy the SmartSEM sotware into a Folder on the Desktop. Name it SmartSEM 8.* Software
403 | 10467 | SmartSEM Version
403 | 10467 | SmartSEM Service Pack
403 | 10467 | SmartSEM Hotfix
403 | 10467 | ASAP Engine Version
403 | 10467 | Installation of Zen Core EM if Hasap is fitted
8 | 1274 | 20 Min plasma cycle - Ceramic disc, Bean Can, Gehause Mont
8 | 1274 | 20 Min plasma cycle - Ceramic Disc, Bean Can, Gehause Mont
8 | 1274 | 20 Min plasma cycle - Ceramic Disc, Bean Can, Gehause Mont
8 | 1274 | 20 Min plasma cycle carried out - Ceramic Disc, Bean Can, Gehause Mont
8 | 1274 | 20 min plasma cycle carried out - Ceramic Disc, Bean Can, Gehause Mont
9 | 1275 | 20 min plasma cycle carried out - Ceramic Disc, Bean Can, Gehause Mont
9 | 1275 | 20 min plasma cycle carried out - Ceramic Disc, Bean Can
376 | 3257 | In SEM status set 'Mains Sync' to off
643 | 1273 | Mains Sync set to OFF
643 | 1273 | Mains Sync set to OFF
376 | 3257 | In SEM status set 'Mains Sync' to off
9 | 1275 | 20 min plasma cycle carried out - Ceramic Disc, Bean Can
9 | 1275 | 20 min plasma cycle carried out - Ceramic Disc, Bean Can
656 | 10664 | Table Top is Correct Version
656 | 10664 | Table Top is Free from Damage
656 | 10664 | Table Top is Correct Version (Airlock punchout hole present)
656 | 10664 | Table Top is Free from Damage
376 | 3254 | F/W Version
376 | 3254 | Hardware Version
376 | 3254 | F/W Version
376 | 3254 | Hardware Version
403 | 10467 | Install SmartSEM + ASAP Engine
403 | 10467 | Copy the SmartSEM sotware into a Folder on the Desktop. Name it SmartSEM 8.* Software
403 | 10467 | SmartSEM Version
403 | 10467 | SmartSEM Service Pack
403 | 10467 | SmartSEM Hotfix
403 | 10467 | ASAP Engine Version
403 | 10467 | Installation of Zen Core EM if Hasap is fitted
403 | 1167 | Check that Time & Date on PC are set to current UK
403 | 1167 | Download Licence from SharePoint
403 | 1167 | Delete the Licence in SharePoint
403 | 1167 | Install the licence using the Zeiss License Management Tool.
403 | 1167 | Installation of ZEN Core EM licence using the Zeiss activation tool
376 | 3251 | F/W Version
376 | 3252 | F/W Version
376 | 3252 | Hardware Version
376 | 3253 | F/W Version
376 | 3253 | Hardware Version
403 | 1167 | Check that Time & Date on PC are set to current UK
403 | 1167 | Download Licence from SharePoint
403 | 1167 | Delete the Licence in SharePoint
403 | 1167 | Install the licence using the Zeiss License Management Tool.
403 | 1167 | Installation of ZEN Core EM licence using the Zeiss activation tool
398 | 1180 | Download the CEU folder and copy to the Desktop.
398 | 1180 | Install RClientinstaller.exe (if not already installed)
398 | 1180 | Run SEM_Calibration Extraction Tool.
398 | 1180 | Store the SEM_CEU.txt and Sigma***-**-**.csv files
398 | 1180 | copy ceu files to the Engineer Folder.
398 | 1180 | If the RClientinstaller has been installed for the the CEU only, then uninstall
646 | 10529 | BSD Diode Pre-test
646 | 10529 | Select Type
646 | 10529 | Select Amplifier Type
646 | 10529 | ES-BSD Only - Using ASAP Test Panel Update the ID to 336
646 | 10529 | Black Level Calibration
646 | 10529 | Check Diode Segment Function
646 | 10529 | Check All Gain ranges
646 | 10529 | Check Auto range
646 | 10529 | Check Auto Level (Does Not Function With BSD1)
646 | 10529 | Setup and Check TOPO
646 | 10529 | Take 3 images - Atomic, 5pA, 2KV
646 | 10529 | Store BSD Functional Images-Atomic, 5pA, 2KV
646 | 10529 | BSD Diode Pre-test
646 | 10529 | Select Type
646 | 10529 | Select Amplifier Type
646 | 10529 | ES-BSD Only - Using ASAP Test Panel Update the ID to 336
646 | 10529 | Black Level Calibration
646 | 10529 | Check Diode Segment Function
646 | 10529 | Check All Gain ranges
646 | 10529 | Check Auto range
646 | 10529 | Check Auto Level (Does Not Function With BSD1)
646 | 10529 | Setup and Check TOPO
646 | 10529 | Take 3 images - Atomic, 5pA, 2KV
646 | 10529 | Store BSD Functional Images-Atomic, 5pA, 2KV
646 | 10529 | BSD Diode Pre-test
646 | 10529 | Select Type
646 | 10529 | Select Amplifier Type
646 | 10529 | ES-BSD Only - Using ASAP Test Panel Update the ID to 336
646 | 10529 | Black Level Calibration
646 | 10529 | Check Diode Segment Function
646 | 10529 | Check All Gain ranges
646 | 10529 | Check Auto range
646 | 10529 | Check Auto Level (Does Not Function With BSD1)
646 | 10529 | Setup and Check TOPO
646 | 10529 | Take 3 images - Atomic, 5pA, 2KV
646 | 10529 | Store BSD Functional Images-Atomic, 5pA, 2KV
650 | 10565 | Create C:Support folder
650 | 10565 | Create a Data folder within the Support folder
650 | 10565 | In the Data folder create the following new sub-folders- Images, Licences, Log files, SmartSEM Software.
650 | 10565 | Create C:Support folder
650 | 10565 | Create a Data folder within the Support folder
650 | 10565 | In the Data folder create the following new sub-folders- Images, Licences, Log files, SmartSEM Software.
650 | 10565 | Copy the standard resolution images into theImages folder.
285 | 3242 | Select PC type
285 | 3242 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
285 | 3242 | PC S/No.
285 | 3242 | Is SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted,
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Upgrade ASAP Firmware (check Service Pack for latest version) and API
285 | 3242 | Copy the SmartSEM sotware into a Folder on the Desktop. Name it SmartSEM 8.0 Software
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Set the Time and Date on the PC to current UK
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
40 | 916 | 1. Unpack Volutome and check all parts on the packing list are present (packing list link can be found in attachments)
40 | 916 | 1. Upload data from USB stick to the O drive product folder
40 | 916 | 1. Remove transit clamps from Volutome stage
40 | 916 | 1. Ensure both safety switches are fitted to the Volutome door
40 | 916 | 2. Ensure hardware has been configured according to work instructions
40 | 916 | 2. Fit Volutome stage
40 | 916 | 2. Check protection circuit for Piezo is fitted (349567-9013-000)
40 | 916 | 2. Connect Volutome stage and controller to the SEM
40 | 916 | 2. Fit Stemi holder bracket to Volutome stage door (STEMI adjustment bracket needs to fit / disconnect without force from door)
40 | 916 | 3. Ethernet Connection Setup
40 | 916 | 4. Install the latest version of Zen - Ensure to configure it for Volutome (Refer to Zen installation WI)
40 | 916 | 4. Ensure Athena inspector manager has all green lights, Physik microtome and Physik stage are present - save a screenshot in the Volutome folder
40 | 916 | 4. Install PI Micromove
40 | 916 | 4. Run getversions.exe and save text file in the Volutome folder
40 | 916 | 4. Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator
40 | 916 | 4. Set up the Volutome workbench in Zen
40 | 916 | 5. PID parameters configured
40 | 916 | 6. Record VolumeBSD serial number
40 | 916 | 6. Record Volutome stage serial number
40 | 916 | 6. Record VolumeBSD amplifier serial number
40 | 916 | 6. Record VolumeBSD mechanics serial number
40 | 916 | 6. Record C-885 PI controller serial number
40 | 916 | 6. Record E-727 PI controller serial number
40 | 916 | 7. Volutome 'internal LED light' function check
40 | 916 | 8. Volutome 'External cold light' function check
40 | 916 | 9. Adjust Volutome stage centre (Crosshairs should be within 4 central squares)
40 | 916 | 10. Install Nav cam in-line connections (If applicable)
40 | 916 | 11. Perform the microtome and stage functionality checks
40 | 916 | 12. Fit and align the VolumeBSD diode
40 | 916 | 12. Check VolumeBSD alignment repeatability
40 | 916 | 12. Check all gain ranges of the VolumeBSD (normal/inverted)
40 | 916 | 12. Perform Signal/noise ratio check
40 | 916 | 12. Ensure BSD diode protector is fitted and adjusted in chamber
40 | 916 | 12. Install cable tie to loop BSD signal cable
40 | 916 | 12. Test BSD functionality in Zen - Insert, retract and gain ranges all functional
40 | 916 | 13. Check SCM with Volutome stage
40 | 916 | 14. Oscillation check - Check there is no visible oscillation when imaging with a pixel size of 5nm
40 | 916 | 15. Perform X-ray check
40 | 916 | 16. Pump Down Time
40 | 916 | 16. Ultimate Vacuum
40 | 916 | 17. Focal CC fitted to Volutome stage
40 | 916 | 17. Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar
40 | 916 | 18. Generate live image from Axiocam
40 | 916 | 19. Dummy experiment completed with no erros reported in Zen - Results saved in Volutome folder
40 | 916 | 20. Re-package Volutome and check it's completeness
40 | 916 | 21. Check functionality of all additional software installed e.g Atlas
40 | 916 | 22. Ensure Arivis Software Licence is available (1 year or permanent)
40 | 916 | 23. ManEng/ Senior Eng/ Production Manger Sign off
403 | 1276 | Create C:Support folder
403 | 1276 | Create a Data folder within the Support folder
403 | 1276 | In the Data folder create the following new sub-folders- Images, Licences, Log files, SmartSEM Software.
389 | 917 | 1. Unpack Volutome and check all parts on the packing list are present (packing list link can be found in attachments)
389 | 917 | 1. Upload data from USB stick to the O drive product folder
389 | 917 | 1. Remove transit clamps from Volutome stage
389 | 917 | 1. Ensure both safety switches are fitted to the Volutome door
389 | 917 | 2. Ensure hardware has been configured according to work instructions
389 | 917 | 2. Fit Volutome stage
389 | 917 | 2. Check protection circuit for Piezo is fitted (349567-9013-000)
389 | 917 | 2. Connect Volutome stage and controller to the SEM
389 | 917 | 2. Fit Stemi holder bracket to Volutome stage door (STEMI adjustment bracket needs to fit / disconnect without force from door)
389 | 917 | 3. Ethernet Connection Setup
389 | 917 | 4. Install the latest version of Zen - Ensure to configure it for Volutome (Refer to Zen installation WI)
389 | 917 | 4. Ensure Athena inspector manager has all green lights, Physik microtome and Physik stage are present - save a screenshot in the Volutome folder
389 | 917 | 4. Install PI Micromove
389 | 917 | 4. Run getversions.exe and save text file in the Volutome folder
389 | 917 | 4. Ensure 'Volutome stage fitted' and 'Volutome stage visible' boxes are ticked in administrator
389 | 917 | 4. Set up the Volutome workbench in Zen
389 | 917 | 5. PID parameters configured
389 | 917 | 6. Record VolumeBSD serial number
389 | 917 | 6. Record Volutome stage serial number
389 | 917 | 6. Record VolumeBSD amplifier serial number
389 | 917 | 6. Record VolumeBSD mechanics serial number
389 | 917 | 6. Record C-885 PI controller serial number
389 | 917 | 6. Record E-727 PI controller serial number
389 | 917 | 7. Volutome 'internal LED light' function check
389 | 917 | 8. Volutome 'External cold light' function check
389 | 917 | 9. Adjust Volutome stage centre (Crosshairs should be within 4 central squares)
389 | 917 | 10. Install Nav cam in-line connections (If applicable)
389 | 917 | 11. Perform the microtome and stage functionality checks
389 | 917 | 11. Sigma cartesian stage angle correction correct
389 | 917 | 12. Fit and align the VolumeBSD diode
389 | 917 | 12. Check VolumeBSD alignment repeatability
389 | 917 | 12. Check all gain ranges of the VolumeBSD (normal/inverted)
389 | 917 | 12. Perform Signal/noise ratio check
389 | 917 | 12. Ensure BSD diode protector is fitted and adjusted in chamber
389 | 917 | 12. Install cable tie to loop BSD signal cable
389 | 917 | 12. Test BSD functionality in Zen - Insert, retract and gain ranges all functional
389 | 917 | 13. Check SCM with Volutome stage
389 | 917 | 14. Oscillation check - Check there is no visible oscillation when imaging with a pixel size of 5nm
389 | 917 | 15. Perform X-ray check
389 | 917 | 16. Pump Down Time
389 | 917 | 16. Ultimate Vacuum
389 | 917 | 17. Focal CC fitted to Volutome stage
389 | 917 | 17. Focal CC @ 100% pressue vacuum doesnt exceed 7e-3 mBar
389 | 917 | 18. Generate live image from Axiocam
389 | 917 | 19. Dummy experiment completed with no erros reported in Zen - Results saved in Volutome folder
389 | 917 | 20. Re-package Volutome and check it's completeness
389 | 917 | 21. Check functionality of all additional software installed e.g Atlas
389 | 917 | 22. Ensure Arivis Software Licence is available (1 year or permanent)
389 | 917 | 23. ManEng/ Senior Eng/ Production Manger Sign off
651 | 10580 | SmartSEM Version
651 | 10580 | Check SmartSEM is in the folder C:Support/Data/SmartSEM Software
650 | 10565 | Copy the standard resolution images into the C:Support - Data - Images folder.
403 | 1276 | Create C:Support folder
403 | 1276 | Create a Data folder within the Support folder
403 | 1276 | In the Data folder create the following new sub-folders- Images, Licences, Log files, SmartSEM Software.
403 | 1276 | Create C:Support folder
403 | 1276 | Create a Data folder within the Support folder
403 | 1276 | In the Data folder create the following new sub-folders- Images, Licences, Log files, SmartSEM Software.
403 | 10467 | Install SmartSEM + ASAP Engine
403 | 10467 | Copy the SmartSEM sotware into the Image folder (C:Support/Data/Images)
403 | 10467 | SmartSEM Version
403 | 10467 | SmartSEM Service Pack
403 | 10467 | SmartSEM Hotfix
403 | 10467 | ASAP Engine Version
403 | 10467 | Installation of Zen Core EM if Hasap is fitted
403 | 1276 | Create C:Support folder
403 | 1276 | Create a Data folder within the Support folder
403 | 1276 | In the Data folder create the following new sub-folders- Images, Licences, Log files, SmartSEM Software.
45 | 363 | Em Server.mdb
45 | 363 | License File
45 | 363 | Saved Registry
45 | 363 | L500.ini
646 | 10652 | Type of Detector
646 | 10652 | Detector Serial number
646 | 10652 | DDP Module Serial Number
646 | 10652 | Fit Flange & Mount Detector
646 | 10652 | Install EDS PCB's Into PC
646 | 10652 | Installation of DPP Module & Cables
646 | 10652 | Configure PC Network & Test
646 | 10652 | Disable Windows10 Fast Startup
646 | 10652 | Install ApexKey.edx Licences
646 | 10652 | Install Genesis.edx Licences (if Applicable)
646 | 10652 | Fit Dongle
646 | 10652 | APEX Software Version
646 | 10652 | Install APEX Software
646 | 10652 | Install Remoting
646 | 10652 | Setup RConfigure
646 | 10652 | Run API Test
646 | 10652 | Configure Imaging-Run SGShell
646 | 10652 | Calibrate Detector-Run eDPP Shell
646 | 10652 | Save Calibration Report
646 | 10652 | Detector Resolution (measured using Cu) ≈ 129 eV
646 | 10652 | Store Calibration Report
646 | 10652 | Save Calibration on Exit
646 | 10652 | Run APEX
646 | 10652 | Check Spectrum
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected
646 | 10652 | Check Mapping
646 | 10652 | Check Line Scanning
646 | 10652 | Save Aquired Spectrum, Mappimg & Line Scanning to the Image Folder
646 | 10652 | Store Aquired Spectrum (STD, Quick, HD), Mappimg & Line Scanning Images
646 | 10652 | Save Calibration Data to the folder C: Support/Data/ Log Files
646 | 10652 | Store Calibration Data File
646 | 10652 | Save Calibration Data to the Image Folder
651 | 10907 | Atlas Hardware Installed
651 | 10907 | Atlas Software Installed
651 | 10907 | Atlas Licence Saved to the C:Support/Data/Log Files
45 | 363 | License file
45 | 363 | Em Server.mdb
45 | 363 | Saved Registry
45 | 363 | L500.ini
45 | 363 | Calibxx.dat
403 | 1276 | Create the Support folder in the root dir (c:)
403 | 1276 | Create a Data folder within the Support folder
403 | 1276 | In the Data folder create the following new sub-folders- Images, Licences, Log files, SmartSEM Software.
403 | 10467 | Install SmartSEM + ASAP Engine
403 | 10467 | Copy the SmartSEM sotware into the Image folder (C:Support/Data/SmartSEM Software)
403 | 10467 | SmartSEM Version
403 | 10467 | SmartSEM Service Pack
403 | 10467 | SmartSEM Hotfix
403 | 10467 | ASAP Engine Version
403 | 10467 | Installation of Zen Core EM if Hasap is fitted
403 | 10467 | Install SmartSEM + ASAP Engine
403 | 10467 | Copy the SmartSEM sotware into the SmartSEM Software folder (C:Support/Data/SmartSEM Software)
403 | 10467 | SmartSEM Version
403 | 10467 | SmartSEM Service Pack
403 | 10467 | SmartSEM Hotfix
403 | 10467 | ASAP Engine Version
403 | 10467 | Installation of Zen Core EM if Hasap is fitted
646 | 10652 | Type of Detector
646 | 10652 | Detector Serial number
646 | 10652 | DDP Module Serial Number
646 | 10652 | Fit Flange & Mount Detector
646 | 10652 | Install EDS PCB's Into PC
646 | 10652 | Installation of DPP Module & Cables
646 | 10652 | Configure PC Network & Test
646 | 10652 | Disable Windows10 Fast Startup
646 | 10652 | Install ApexKey.edx Licences
646 | 10652 | Install Genesis.edx Licences (if Applicable)
646 | 10652 | Fit Dongle
646 | 10652 | APEX Software Version
646 | 10652 | Install APEX Software
646 | 10652 | Install Remoting
646 | 10652 | Setup RConfigure
646 | 10652 | Run API Test
646 | 10652 | Configure Imaging-Run SGShell
646 | 10652 | Calibrate Detector-Run eDPP Shell
646 | 10652 | Save Calibration Report
646 | 10652 | Detector Resolution (measured using Cu) ≈ 129 eV
646 | 10652 | Store Calibration Report
646 | 10652 | Save Calibration on Exit
646 | 10652 | Run APEX
646 | 10652 | Check Spectrum
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with SD Quality Selected
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with Quick Quality Selected
646 | 10652 | Check Spectrum Aligns with Al/Cu Markers with HD Quality Selected
646 | 10652 | Check Mapping
646 | 10652 | Check Line Scanning
646 | 10652 | Save Aquired Spectrum, Mappimg & Line Scanning to the Image Folder
646 | 10652 | Store Aquired Spectrum (STD, Quick, HD), Mappimg & Line Scanning Images
646 | 10652 | Save Calibration Data to the folder C: Support/Data/ Log Files
646 | 10652 | Store Calibration Data File
646 | 10652 | Save Calibration Data to the SmartSEM Image Folder
650 | 10565 | Copy the standard resolution images into the C:Support/Data/Images folder.
650 | 10567 | Delete the Production Test Licences
650 | 10567 | Obtain the customer Entitlement Certificate
650 | 10567 | Download into the C:Support/Data/Log Files folder
650 | 10567 | Install the Customer Licences
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer.db to the Product Folder
650 | 10567 | Print out the Customer Entitlement Certicate - To be sent to the customer
650 | 10567 | Store EM Server.db
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
650 | 10567 | Delete the Production Test Licences
650 | 10567 | Obtain the customer Entitlement Certificate
650 | 10567 | Download into the C:Support/Data/Licence folder
650 | 10567 | Install the Customer Licences
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer.db to the Product Folder
650 | 10567 | Print out the Customer Entitlement Certicate - To be sent to the customer
650 | 10567 | Store EM Server.db
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
650 | 10567 | Delete the Production Test Licences
650 | 10567 | Obtain the customer Entitlement Certificate
650 | 10567 | Download into the C:Support/Data/Licence folder
650 | 10567 | Install the Customer Licences
650 | 10567 | Copy Contents of Image Folder to the Product Folder
650 | 10567 | Copy EMServer.db to the Product Folder
650 | 10567 | Copy EMServer.db to the C: Support/Data/Log Files folder
650 | 10567 | Print out the Customer Entitlement Certicate - To be sent to the customer
650 | 10567 | Store EM Server.db
650 | 10567 | Copy EMServer.log to the Product Folder
650 | 10567 | Store EM Server Log
650 | 10567 | Copy EMServer Screening.txt to the Product Folder
650 | 10567 | Store EM Server Screening
650 | 10568 | Calibration File Added to C: Support/Data/Log Files
650 | 10568 | Copy Calibration File to the Product Folder
650 | 10568 | Store SmartEDX Calibration File
403 | 1276 | Select the Support folder in the root dir (c:)
403 | 1276 | Create a Data folder within the Support folder
403 | 1276 | In the Data folder create the following new sub-folders- Images, Licences, Log files, SmartSEM Software.
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50
646 | 1278 | Install ZEN core EM
646 | 1278 | Upgrade Pneumatic Detectot Firmware to 3.4.0
645 | 10929 | Calibrate Stage Centre (Right Click on Stgae tab)
645 | 10929 | Copy the Centre Pos X and Y values into the Adminitrator
645 | 10929 | Tilt Centre Calibration (Located in Docking panel)
651 | 10884 | Check Aquire Spectrum
651 | 10884 | Check Line Scanning
651 | 10884 | Check Mapping
651 | 10884 | Check EDX Icon Is Present In All User Toolbars
651 | 10884 | EDAX Calibration File Saved in the folder C:Support/Data/Log Files
651 | 669 | The Customer Entitlement Certificate is present in C:Support/Data/Licence folder
651 | 669 | The EM Server.db is present in C:Support/Data/Log Files folder
651 | 1179 | Download the CEU folder and copy to the Desktop.
651 | 1179 | Install RClientinstaller.exe (if not already installed)
651 | 1179 | Run SEM_Calibration Extraction Tool.
651 | 1179 | Store the SEM_CEU.txt and EVO**-**-**.csv files
651 | 1179
651 | 1179 | If the RClientinstaller has been installed for the the CEU only, then uninstall
651 | 1179 | Copy the SEM_CEU.txt and EVO**-**-**.csv files to the C:Support/Data/Log Files folder
651 | 1179 | Download the CEU folder and copy to the Desktop.
651 | 1179 | Install RClientinstaller.exe (if not already installed)
651 | 1179 | Run SEM_Calibration Extraction Tool.
651 | 1179 | Store the SEM_CEU.txt and EVO**-**-**.csv files
651 | 1179 | Copy the SEM_CEU.txt and EVO**-**-**.csv files to the C:Support/Data/Log Files folder
651 | 1179 | If the RClientinstaller has been installed for the the CEU only, then uninstall
651 | 915 | Check that the 200V image is of a good quality.
651 | 915 | Check 1KV Resolution Image meets specification and of a good quality.
651 | 915 | Check 3KV Resolution Image meets specification and of a good quality.
651 | 915 | Check 30KV Resolution Image meets specification and of a good quality.
651 | 669 | The Customer Entitlement Certificate is present in C:Support/Data/Licence folder
651 | 669 | The EM Server.db is present in C:Support/Data/Log Files folder
651 | 669 | The Customer Entitlement Certificate is present in C:Support/Data/Licence folder
651 | 669 | The EM Server.db is present in C:Support/Data/Log Files folder
403 | 11272 | Predictive Service S/W installed on Customer PC.
403 | 11272 | Copy the Predictive Service .json file into C:Support/Data/Licence folder.
403 | 11272 | Predictive Service S/W installed on Customer PC.
403 | 11272 | Copy the Predictive Service.json file into C:Support/Data/Licence folder.
21 | 895 | Set Water Flow Meter to 0.9 lpm
651 | 1277 | Verify that all the folders and their contents have been copied
651 | 1277 | Copy the Support folder and its sub folders to the (D:) drive on the PC
651 | 1277 | Verify that all the folders and their contents have been copied
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50
646 | 1278 | Install ZEN core EM
646 | 1278 | Upgrade Pneumatic Detectot Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectord)
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50
646 | 1278 | Install ZEN core EM
646 | 1278 | Upgrade Pneumatic Detectot Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectord)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SMartSEM Toggle Button
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50
646 | 1278 | Install ZEN core EM
646 | 1278 | Upgrade Pneumatic Detectot Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectord)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50
646 | 1278 | Install ZEN core EM
646 | 1278 | Upgrade Pneumatic Detectot Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectord)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Set ZEN APIGateway Port to 5002
646 | 1278 | Configue NacCam and Setup Position $NavCam
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50
646 | 1278 | Install ZEN core EM
646 | 1278 | Upgrade Pneumatic Detectot Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectord)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Set ZEN APIGateway Port to 5002
646 | 1278 | Configue NacCam and Setup Position $NavCam (Capital N and C)
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50
646 | 1278 | Install ZEN core EM
646 | 1278 | Upgrade Pneumatic Detectot Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectord)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Set ZEN APIGateway Port to 5002
646 | 1278 | Configue NacCam and Setup Position $NavCam (Capital N and C)
646 | 1278 | Configue Airlock Cam and Setup Exchange Position £Exchange
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50
646 | 1278 | Install ZEN core EM
646 | 1278 | Upgrade Pneumatic Detectot Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectord)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Set ZEN APIGateway Port to 5002
646 | 1278 | Configue NacCam and Setup Position $NavCam (Capital N and C)
646 | 1278 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50
646 | 1278 | Install ZEN core EM
646 | 1278 | Upgrade Pneumatic Detectot Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectord)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Set ZEN APIGateway Port to 5002
646 | 1278 | Configue NacCam and Setup Position $NavCam (Capital N and C)
646 | 1278 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
646 | 1278 | Mains Sync Set To OFF
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50
646 | 1278 | Install ZEN core EM
646 | 1278 | Upgrade Pneumatic Detectot Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectord)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Set ZEN APIGateway Port to 5002
646 | 1278 | Configue NacCam and Setup Position $NavCam (Capital N and C)
646 | 1278 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
646 | 1278 | Mains Sync Set To OFF
646 | 1278 | General Athena Setting - Tick Correct NavCam if Fitted
646 | 1278 | Snap an Image and Save To Images Folder in SmartSEM
646 | 1278 | Run a 3 x3 Tile Aquisition
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50
646 | 1278 | Install ZEN core EM
646 | 1278 | Upgrade Pneumatic Detectot Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectord)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Set ZEN APIGateway Port to 5002
646 | 1278 | Configue NacCam and Setup Position $NavCam (Capital N and C)
646 | 1278 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
646 | 1278 | Mains Sync Set To OFF
646 | 1278 | General Athena Setting - Tick Correct NavCam if Fitted
646 | 1278 | Snap an Image and Save To Images Folder in SmartSEM
646 | 1278 | Run a 3 x3 Tile Aquisition
646 | 1278 | Remove the Static and Dynamic Calibration Workbenches
645 | 1226 | carry out scanning calibration (Hasap only)
651 | 10906 | ZEN Software Installed
393 | 3346 | Ensure all customer licences are loaded in the Zeiss Licence Management Tool and functions correctly
393 | 3346 | Check thst the SmartSEM 8 folder is present in the support folder
651 | 631 | VP/EP Function Checked
651 | 631 | Check Beam Is On at Max VP/EP Pressure
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50)
646 | 1278 | Install ZEN core EM
646 | 1278 | Upgrade Pneumatic Detectot Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectord)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Set ZEN APIGateway Port to 5002
646 | 1278 | Configue NacCam and Setup Position $NavCam (Capital N and C)
646 | 1278 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
646 | 1278 | Mains Sync Set To OFF
646 | 1278 | General Athena Setting - Tick Correct NavCam if Fitted
646 | 1278 | Snap an Image and Save To Images Folder in SmartSEM
646 | 1278 | Run a 3 x3 Tile Aquisition
646 | 1278 | Remove the Static and Dynamic Calibration Workbenches
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50)
646 | 1278 | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure
646 | 1278 | Upgrade Pneumatic Detectot Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectord)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Set ZEN APIGateway Port to 5002
646 | 1278 | Configue NacCam and Setup Position $NavCam (Capital N and C)
646 | 1278 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
646 | 1278 | Mains Sync Set To OFF
646 | 1278 | General Athena Setting - Tick Correct NavCam if Fitted
646 | 1278 | Snap an Image and Save To Images Folder in SmartSEM
646 | 1278 | Run a 3 x3 Tile Aquisition
646 | 1278 | Remove the Static and Dynamic Calibration Workbenches
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50)
646 | 1278 | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure
646 | 1278 | Upgrade Pneumatic Detector Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectors)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Set ZEN APIGateway Port to 5002
646 | 1278 | Configue NacCam and Setup Position $NavCam (Capital N and C)
646 | 1278 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
646 | 1278 | Mains Sync Set To OFF
646 | 1278 | General Athena Setting - Tick Correct NavCam if Fitted
646 | 1278 | Snap an Image and Save To Images Folder in SmartSEM
646 | 1278 | Run a 3 x3 Tile Aquisition
646 | 1278 | Remove the Static and Dynamic Calibration Workbenches
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50)
646 | 1278 | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure
646 | 1278 | Upgrade Pneumatic Detector Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectors)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Set ZEN API Gateway Port to 5002
646 | 1278 | Configue NacCam and Setup Position $NavCam (Capital N and C)
646 | 1278 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
646 | 1278 | Mains Sync Set To OFF
646 | 1278 | General Athena Settings - Tick Correct NavCam if Fitted
646 | 1278 | Snap an Image and Save To Images Folder in SmartSEM
646 | 1278 | Run a 3 x3 Tile Aquisition
646 | 1278 | Remove the Static and Dynamic Calibration Workbenches
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50)
646 | 1278 | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure
646 | 1278 | Upgrade Pneumatic Detector Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectors)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Set ZEN API Gateway Port to 5002
646 | 1278 | Configue NacCam and Setup Position $NavCam (Capital N and C)
646 | 1278 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
646 | 1278 | Mains Sync Set To OFF
646 | 1278 | General Athena Settings - Tick Correct NavCam if Fitted
646 | 1278 | Snap an Image and Save To Images Folder in SmartSEM
646 | 1278 | Run a 3 x3 Tile Aquisition
646 | 1278 | Remove the Static and Dynamic Calibration Workbenches
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50)
646 | 1278 | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure
646 | 1278 | Upgrade Pneumatic Detector Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectors)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Set ZEN API Gateway Port to 5002
646 | 1278 | Configue NacCam and Setup Position $NavCam (Capital N and C)
646 | 1278 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
646 | 1278 | Mains Sync Set To OFF
646 | 1278 | General Athena Settings - Tick Correct NavCam if Fitted
646 | 1278 | Snap an Image and Save To Images Folder in SmartSEM
646 | 1278 | Run a 3 x3 Tile Aquisition
646 | 1278 | Remove the Static and Dynamic Calibration Workbenches
403 | 1167 | Check that Time & Date on PC are set to current UK
403 | 1167 | Download Licence from SharePoint
403 | 1167 | Delete the Licence in SharePoint
403 | 1167 | Install the licence using the Zeiss License Management Tool.
403 | 1167 | Installation of ZEN Core EM licence using the Zeiss activation tool
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50)
646 | 1278 | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure
646 | 1278 | Install Customer's ZEN Licence
646 | 1278 | Upgrade Pneumatic Detector Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectors)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Set ZEN API Gateway Port to 5002
646 | 1278 | Configue NacCam and Setup Position $NavCam (Capital N and C)
646 | 1278 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
646 | 1278 | Mains Sync Set To OFF
646 | 1278 | General Athena Settings - Tick Correct NavCam if Fitted
646 | 1278 | Snap an Image and Save To Images Folder in SmartSEM
646 | 1278 | Run a 3 x3 Tile Aquisition
646 | 1278 | Remove the Static and Dynamic Calibration Workbenches
45 | 363 | Customer Images
45 | 363 | License file
45 | 363 | Em Server.mdb
45 | 363 | Saved Registry
45 | 363 | L500.ini
45 | 363 | Calibxx.dat
645 | 1284 | X/Y Macro Run For 30 Minutes
389 | 67 | Install ZEN core EM.Exe
389 | 67 | ZEN core Software Version
389 | 67 | Setup Exchange Position
389 | 67 | Capture Image
645 | 10929 | Calibrate Stage Centre (Right Click on Stgae tab)
645 | 10929 | Copy the Centre Pos X and Y values into the Adminitrator, Stage Post Init. Position.
645 | 10929 | Tilt Centre Calibration (Located in Docking panel)
389 | 1279 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50)
389 | 1279 | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure
389 | 1279 | Install Customer's ZEN Licence
389 | 1279 | Upgrade Pneumatic Detector Firmware to 3.4.0
389 | 1279 | Calibrate and Configure All Detectors (including legacy detectors)
389 | 1279 | Install Latest Opal Kelly Driver (5.3.3)
389 | 1279 | Setup SmartSEM Toggle Button
389 | 1279 | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
389 | 1279 | Set ZEN API Gateway Port to 5002
389 | 1279 | Configue NacCam and Setup Position $NavCam (Capital N and C)
389 | 1279 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
389 | 1279 | Mains Sync Set To OFF
389 | 1279 | General Athena Settings - Tick Correct NavCam if Fitted
389 | 1279 | Snap an Image and Save To Images Folder in SmartSEM
389 | 1279 | Run a 3 x3 Tile Aquisition
389 | 1279 | Remove the Static and Dynamic Calibration Workbenches
40 | 1280 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50)
40 | 1280 | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure
40 | 1280 | Install Customer's ZEN Licence
40 | 1280 | Upgrade Pneumatic Detector Firmware to 3.4.0
40 | 1280 | Calibrate and Configure All Detectors (including legacy detectors)
40 | 1280 | Install Latest Opal Kelly Driver (5.3.3)
40 | 1280 | Setup SmartSEM Toggle Button
40 | 1280 | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
40 | 1280 | Set ZEN API Gateway Port to 5002
40 | 1280 | Configue NacCam and Setup Position $NavCam (Capital N and C)
40 | 1280 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
40 | 1280 | Mains Sync Set To OFF
40 | 1280 | General Athena Settings - Tick Correct NavCam if Fitted
40 | 1280 | Snap an Image and Save To Images Folder in SmartSEM
40 | 1280 | Run a 3 x3 Tile Aquisition
40 | 1280 | Remove the Static and Dynamic Calibration Workbenches
40 | 1280 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50)
40 | 1280 | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure
40 | 1280 | Install Customer's ZEN Licence
40 | 1280 | Upgrade Pneumatic Detector Firmware to 3.4.0
40 | 1280 | Calibrate and Configure All Detectors (including legacy detectors)
40 | 1280 | Install Latest Opal Kelly Driver (5.3.3)
40 | 1280 | Setup SmartSEM Toggle Button
40 | 1280 | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
40 | 1280 | Set ZEN API Gateway Port to 5002
40 | 1280 | Configue NacCam and Setup Position $NavCam (Capital N and C)
40 | 1280 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
40 | 1280 | Mains Sync Set To OFF
40 | 1280 | General Athena Settings - Tick Correct NavCam if Fitted
40 | 1280 | Snap an Image and Save To Images Folder in SmartSEM
40 | 1280 | Run a 3 x3 Tile Aquisition
40 | 1280 | Remove the Static and Dynamic Calibration Workbenches
389 | 1279 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50)
389 | 1279 | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure
389 | 1279 | Install Customer's ZEN Licence
389 | 1279 | Upgrade Pneumatic Detector Firmware to 3.4.0
389 | 1279 | Calibrate and Configure All Detectors (including legacy detectors)
389 | 1279 | Install Latest Opal Kelly Driver (5.3.3)
389 | 1279 | Setup SmartSEM Toggle Button
389 | 1279 | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
389 | 1279 | Set ZEN API Gateway Port to 5002
389 | 1279 | Configue NacCam and Setup Position $NavCam (Capital N and C)
389 | 1279 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
389 | 1279 | Mains Sync Set To OFF
389 | 1279 | General Athena Settings - Tick Correct NavCam if Fitted
389 | 1279 | Snap an Image and Save To Images Folder in SmartSEM
389 | 1279 | Run a 3 x3 Tile Aquisition
389 | 1279 | Remove the Static and Dynamic Calibration Workbenches
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50)
646 | 1278 | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure
646 | 1278 | Install Customer's ZEN Licence
646 | 1278 | Upgrade Pneumatic Detector Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectors)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Set ZEN API Gateway Port to 5002
646 | 1278 | Configue NacCam and Setup Position $NavCam (Capital N and C)
646 | 1278 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
646 | 1278 | Mains Sync Set To OFF
646 | 1278 | General Athena Settings - Tick Correct NavCam if Fitted
646 | 1278 | Snap an Image and Save To Images Folder in SmartSEM
646 | 1278 | Run a 3 x3 Tile Aquisition
646 | 1278 | Remove the Static and Dynamic Calibration Workbenches
| 439 | Ion Pump Current draw < 1.0e-07 A
403 | 1167 | Check that Time & Date on PC are set to current UK
403 | 1167 | Download Licence from SharePoint
403 | 1167 | Delete the Licence in SharePoint
403 | 1167 | Install the licence using the Zeiss License Management Tool.
403 | 1167 | hASAP Only - Installation of ZEN Core EM licence using the Zeiss activation tool
| 439 | Ion Pump Current draw < 1.0e-07 A
| 768 | Cable Megger reads >1000Mohms at 1000V to ground
| 441 | S/N =
403 | 10467 | Install SmartSEM + ASAP Engine
403 | 10467 | Copy the SmartSEM sotware into the SmartSEM Software folder (C:Support/Data/SmartSEM Software)
403 | 10467 | SmartSEM Version
403 | 10467 | SmartSEM Service Pack
403 | 10467 | SmartSEM Hotfix
403 | 10467 | ASAP Engine Version
403 | 10467 | Installation of Zen Core EM if Hasap is fitted
642 | 10485 | Approx. No. of Hours
| 439 | Ion Pump Current draw < 1.0e-07 A
| 768 | Cable Megger reads >1000Mohms at 1000V to ground
| 441 | S/N =
403 | 11272 | Predictive Service S/W installed on Customer PC.
403 | 11272 | Do NOT Activate
403 | 11272 | Copy the Predictive Service.json file into C:Support/Data/Licence folder.
403 | 11272 | Predictive Service S/W installed on Customer PC.
403 | 11272 | Do Not Activate
403 | 11272 | Copy the Predictive Service.json file into C:Support/Data/Licence folder.
285 | 11271 | Predictive Service S/W installed on Customer PC.
285 | 11271 | Do Not Activate
285 | 11271 | Copy the Predictive Service.json file into C:Support/Data/Licence folder.
645 | 1283 | X Axis
645 | 1283 | Y Axis
645 | 1283 | Z Axis
645 | 1283 | T Axis
645 | 1283 | X Axis
645 | 1283 | Y Axis
645 | 1283 | Z Axis
645 | 1283 | T Axis
644 | 10656 | Vp Functional Check
644 | 10656 | Vac Status = Ready and EHT Vac Ready = Yes, are always present at all VP values. (i.e. The beam remains on/can be switched on).
645 | 1250 | Select Stage Type
645 | 1250 | Macro Result 1 ($[StageRepeatability])
645 | 1250 | Macro Result 2 ($[StageRepeatability])
645 | 1250 | Macro Result 3 ($[StageRepeatability])
645 | 1250 | Image of Macro Results
645 | 1250 | Select Stage Type
645 | 1250 | Macro Result 1 ($[StageRepeatability])
645 | 1250 | Macro Result 2 ($[StageRepeatability])
645 | 1250 | Macro Result 3 ($[StageRepeatability])
645 | 1250 | Image of Macro Results
645 | 1250 | Select Stage Type
645 | 1250 | Macro Result 1 ($[StageRepeatability])
645 | 1250 | Macro Result 2 ($[StageRepeatability])
645 | 1250 | Macro Result 3 ($[StageRepeatability])
645 | 1250 | Image of Macro Results
645 | 1250 | Macro Result 1 ($[StageRepeatability])
645 | 1250 | Macro Result 2 ($[StageRepeatability])
645 | 1250 | Macro Result 3 ($[StageRepeatability])
645 | 1250 | Image of Macro Results
645 | 1250 | Download Macro folder
645 | 1250 | Run Kompass 3.08.exe
645 | 1250 | Macro Result 1 ($[StageRepeatability])
645 | 1250 | Macro Result 2 ($[StageRepeatability])
645 | 1250 | Macro Result 3 ($[StageRepeatability])
645 | 1250 | Image of Macro Results
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50)
646 | 1278 | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure
646 | 1278 | Install Customer's ZEN Licence
646 | 1278 | Upgrade Pneumatic Detector Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectors)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Set ZEN API Gateway Port to 5002
646 | 1278 | Configue NacCam and Setup Position $NavCam (Capital N and C)
646 | 1278 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
646 | 1278 | Mains Sync Set To OFF
646 | 1278 | General Athena Settings - Tick Correct NavCam if Fitted
646 | 1278 | Snap an Image and Save To Images Folder in SmartSEM
646 | 1278 | Run a 3 x3 Tile Aquisition
646 | 1278 | Remove the Static and Dynamic Calibration Workbenches
646 | 1278 | Create User Supervisor EM
646 | 1278 | Create User Supervisor LM
646 | 1278 | Place both users as “Member of” Supervisor group
646 | 1278 | Check ASAP 12V PSU's. Read values in ASAP Test Panel(12.0 to 12.50)
646 | 1278 | Install ZEN core EM - open the hamburger menu> open the Settings> Choose “ZEISS Electron Microscopy Products”> installation tab> choose the ZEN core EM, accept the T&Cs and “Configure
646 | 1278 | Install Customer's ZEN Licence
646 | 1278 | Upgrade Pneumatic Detector Firmware to 3.4.0
646 | 1278 | Calibrate and Configure All Detectors (including legacy detectors)
646 | 1278 | Install Latest Opal Kelly Driver (5.3.3)
646 | 1278 | Setup SmartSEM Toggle Button
646 | 1278 | Select System Type and Startup Tick Box - Do Not Select the Tick Box for Volutone Stage Systems.
646 | 1278 | Create User Supervisor EM
646 | 1278 | Create User Supervisor LM
646 | 1278 | Place both users as “Member of” Supervisor group
646 | 1278 | Set ZEN API Gateway Port to 5002
646 | 1278 | Configue NacCam and Setup Position $NavCam (Capital N and C)
646 | 1278 | Configue Airlock Cam and Setup Exchange Position $Exchange (use lower case e in ZEN)
646 | 1278 | Mains Sync Set To OFF
646 | 1278 | General Athena Settings - Tick Correct NavCam if Fitted
646 | 1278 | Snap an Image and Save To Images Folder in SmartSEM
646 | 1278 | Run a 3 x3 Tile Aquisition
646 | 1278 | Remove the Static and Dynamic Calibration Workbenches
644 | 10506 | Vp Functional Check
644 | 10506 | Vac Status = Ready and EHT Vac Ready = Yes, are always present at all VP values. (i.e. The beam remains on/can be switched on).
644 | 10506 | Fit Label to the Black Plastic Aperture Box
644 | 10510 | 500um Lower Aperture Fitted
644 | 10510 | 2nd Rotary Pump is ON in EP Mode
644 | 10510 | 10Pa Chamber Vacuum Achieved
644 | 10510 | 3000Pa Vacuum Achieved (Air).
644 | 10510 | Vac Status = Ready and EHT Vac Ready = Yes, at 3000Pa (i.e. The beam remains on/can be switched on)
644 | 10510 | Water Function Checked
644 | 10510 | Approx. 2000Pa Chamber Vacuum Achieved with Water
644 | 10510 | Go To HV, 2nd Rotary Pump is OFF
644 | 10510 | 500um Lower Aperture Fitted
644 | 10510 | 2nd Rotary Pump is ON in EP Mode
644 | 10510 | 10Pa Chamber Vacuum Achieved
644 | 10510 | 3000Pa Vacuum Achieved (Air).
644 | 10510 | Vac Status = Ready and EHT Vac Ready = Yes, at 3000Pa (i.e. The beam remains on/can be switched on)
644 | 10510 | Water Function Checked
644 | 10510 | Approx. 2000Pa Chamber Vacuum Achieved with Water
644 | 10510 | Vac Status = Ready and EHT Vac Ready = Yes, at 2000Pa (i.e. The beam remains on/can be switched on)
644 | 10510 | Go To HV, 2nd Rotary Pump is OFF
646 | 10530 | Calibration of Bias
646 | 10530 | Functional Check
646 | 10530 | Function Image
646 | 10530 | Check Beam Is On at Max VP Pressure
646 | 10530 | Store VPSE Functional Image
646 | 10531 | Select Type
646 | 10531 | Black Level Cal
646 | 10531 | C2D Bias Target Check
646 | 10531 | Gain Range Check
646 | 10531 | Auto Level Check
646 | 10531 | Auto Range Check
646 | 10531 | Function Image
646 | 10531 | Check Beam Is On at Max VP Pressure
646 | 10531 | Store C2D Functional Image
642 | 10485 | Approx. No. of Hours
642 | 10485 | Approx. No. of Hours
642 | 10485 | Approx. No. of Hours
645 | 1284 | X/Y Macro Run For 30 Minutes
642 | 1285 | Set EO Temp Limit to 40 deg C
376 | 1286 | Set EO Temp Limit to 40 deg C
382 | 1287 | X/Y Macro Run For 30 Minutes
382 | 6927 | Macro run time
21 | 424 | Install Customer Quiet Mode
21 | 424 | Set Test tool Quiet Mode to 'OFF' (In service cupboard)
21 | 424 | Install customer quiet mode
21 | 424 | Set Test tool Quiet Mode to 'OFF' (In service cupboard)
642 | 1285 | Set EO Temp Limit to 40deg C
642 | 1285 | Set EO Temp Limit to 40 deg C
376 | 1286 | Set EO Temp Limit to 40 deg C
645 | 1284 | X/Y Macro Run For 30 Minutes
403 | 10467 | Install SmartSEM + ASAP Engine
403 | 10467 | Copy the SmartSEM sotware into the SmartSEM Software folder (C:Support/Data/SmartSEM Software)
403 | 10467 | SmartSEM Version
403 | 10467 | SmartSEM Service Pack
403 | 10467 | SmartSEM Hotfix
403 | 10467 | ASAP Driver Scan Control Service Version
403 | 10467 | Installation of Zen Core EM if Hasap is fitted
21 | 255 | Air (Pressure = 6 Bar)
21 | 255 | Water
21 | 255 | Confirm that the Water Flow Switch is set to 0.9 LPM
21 | 255 | Mains
21 | 255 | Safety Earth
21 | 255 | N2
21 | 255 | Ser
21 | 255 | Air (Pressure = 6 Bar)
21 | 255 | Water
21 | 255 | Confirm that the Water Flow Switch is set to 0.9 LPM
21 | 255 | Mains
21 | 255 | Safety Earth
21 | 255 | N2
21 | 255 | Ser
21 | 255 | Air (Pressure = 6 Bar)
21 | 255 | Water
21 | 255 | Confirm that the Water Flow Switch is set to 0.9 LPM
21 | 255 | Mains
21 | 255 | Safety Earth
21 | 255 | N2
21 | 255 | Ensure Test tool Quiet Mode to 'ON' in service cupboard (Unless customer Quiet mode fitted)
21 | 255 | Air (Pressure = 6 Bar)
21 | 255 | Water
21 | 255 | Confirm that the Water Flow Switch is set to 0.9 LPM
21 | 255 | Mains
21 | 255 | Safety Earth
21 | 255 | N2
21 | 255 | Ensure Test tool Quiet Mode is set to 'ON' in service cupboard (Unless customer Quiet mode fitted)
8 | 1274 | 20 min plasma cycle carried out - Ceramic Disc, Bean Can, Gehause Mont
9 | 1275 | 20 min plasma cycle carried out - Ceramic Disc, Bean Can
642 | 1285 | Set EO Temp Limit to 40 deg C
642 | 1285 | Set EO Temp Limit to 40 deg C
642 | 1285 | Set EO Temp Limit to 40 deg C
642 | 10485 | Approx. No. of Hours
642 | 10485 | Minimum of 8 Hours (Cartesian Stage)
645 | 1283 | X Axis
645 | 1283 | Y Axis
645 | 1283 | Z Axis
645 | 1283 | T Axis
645 | 1284 | X/Y Macro Run For 30 Minutes
642 | 10485 | Minimum of 8 Hours (Cartesian Stage)
645 | 1250 | Download Macro folder
645 | 1250 | Run Kompass 3.08.exe
645 | 1250 | Macro Result 1 ($[StageRepeatability])
645 | 1250 | Macro Result 2 ($[StageRepeatability])
645 | 1250 | Macro Result 3 ($[StageRepeatability])
645 | 1250 | Image of Macro Results
652 | 1288 | Fit Right Hand Foam
652 | 1288 | Fit Lefy Hand Foam
645 | 10516 | X Axis
645 | 10516 | Y Axis
645 | 10516 | X+Y Axis
645 | 10516 | Y+X Axis
645 | 10516 | 10mm Square
645 | 10517 | Z Axis
645 | 10518 | X Axis
645 | 10518 | Y Axis
652 | 1288 | Fit Right Hand Foam
652 | 1288 | Fit Lefy Hand Foam
652 | 1288 | Fit Right Hand Foam
652 | 1288 | Fit Left Hand Foam
647 | 10550 | Fit All Panels
647 | 10550 | Fit All Earth Leads
647 | 10550 | Front Column Assy. Fits Correctly
651 | 1289 | Correct Serial Number label fitted
651 | 1289 | Correct Serial Number label fitted and correct.
651 | 1289 | Correct Serial Number label fitted and correct.
651 | 1290 | Cladding is secured correctly and not able to move.
651 | 1290 | Safety stickers are applied correctly and are the correct language.
651 | 1290 | Ensure paint finish matches and is free from chips and scratches
651 | 1290 | Cladding is secured correctly and not able to move.
651 | 1290 | Safety stickers are applied correctly and are the correct language.
651 | 1290 | Ensure paint finish matches and is free from chips and scratches
651 | 1290 | Cladding is secured correctly and not able to move.
651 | 1290 | Cladding Front Hood Aasembly Fits Correctly.
651 | 1290 | Safety stickers are applied correctly and are the correct language.
642 | 10470 | Enter Syatem Serial number into Administrator
642 | 10470 | Enter System Serial number into Administrator
642 | 10470 | Check the Correct Serial number is on the Rear Silver S/N Label
651 | 1289 | Correct Serial Number label fitted and correct.
642 | 10470 | Enter System Serial number into Administrator.
642 | 10470 | Check the Correct Serial number is on the Rear Silver S/N Label
285 | 3242 | Select PC type
285 | 3242 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
285 | 3242 | PC S/No.
285 | 3242 | Is SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted,
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Upgrade ASAP Firmware (check Service Pack for latest version) and API
285 | 3242 | Copy the SmartSEM sotware into a Folder on the Desktop. Name it SmartSEM 8.0 Software
285 | 3242 | Installation of Zen Core EM if Hasap is fitted
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Set the Time and Date on the PC to current UK
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
285 | 3242 | Select PC type
285 | 3242 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
285 | 3242 | PC S/No.
285 | 3242 | Is SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted,
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Upgrade ASAP Firmware (check Service Pack for latest version) and API
285 | 3242 | Copy the SmartSEM sotware into a Folder on the Desktop. Name it SmartSEM 8.0 Software
285 | 3242 | Installation of Zen Core EM if Hasap is fitted
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Set the Time and Date on the PC to current UK
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables
403 | 1167 | Check that Time & Date on PC are set to current UK
403 | 1167 | Download Licence from SharePoint
403 | 1167 | Delete the Licence in SharePoint
403 | 1167 | Install the licence using the Zeiss License Management Tool.
403 | 1167 | hASAP Only - Installation of ZEN Core EM licence using the Zeiss activation tool
382 | 6931 | Drift
6 | 64 | Start SmartSEM and confirm there are no errors in EM server.
6 | 64 | Initialise Stage can be initialised with no errors.
6 | 64 | Water flow test - Confirm SmartSEM reports 'Water loss error' when water switched off at services.
6 | 921 | Record Water Flow Temperature (20ºC +/- 5ºC)
6 | 921 | Water flow test - Confirm SmartSEM reports 'Water loss error' when water switched off at services.
6 | 64 | Start SmartSEM and confirm there are no errors in EM server.
6 | 64 | Initialise Stage can be initialised with no errors.
678 | 1232 | System Vacuum Reading - Confirm system vacuum reads < 1x10-5mbar
678 | 1232 | Gun Vacuum Reading OK - Confirm gun vacuum <3x10-6 mbar (If vacuum showing no improvment please contact ME)
678 | 1232 | Initialise Stage - Confirm stage initialises without errors
678 | 1232 | Water flow test - Confirm SmartSEM reports 'Water loss error' when water switched off at services.
2 | 5 | Fit the pre-assembled Air and Water assembly
2 | 5 | Loctite 243 applied to water fittings (See WI Page **)
2 | 5 | Fit the pre-assembled Air and Water assembly
2 | 5 | Loctite 243 applied to water fitting nuts on back panel.
12 | 162 | Tip Re-measured Height Avg: Z =
12 | 162 | Actual Tip Height: Y - Z =
12 | 162 | Tip Gap: Actual Tip Height – 0.75 =
12 | 162 | Final Shim Value (Please ensure to always round up the shim value) =
12 | 162 | Denka Torque Setting =
285 | 3242 | Select PC type
285 | 3242 | Exchange USB card (Renesas USB card must be exchanged with ASMedia card, see WI for details)
285 | 3242 | PC S/No.
285 | 3242 | Is SEM Basic PC(349535-9001-010) or SEM Enhanced PC(349535-9001-020) Fitted,
285 | 3242 | Hard Drive 1 Size
285 | 3242 | Hard Drive 2 Size
285 | 3242 | RAM
285 | 3242 | Processor
285 | 3242 | Operating System
285 | 3242 | Change Desktop Settings
285 | 3242 | Install SmartSEM and record Version
285 | 3242 | Check for Service Packs and install them
285 | 3242 | Check for Hotfixes and install them
285 | 3242 | Upgrade ASAP Firmware (check Service Pack for latest version) and API
285 | 3242 | Copy the SmartSEM sotware into a Folder in C:Support and Name it SmartSEM X.x (i.e 8.2) Software
285 | 3242 | Installation of Zen Core EM if Hasap is fitted
285 | 3242 | Set BIOS AC Power loss to ‘ON’
285 | 3242 | Set the Time and Date on the PC to current UK
285 | 3242 | Stick PC Configuration Sheet to Side of PC
285 | 3242 | Attach Cable Markers to USB Cables